## Sheet1 (hidden)
Atellica® CH;Atellica® IM; – OUS;Global;
Platform | SMN | Catalog Number | Index | GTIN | Availability | Discipline/Disease state | Alliance Application | Product Name | Abbreviation Product/Test Name | TDef Code | Number of Tests | Kit Configuration | Materials Required But Not Provided | Optional Materials | TTFR (min) Time To First Result | Reagent On Board Stability (days) For CH assay - per well | Reagent Material Origin | Sample Types | Sample Stability | Sample Volume | Sample Pretreatment | Dilution Factor | Assay Principle | Analytical Measuring Range (Common Units) | Analytical Measuring Range (SI Units) | Extended Range (Common Units) | Extended Range (SI Units) | Expected Values (Common Units) | Expected Values (SI Units) | Expected Values #1 | Expected Values #2 | Expected Values #3 | Calibrator Name | Calibrator SMN | Calibrator Open Vial Stability | Calibrator (Lyophilized/Liquid) | Calibrator Preparation | Calibration: Number of replicates | Levels used in Calibration curve | Calibrator Material Origin | Lot Calibration Interval (days) | Pack Calibration Interval (days) | Calibration Interval (days) | IMT | Diluent | Automatic Dilution | Sample Diluent | Diluent SMN | Dilution Factor2 | Reportable Range With Dilution (Common Units) | Reportable Range With Dilution (SI Units) | Siemens QC avaiable? | Quality Control Name | Quality Control SMN | Analytical Sensitivity (Common Units) | Analytical Sensitivity (SI Units) | Functional Sensitivity (Common Units) | Functional Sensitivity (SI Units) | LoD (Common Units) | LoB (Common Units) | LoQ (Common Units) | LoB (SI Units) | LoD (SI Units) | LoQ (SI Units) | Mean Low Value Repeatability (Common Units) | Mean Low Value Repeatability (SI Units) | Low Repeatability (%) | Low Within-Lab Precision (%) | Mean Mid Value Repeatability (Common Units) | Mean Mid Value Repeatability (SI Units) | Mid Repeatability (%) | Mid Within-Lab Precision (%) | Mean High Value Repeatability (Common Units) | Mean High Value Repeatability (SI Units) | High Repeatability (%) | High Within-Lab Precision (%) | Method Comparison (Common Units) | Method Comparison (SI Units) | Hemolysis | Icterus | Lipemia | Biotin | Hook Effect (Common Units) | Hook Effect (SI Units) | High-Dose Hook Effect (Common Units) | High-Dose Hook Effect (SI Units) | Standardization Reference Method/Material | IFU Reference | IFU Document Link | Weight (in LB) | Length (in IN) | Heigth (in IN) | Width (in IN) | Weight (in KG) | Length (in CM) | Heigth (in CM) | Width (in CM)
Atellica® CH | 11097619 |  | pl22 | 00630414595610 | Global | Specific Proteins | No | Atellica CH Immunoglobulin A_2 (IgA_2) | IgA_2IgA_2 | 0907 | 600 | 4 x 150 | N/A | N/A | 10 | 30 | antihuman IgA (goat), lot specific | ;#Serum;#Lithium Heparin Plasma;# | up to 3 days at room temperature up to 7 days Frozen up to 6 months at -20 °C | 4 μL | No |  | PEG-enhanced immunoturbidimetric | 33.0 - 540.0 mg/dL | 0.33 - 5.40 g/L | 15.0-5,400.0 mg/dL | 0.15-54.00 g/L | 40.0-350.0 mg/dL for adults | 0.40-3.50 g/L for adults | N/A | N/A | N/A | Atellica CH LSP CAL | 11099434 | 28 days | Liquid | Ready to Use | 3 | 6 | Proteins derived from human serum | 60 | 30 |  | No |  | Yes | NA | 11099300 |  | 15.0-5,400.0 mg/dL | 0.15-54.00 g/L | No | NA | NA | 1.8 mg/dL | 0.6 mg/dL | NA | NA | 1.8 mg/dL | 0.6 mg/dL | NA | 0.01 g/L | 0.02 g/L | NA | 141.7 mg/dL | 1.42 g/L | 1.3 | 1.9 | 266.8 mg/dL | 2.67 g/L | 0.5 | 0.7 | 443.7 mg/dL | 4.44 g/L | 0.6 | 1 | ADVIA Chemistry 1800 IgA_2 y = 0.99x - 5.9 mg/dL r=1.000 n=80 | ADVIA Chemistry 1800 IgA_2 y = 0.99x - 0.06 g/L r=1.000 n=80 | <10% up to 1000 mg/dL of hemoglobin | <10% up to 50 mg/dL of bilirubin | <10% up to 922 mg/dL of Trig; <10% up to 500 mg/dL of Intralipid. | No Biotin in assay formulation | 9000.0 mg/dL | 90.00 g/L | 9,000.0 mg/dL | 90.00 g/L | IRMM reference material CRM-470 | 11110182_EN Rev. 03, 2019-05 | IFU Document | 0.572 | 5.512 | 3.543 | 2.756 | 0.26 | 14 | 9 | 7
Atellica® CH | 11097616 |  | pl24 | 00630414596266 | Global | Specific Proteins | No | Atellica CH Immunoglobulin G_2 (IgG_2) | IgG_2IgG_2 | 0910 | 720 | 4 x 180 | N/A | N/A | 10 | 30 | antihuman IgG (goat), lot-specific | ;#Serum;#Lithium Heparin Plasma;# | up to 7 days at 2-8 °C Frozen up to 3 months at -20°C | 2 μL | No |  | PEG-enhanced immunoturbidimetric | 140 - 3400 mg/dL | 1.40 – 34.00 g/L | 70-34,000 mg/dL | 0.70-340.00 g/L | 650-1,600 mg/dL for adults | 6.50-16.00 g/L for adults | N/A | N/A | N/A | Atellica CH LSP CAL | 11099434 | 28 days | Liquid | Ready to Use | 3 | 6 | Proteins derived from human serum | 60 | 30 |  | No |  | Yes | NA | 11099300 |  | 70-34,000 mg/dL | 0.70-340.00 g/L | No | NA | NA | 21 mg/dL | 7 mg/dL | NA | NA | 21 mg/dL | 7 mg/dL | NA | 0.07 g/L | 0.21 g/L | NA | 806 mg/dL | 8.06 g/L | 1.6 | 1.8 | 1112 mg/dL | 11.12 g/L | 1.2 | 1.4 | 2309 mg/dL | 23.09 g/L | 0.8 | 1.6 | ADVIA Chemistry 1800 IgG_2 y = 0.99x + 4 mg/dL r = 0.999 n = 80 | ADVIA Chemistry 1800 IgG_2 y = 0.99x + 0.04 g/L)\ r = 0.999 n = 80 | <10% up to 1000 mg/dL of hemoglobin | <10% up to 50 mg/dL of bilirubin | <10% up to 1000 mg/dL of Intralipid. | No Biotin in assay formulation | 70,000 mg/dL | 700 g/L | 70,000 mg/dL | 700 g/L | IRMM reference material CRM-470 | 11110179_EN Rev. 03, 2019-05 | IFU Document | 0.614 | 5.625 | 3.5 | 2.75 | 0.279 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097620 |  | pl04 | 00630414595627 | Global | Specific Proteins | No | Atellica CH Immunoglobulin M_2 (IgM_2) | IgM_2IgM_2 | 0911 | 720 | 4 x 180 | N/A | N/A | 10 | 30 | antihuman IgM (goat), lot-specific | ;#Serum;#Lithium Heparin Plasma;#Potassium EDTA Plasma;# | up to 7 days at 2-8 °C Frozen up to 3 months at -20°C | 9 μL | No |  | PEG-enhanced immunoturbidimetric | 21.0–330.0 mg/dL | 0.21–3.30 g/L | 8.0-3,300.0 mg/dL | 0.08-33.00 g/L | 50.0-300.0 mg/dL for adults | 0.50-3.00 g/L for adults | N/A | N/A | N/A | Atellica CH LSP CAL | 11099434 | 28 days | Liquid | Ready to Use | 3 | 6 | Proteins derived from human serum | 60 | 30 |  | No |  | Yes | NA | 11099300 |  | 8.0-3,300.0 mg/dL | 0.08-33.00 g/L | No | NA | NA | 3.9 mg/dL | 1.9 mg/dL | NA | NA | 3.9 mg/dL | 1.9 mg/dL | NA | 0.02 g/L | 0.04 g/L | NA | 85.8 mg/dL | 0.86 g/L | 1.1 | 2.5 | 151.9 mg/dL | 1.52 g/L | 0.6 | 1.8 | 260.5 mg/dL | 1.52 g/L | 0.6 | 1.3 | ADVIA Chemistry 1800 IGM_2 y = 1.01x + 2.5 mg/dL r = 0.999 n = 80 | ADVIA Chemistry 1800 IGM_2 y = 1.01x + 0.02 g/L r = 0.999 n = 80 | <10% up to 1000 mg/dL of hemoglobin | <10% up to 60 mg/dL of bilirubin | <10% up to 1000 mg/dL of Intralipid. | No Biotin in assay formulation | 7,000.0 mg/dL | 70.00 g/L | 7,000.0 mg/dL | 70.00 g/L | IRMM reference material ERMDA-470 | 11110183_EN Rev. 03, 2019-05 | IFU Document | 0.605 | 5.625 | 3.5 | 2.75 | 0.275 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097601 |  | ge33 | 00630414596402 | Global | General Chemistry | No | Atellica CH Iron_2 (Iron_2) | Iron_2Iron_2 | 0912 | 1792 | 4 x 448 | N/A | N/A | 10 | 30 | Acetate buffer; Ferrozine; ascorbic acid; guanidinium chloride | ;#Serum;#Lithium Heparin Plasma;# | up to 4 days at 25 °C, up to 7 days at 2-8 °C or stored frozen up to 60 days at -20 °C | 25 μL | No |  | Ferrozine | 2 - 1000 µg/dL | 0.4 - 179.0 µmol/L | 2000 μg/dL | 358.0 μmol/L | Male serum&nbsp; 65-175 μg/dL Female serum: 50-170 μg/dL | Male serum:11.6 - 31.3 μmol/L Female serum; 9.0 - 30.4 μmol/L | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 days | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 180 | 30 |  | No |  | Yes | N/A | 11099300 |  | 2 - 2000 μg/dL | 0.4 - 358.0 μmol/L | No | N/A | N/A | 2 μg/dL | 0.4 μmol/L | N/A | N/A | 2 μg/dL | 1 μg/dL | N/A | 0.2 μmol/L | 0.4 μmol/L | N/A | 62 μg/dL | 11.1 μmol/L | 0.8 | 2.3 | 253 μg/dL | 45.3 μmol/L | 0.3 | 0.6 | 813 μg/dL | 145.5 μmol/L | 0.3 | 0.4 | ADVIA Chemistry 1800 Iron_2 y = 0.99x - 2 μg/dL r=0.994 n=103 | ADVIA Chemistry 1800 Iron_2 y = 0.99x - 0.4 μmol/L r=0.994 n=103 | N/A | <10% up to 50 mg/dL of bilirubin | <10% up to 500 mg/dL of Intralipid. | No Biotin in assay formulation | N/A | N/A | N/A | N/A | NIST SRM 937 | 11110164_EN Rev. 02, 2019-07 | IFU Document | 0.977 | 5.625 | 3.5 | 2.75 | 0.444 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097614 |  | ge35 | 00630414596280 | Global | General Chemistry | No | Atellica CH Lactate (Lac) | LacLac | 1201 | 380 | 4 x 95 | N/A | N/A | 4 | 30 per pack | PIPES buffer, TOOS, Peroxidase (plant), lactate oxidase and ascorbate oxidase (Microorganism);4-aminoantipyrine | ;#Lithium Heparin Plasma;#Potassium EDTA Plasma;#Sodium fluoride/potassium oxalate Plasma;# | up to 1 day at 2-8 °C or stored frozen up to 30 days at -20 °C | 4 μL | No |  | lactate oxidase to pyruvate and hydrogen peroxide | 0.6 - 110.0 mg/dL | 0.07 - 12.21 mmol/L | N/A | N/A | 4.5-19.8 mg/dL | 0.50-2.20 mmol/L | N/A | N/A | N/A | Atellica CH SPCL CHEM CAL | 11099438 | 7 days | Lyophilized | 5.0 mL of distilled water; Stand for 30 minutes at room temperature; Invert to mix Prior to use mix the contents by gently inverting the vials | 3 | 2 | Human serum | 62 | 30 |  | No |  | No | N/A | 11099300 |  | N/A | N/A | No | N/A | N/A | 0.6 mg/dL | 0.07 mmol/L | N/A | N/A | 0.6 mg/dL | 0.4 mg/dL | n/a | 0.04 mmol/L | 0.07 mmol/L | n/a | 12.5 mg/dL | 1.39 mmol/L | 1.9 | 4.9 | 48.7 mg/dL | 5.41 mmol/L | 0.9 | 4.8 | 99.0 mg/dL | 10.99 mmol/L | 0.8 | 4.6 | ADVIA Chemistry 1800 Lac y = 1.01x - 0.2 mg/dL r=0.995 n =101 | ADVIA Chemistry 1800 Lac y = 1.01x - 0.02 mmol/L r=0.995 n =101 | <10% up to 300 mg/dL of hemoglobin | <10% up to 3.75 mg/dL of bilirubin | <10% up to 1000 mg/dL of Intralipid. | No Biotin in assay formulation | N/A | N/A | N/A | N/A | Internal standard using highly purified material | 11110177_EN Rev. 02, 2019-07 | IFU Document | 0.442 | 5.625 | 3.5 | 2.75 | 0.201 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097606 |  | ge36 | 00630414596358 | Global | General Chemistry | No | Atellica CH Lipase (Lip) | LipLip | 1205 | 1280 | 4 x 320 | N/A | N/A | 8 | 9 | TAPS; -tartaric acid; sodium hydroxide; colipase (porcine); 2-Propanol; DGGMR; | ;#Serum;#Lithium Heparin Plasma;# | up to 24 hours at room temperature, up to 7 days at 2-8 °C or stored Frozen up to a year at -20 °C or colder | 3 μL | No |  | Colorimetric rate | 8 - 700 U/L | 8 - 700 U/L | 9- 3500 U/L | 9- 3500 U/L | 12 - 53 U/L | 12 - 53 U/L | N/A | N/A | N/A | Atellica CH ENZ 1 CAL | 11099317 | 30 days | Liquid | No preparation required | 3 | 2 | Bovine Serum, (lipase (porcine pancreas) | 61 | 9 |  | No |  | Yes | N/A | 11099300 |  | 9- 3500 U/L | 9- 3500 U/L | No | N/A | N/A | 6 U/L | 3 U/L | N/A | N/A | 6 U/L | 3 U/L | N/A | 3 U/L | 6 U/L | N/A | 36 U/L | 36 U/L | 3.1 | 4.1 | 204 U/L | 204 U/L | 0.9 | 2.8 | 655 U/L | 655 U/L | 0.6 | 2.5 | ADVIA Chemistry 1800 y = 0.97x - 1 U/L r=1.000 | ADVIA Chemistry 1800 y = 0.97x - 1 U/L r=1.000 | <10% up to 200 mg/dL of hemoglobin | <10% up to 25 mg/dL of bilirubin | <10% up to 1000 mg/dL of Intralipid. | No Biotin in assay formulation | N/A | N/A | N/A | N/A | Internal standard | 11110169_EN Rev. 03, 2019-06 | IFU Document | 0.862 | 5.625 | 3.5 | 2.75 | 0.392 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097612 |  | ge37 | 00630414595597 | Global | General Chemistry | No | Atellica CH Magnesium (Mg) (serum/plasma) | MgMg | 1301 | 1200 | 3 x 400 | N/A | N/A | 8 | 14 | Xylidyl blue | ;#Serum;#Lithium Heparin Plasma;#Urine;# | Up to 7 days at 2-8 °C; Frozen up to 12 months at -20 °C or colder | 12 μL | No |  | Xylidyl Blue | 0.50 - 5.00 mg/dL | 0.21 - 2.06 mmol/L | 10.00 mg/dL | 4.11 mmol/L | Adult 1.60 - 2.60 mg/dL | Adult 0.66 - 1.07 mmol/L | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 days | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 180 | 3 |  | No |  | Yes | Atellica CH Diluent | 11099300 |  | 0.50 - 10.00 mg/dL | 0.21 - 4.11 mmol/L | No | N/A | N/A | 0.02 mg/dL | 0.01 mmol/L | N/A | N/A | 0.02 mg/dL | 0.00 mg/dL | 0.46 mg/dL | 0.00 mmol/L | 0.01 mmol/L | 0.19 mmol/L | 0.78 mg/dL | 0.32 mmol/L | 3 | 3.9 | 1.51 mg/dL | 0.62 mmol/L | 2.3 | 3.4 | 4.22 mg/dL | 1.74 mmol/L | 0.6 | 1.1 | Dimension y = 0.94x + 0.09 mg/dL (y = 0.94x + 0.04 mmol/L) r=0.996 n=108 | Dimension y = 0.94x + 0.09 mg/dL (y = 0.94x + 0.04 mmol/L) r=0.996 n=108 | <10% up to 750 mg/dL of hemoglobin | <10% up to 30 mg/dL of bilirubin | <10% up to 500 mg/dL of Intralipid. | No Biotin in assay formulation | N/A | N/A | N/A | N/A | Atomic absorption reference method/ NIST materials | 11110175_EN Rev. 02, 2019-06 | IFU Document | 0.502 | 4.187 | 3.5 | 2.75 | 0.228 | 10.636 | 8.89 | 6.985
Atellica® CH | 11097612 |  | ge37 | 00630414595597 | Global | General Chemistry | No | Atellica CH Magnesium (Mg) (urine) | MgMg | 1301 | 1200 | 3 x 400 | N/A | N/A | 8 | 14 | Xylidyl blue | ;#Serum;#Lithium Heparin Plasma;#Urine;# | Up to 7 days at 2-8 °C: | 12 μL | No |  | Xylidyl Blue | Urine 1.00 - 14.00 mg/dL | 0.41 - 5.75 mmolL | 28.00 mg/dL | 11.51 mmol/L | Adult: 24 - 255 mg/24 hours | Adult: 0.99 - 10.45 mmol/24 hours | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 days | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 180 | 3 |  | No |  | Yes | Atellica CH Diluent | 11099300 |  | 1.00 - 28.00 mg/dL | 0.41 - 11.51 mmol/L | No | N/A | N/A | 0.04 mg/dL | 0.02 mmol/L | N/A | N/A | 0.04 mg/dL | 0.00 mg/dL | 0.57 mg/dL | 0.00 mmol/L | 0.02 mmol/L | 0.23 mmol/L | 4.61 mg/dL | 1.89 mmol/L | 0.8 | 2.1 | 11.19 mg/dL | 4.60 mmol/L | 1 | 1.3 | N/A | N/A |  |  | Dimension y = 0.96x - 0.06 mg/dL (y = 0.96x - 0.02 mmol/L) r=0.998 n=100 | Dimension y = 0.96x - 0.06 mg/dL (y = 0.96x - 0.02 mmol/L) r=0.998 n=100 | N/A | N/A | N/A | No Biotin in assay formulation | N/A | N/A | N/A | N/A | Atomic absorption reference method/ NIST materials | 11110175_EN Rev. 02, 2019-06 | IFU Document | 0.502 | 4.187 | 3.5 | 2.75 | 0.228 | 10.636 | 8.89 | 6.985
Atellica® CH | 11097531 |  | ge42 | 00630414595818 | Global | General Chemistry | No | Atellica CH Total Bilirubin_2 (TBil_2) | TBil_2TBil_2 | 0207 | 1792 | 4 x 448 | N/A | N/A | 10 | 30 | sodium metavanadate | ;#Serum;#Lithium Heparin Plasma;# | Protect daylight and fluorescent light 4 °C and analyzed within 5 days Stored frozen up to 3 months at -70 °C | 14.3 μL | No |  | Vanadate oxidation | 0.15 - 35.0 mg/dL | 3 - 599 µmol/L | 70.0 mg/dL | 1197 μmol/L | 0–1 day: <8.0 mg/dL 1–2 days: <12.0 mg/dL 3–5 days: <16.0 mg/dL > 5 days–60 years: 0.3–1.2 mg/dL 60–90 years: 0.2–1.1 mg/dL >90 years: 0.2–0.9 mg/dL | 0–1 day: <137 μmol/L 1–2 days: <205 μmol/L 3–5 days: <274 μmol/L > 5 days–60 years: 5 – 21 μmol/L 60–90 years: 3 –19 μmol/L >90 years: 3 – 15 μmol/L | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 8 hours | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 60 | 30 |  | No |  | Yes | N/A | 11099300 |  | 0.15 - 70.0 mg/dL | 3 - 1197 μmol/L | No | N/A | N/A | 0.0 mg/dL | 0 μmol/L | N/A | N/A | 0.0 mg/dL | 0.0 mg/dL | 0.1 mg/dL | 0 μmol/L | 0 μmol/L | 1.71 μmol/L | 0.8 mg/dL | 14 μmol/L | 2.1 | 6.3 | 1.5 mg/dL | 26 μmol/L | 1.1 | 3.3 | 25.2 mg/dL | 431 μmol/L | 0.2 | 1 | ADVIA Chemistry 1800 TBiL_2 y = 0.96x + 0.0 mg/dL r=1.000 n=109 see IFU for Neonatal | ADVIA Chemistry 1800 TBiL_2 y = 0.96x + 0 μmol/L r=1.000 n=109 see IFU for Neonatal | <10% up to 1000 mg/dL of hemoglobin | N/A | ≤10% up to 750 mg/dL of Lipemia (Triglycerides concentrate) | No Biotin in assay formulation | N/A | N/A | N/A | N/A | AACC reference method which uses reference materials from the National Institute of Standards and Technology (NIST SRM 916 | 11110149_EN Rev. 04, 2020-08 | IFU Document | 0.942 | 5.625 | 3.5 | 2.75 | 0.428 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097525 |  | ge07 | 00630414595535 | Global | General Chemistry | No | Atellica CH Total Iron Binding Capacity (TIBC) | TIBCTIBC | 2016 | 800 | 4 x 200 | N/A | N/A | 10 | 7 | Chromazurol B; cetrimide; ; ferric chloride; acetate buffer; Sodium bicarbonate buffer | ;#Serum;# | Separated specimens may be stored for up to 4 days at room temperature or for up to 7 daysat 2–8°C or stored frozen for up to 2 months at -20°C. | 24 μL | No |  | Sequential release and uptake of iron | 40 – 670 μg/dL | 7.16 –119.93 μmol/L | N/A | N/A | 250–425 μg/dL . | 44.75–76.08 μmol/L. | N/A | N/A | N/A | Atellica CH SPCL CHEM CAL | 11099438 | 7 | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Human serum | 180 | 7 |  | No |  | No | N/A | 11099300 |  | N/A | N/A | No | N/A | N/A | 9 μg/dL | 1.61 μmol/L | N/A | N/A | 9 μg/dL | 6 μg/dL | N/A | 1.07 μmol/L | 1.61 μmol/L | N/A | 222 μg/dL | 39.74 μmol/L | 0.8 | 1.3 | 370 μg/dL | 66.23 μmol/L | 0.4 | 0.7 | 536 μg/dL | 95.94 μmol/L | 0.4 | 0.6 | ADVIA Chemistry TIBC (serum) y = 1.05x - 20 μg/dL r=0.995 n=137 | ADVIA Chemistry TIBC (serum) y = 1.05x - 3.58 μmol/L r=0.995 n=137 | <10% up to 1000 mg/dL of hemoglobin | <10% up to 25 mg/dL of bilirubin | <10% up to 1000 mg/dL of Trig; <10% up to 375 mg/dL of Intralipid. | No Biotin in assay formulation | N/A | N/A | N/A | N/A | internal standard. | 11110143_EN Rev. 02, 2019-11 | IFU Document | 0.678 | 5.625 | 3.5 | 2.75 | 0.308 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097604 |  | ge44 | 00630414596372 | Global | General Chemistry | No | Atellica CH Total Protein II (TP) | TPTP | 2017 | 7400 | 4 x 1850 | N/A | N/A | 10 | 90 | Sodium hydroxide; Na-K-tartrate;potassium iodide; cupric sulfate | ;#Serum;#Lithium Heparin Plasma;# | up to 8 hours at room temperature, up to 3 days at 2-8 °C or stored frozen up to 180 days at -20 °C | 17.5 μL | No |  | biuret reagent (cupric sulfate in an alkaline solution) | 2.0 - 12.0 g/dL | 20 - 120 g/L | 24.0 g/dL | 240 g/L | 5.7-8.2 g/dL | 57-82 g/L | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 days | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 181 | 30 |  | No |  | Yes | N/A | 11099300 |  | N/A | N/A | No | N/A | N/A | 0.7 g/dL | 7 g/L | N/A | N/A | 0.7 g/dL | 0.6 g/dL | N/A | 6 g/L | 7 g/L | N/A | 4.0 g/dL | 40 g/L | 0.8 | 1 | 7.7 g/dL | 77 g/L | 1.7 | 1.7 | 9.6 g/dL | 96 g/L | 0.5 | 1.3 | ADVIA Chemistry 1800 Total Protein II y = 0.98x + 0.1 g/dL r=0.997 n=100 | ADVIA Chemistry 1800 Total Protein II y = 0.98x + 10 g/L r=0.997 n=100 | <10% up to 500 mg/dL of hemoglobin | <10% up to 25 mg/dL of bilirubin | <10% up to 500 mg/dL of Intralipid. | No Biotin in assay formulation | N/A | N/A | N/A | N/A | NIST SRM 927 | 11110167_EN Rev. 03, 2020-08 | IFU Document | 1.223 | 5.625 | 3.5 | 2.75 | 0.556 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097613 |  | pl34 | 00630414596297 | Global | Specific Proteins | No | Atellica CH Transferrin (Trf) | TrfTrf | 2018 | 880 | 4 x 220 | N/A | N/A | 10 | 30 | Antihuman transferrin antibody (goat) | ;#Serum;#Lithium Heparin Plasma;#Potassium EDTA Plasma;# | Separated specimens store for up to 8 hours at room temperature and up to 7 days at 2–8°C Frozen for up to 1 month at -20°C or 1 year at -70°C. | 2 μL | No |  | PEG-enhanced immunoturbidimetric | 1 - 440 mg/dL | 0.01 - 4.40 g/L | 4400 mg/dL | 44.00 g/L | Males Serum/Plasma 215-365 mg/dL Females Serum/Plasma 250-380 mg/dL | Males Serum/Plasma 2.15–3.65 g/L Females Serum/Plasma 2.50–3.80 g/L | N/A | N/A | N/A | Atellica CH LSP CAL | 11099434 | 28 days | Liquid | Ready to Use | 3 | 6 | Proteins derived from human serum | 60 | 14 |  | No |  | Yes | NA | 11099300 |  | 1 - 4400 mg/dL | 0.01 - 44.00 g/L | No | NA | NA | 1.0 mg/dL | 0.01 g/L | NA | NA | 1.0 mg/dL | 0 mg/dL | NA | 0.00 g/L | 0.01 g/L | NA | 130 mg/dL | 1.3 g/L | 0.7 | 1.9 | 257mg/dL | 2.57 g/L | 0.8 | 1.2 | 322 mg/dL | 3.22 g/L | 0.9 | 1.3 | ADVIA Chemistry 1800 TRF y = 0.98x - 3 mg/dL r = 0.997 n = 80 | ADVIA Chemistry 1800 TRF y = 0.98x - 0.03 g/L r = 0.997 n = 80 | <10% up to 525 mg/dL of hemoglobin | <10% up to 30 mg/dL of bilirubin | <10% up to 650 mg/dL of Intralipid. | No Biotin in assay formulation | 7200 mg/dL | 72.00 g/L | 7200 mg/dL | 72.00 g/L | Institute for Reference Materials and Measurements (IRMM) reference material CRM-470 from IFCC. | 11110176_EN Rev. 03, 2019-04 | IFU Document | 0.662 | 5.625 | 3.5 | 2.75 | 0.301 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097591 |  | ge43 | 00630414596495 | Global | General Chemistry | No | Atellica CH Triglycerides (concentrated) (Trig) | TrigTrig | 2020 | 2000 | 4 x 500 | N/A | N/A | 4 | 30 | 4‑Chlorophenol; peroxidase (plant) ; glycerol kinase (microorganism); glycerol‑3‑phosphate oxidase (microorganism); 4‑aminophenazone; lipase (microorganism) | ;#Serum;#Lithium Heparin Plasma;#Potassium EDTA Plasma;# | up to 7 days at 2-8 °C; Frozen up to 3 months at -15 to -25 °C | 4 μL | No |  | GPO | 10 - 550 mg/dL | 0.11 - 6.22 mmol/L | 1100 mg/dL | 12.43 mmol/L | Adult Serum Normal:150 mg/dL Borderline high: 150-199 mg/dL High: 200-499 mg/dL Very high: ≥ 500 mg/dL | Adult Serum: Normal: 1.70 mmol/L Borderline high: 1.70-2.25 mmol/L High: 2.26-5.64 mmol/L Very high: ≥ 5.65 mmol/L | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 days | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 60 | 14 |  | No |  | Yes | N/A | 11099300 |  | 10 - 1100 mg/dL | 0.11 - 12.43 mmol/L | No | N/A | N/A | 4 mg/dL | 0.05 mmol/L | N/A | N/A | 4 mg/dL | 3 mg/dL | 10 mg/dL | 0.03 mmol/L | 0.05 mmol/L | 0.11 mmol/L | 106 mg/dL | 1.2 mmol/L | 0.8 | 2.5 | 228 mg/dL | 2.58 mmol/L | 0.5 | 1.3 | 424 mg/dL | 4.79 mmol/L | 0.5 | 1 | ADVIA Chemistry 1800 TRIG_2 y = 0.98x + 0.5 mg/dL r=0.999 n=100 | ADVIA Chemistry 1800 TRIG_2 y = 0.98x + 0.01 mmol/L r=0.999 n=100 | <10% up to 1000 mg/dL of hemoglobin | <10% up to 30 mg/dL of bilirubin | N/A | No Biotin in assay formulation | N/A | N/A | N/A | N/A | NIST reference material SRM909c | 11110154_EN Rev. 03, 2019-11 | IFU Document | 0.365 | 2.875 | 3.5 | 2.75 | 0.166 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097608 |  | ge45 | 00630414596334 | Global | General Chemistry | No | Atellica CH Uric Acid (UA) | UAUA (serum) | 2102 | 4800 | 4 x 1200 | N/A | N/A | 7 | 30 | N-ethyl-N-(2-hydroxy-3-sulfopropyl)-3-methyl-aniline; 4-aminophenazone; peroxidase (plant); uricase (microorganism) | ;#Serum;#Lithium Heparin Plasma;# | Serum and plasma: up to 3–5 days at 4°C or Frozen for up to 6 months at -20°C. | 11 μL | No |  | Uricase/Peroxidase | 0.5 - 20 mg/dL | 30 - 1190 µmol/L | 100.0 mg/dL | 5950 μmol/L | Serum/Plasma Male: 3.7–9.2 mg/dL Female: 3.1–7.8 mg/dL | Serum/Plasma Male: 220–547 μmol/L Female: 184–464 μmol/L | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 days | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 183 | 7 |  | No |  | Yes | N/A | 11099300 |  | 0.5 - 100.0 mg/dL | 30 - 5950 μmol/L | No | N/A | N/A | 0.0 mg/dL | 0.0 μmol/L | N/A | N/A | 0.0 mg/dL | 0.0 mg/dL | N/A | 0.0 μmol/L | 0.0 μmol/L | N/A | 3.2 mg/dL | 190 μmol/L | 1 | 1.5 | 10.6 mg/dL | 631 μmol/L | 0.3 | 0.6 | 16.0 mg/dL | 952 μmol/L | 1.2 | 2 | ADVIA Chemistry 1800 UA y = 0.99x + 0.0 mg/dL r=1.000 n=100 | ADVIA Chemistry 1800 UA y = 0.99x + 0 μmol/L r=1.000 n=100 | <10% up to 250 mg/dL of hemoglobin | <10% up to 30 mg/dL of bilirubin | <10% up to 1000 mg/dL of Trig. | No Biotin in assay formulation | N/A | N/A | N/A | N/A | CDC candidate reference method which uses SRM 913 and SRM 909 reference materials from the National Institute of Standards and Technology | 11110171_EN Rev. 03, 2019-05 | IFU Document | 1.049 | 5.625 | 3.5 | 2.75 | 0.477 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097608 |  | ge45 | 00630414596334 | Global | General Chemistry | No | Atellica CH Uric Acid (UA) | UAUA (urine) | 2102 | 4800 | 4 x 1200 | N/A | N/A | 7 | 30 | N-ethyl-N-(2-hydroxy-3-sulfopropyl)-3-methyl-aniline; 4-aminophenazone; peroxidase; uricase | ;#Urine;# | 3–4 days at ambient temperature for alkaline urine. | 11 μL | No |  | Uricase/Peroxidase | 0.9 - 180 mg/dL | 54 - 10710 µmol/L | 360.0 mg/dL | 21420 μmol/L | Urine: 250–750 mg/day | Urine: 1.48–4.43 mmol/day | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 days | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 183 | 7 |  | No |  | Yes | N/A | 11099300 |  | 0.9 - 360.0 mg/dL | 54 - 21420 μmol/L | No | N/A | N/A | 0.5 mg/dL | 30 μmol/L | N/A | N/A | 0.5 mg/dL | 0.3 mg/dL | N/A | 18 μmol/L | 30 μmol/L | N/A | 12.9 mg/dL | 768 μmol/L | 2 | 2.9 | 21.1 mg/dL | 1255 μmol/L | 1.2 | 1.6 | N/A |  |  |  | ADVIA Chemistry 1800 UA y = 1.01x - 1.0 mg/dL r=0.997 n=100 | ADVIA Chemistry 1800 UA y = 1.01x - 60 μmol/L r=0.997 n=100 | N/A | N/A | N/A | No Biotin in assay formulation | N/A | N/A | N/A | N/A | CDC candidate reference method which uses SRM 913 and SRM 909 reference materials from the National Institute of Standards and Technology | 11110171_EN Rev. 03, 2019-05 | IFU Document | 1.049 | 5.625 | 3.5 | 2.75 | 0.477 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097593 |  | ge04 | 00630414596471 | Global | General Chemistry | No | Atellica CH Urea Nitrogen (UN_c) (serum/plasma) | UN_cUN_c (serum/plasma) | 2101 | 6240 | 4 x 1560 | N/A | N/A | 7 | 90 | NADH; Urease (plant); α‑ketoglutarate; GLDH (bovine) | ;#Serum;#Lithium Heparin Plasma;# | Separated blood urea nitrogen in separated serum plasma up to 3-5 days at room temperature or up to 7 days at 4 °C or stored frozen indefinitely at -20 °C; urine stored up to 4 days at 4-8 °C when preserved with thymol | 7 μL | No |  | Urease with GLDH | 5 - 150 mg/dL | 1.8 - 53.6 mmol/L | 300 mg/dL | 107.1 mmol/L | Adult 9-23 mg/dL | Adult: 3.2-8.2 mmol/L | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 days | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 75 | 6 |  | No |  | Yes | N/A | 11099300 |  | 5 - 300 mg/dL | 1.8 - 107.1 mmol/L | No | N/A | N/A | 2 mg/dL | 0.7 mmol/L | N/A | N/A | 2 mg/dL | 1 mg/dL | N/A | 0.4 mmol/L | 0.7 mmol/L | N/A | 16 mg/dL | 5.7 mmol/L | 3.1 | 3.4 | 67 mg/dL | 23.9 mmol/L | 0.7 | 1.6 | 118 mg/dL | 42.1 mmol/L | 0.8 | 2.5 | ADVIA Chemistry 1800 UN y = 0.96x + 2 mg/dL r=1.000 n=100 | ADVIA Chemistry 1800 UN y = 0.96x + 0.7 mmol/L r=1.000 n=100 | <10% up to 500 mg/dL of hemoglobin | <10% up to 30 mg/dL of bilirubin | <10% up to 650 mg/dL of Intralipid. | No Biotin in assay formulation | N/A | N/A | N/A | N/A | traceable to the CDC reference method, which uses SRM 912 and 909 reference materials from the National Institute of Standards and Technology (NIST). | 11110156_EN Rev. 03, 2020-08 | IFU Document | 1.21 | 5.625 | 3.5 | 2.75 | 0.55 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097593 |  | ge04 | 00630414596471 | Global | General Chemistry | No | Atellica CH Urea Nitrogen (UN_c) (urine) | UN_cUN_c (urine) | 2101 | 6240 | 4 x 1560 | N/A | N/A | 7 | 90 | NADH; Urease (plant); α‑ketoglutarate; GLDH (bovine) | ;#Urine;# | Separated blood urea nitrogen in separated serum plasma up to 3-5 days at room temperature or up to 7 days at 4 °C or stored frozen indefinitely at -20 °C; urine stored up to 4 days at 4-8 °C when preserved with thymol | 7 μL | No |  | Urease with GLDH | 35 - 1000 mg/dL | 12.5 - 357.0 mmol/L | 2000 mg/dL | 714 mmol/L | Adult 12-20 g/day | Adult: 0.43-0.71 mol/day | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 days | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 75 | 6 |  | No |  | Yes | N/A | 11099300 |  | 35 - 2000 mg/dL | 12.5 - 714 mmol/L | No | N/A | N/A | 7 mg/dL | 2.5 mmol/L | N/A | N/A | 7 mg/dL | 4 mg/dL | N/A | 1.4 mmol/L | 2.5 mmol/L | N/A | 158 mg/dL | 56.4 mmol/L | 1.4 | 3.8 | 538 mg/dL | 192.1 mmol/L | 1.5 | 2.9 | 815 mg/dL | 291.0 mmol/L | 0.9 | 2.5 | ADVIA Chemistry 1800 UN y = 0.95x + 1 mg/dL r=0.997 n=100 | ADVIA Chemistry 1800 UN y = 0.95x + 0.4 mmol/L r=0.997 n=100 | N/A | N/A | N/A | No Biotin in assay formulation | N/A | N/A | N/A | N/A | traceable to the CDC reference method, which uses SRM 912 and 909 reference materials from the National Institute of Standards and Technology (NIST). | 11110156_EN Rev. 03, 2020-08 | IFU Document | 1.21 | 5.625 | 3.5 | 2.75 | 0.55 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097610 |  | di01 | 00630414596310 | Global | General Chemistry | No | Atellica CH Microalbumin_2 (µALB_2) | μALB_2µALB_2 | 1308 | 840 | 4 x 210 | N/A | N/A | 10 | 30 | Antihuman albumin (goat - ploclonal | ;#Urine;# | 14 days at 2–8°C 5 months at -20°C | 13.7 μL | No |  | PEG-enhanced immunoturbidimetric | 0.3 - 38.0 mg/dL | 3 - 380 mg/L | 0.3 - 380.0 mg/dL | 3.0 - 3800 mg/L | < 30 mg/day | < 30 mg/day | N/A | N/A | N/A | Atellica CH μALB_2 CAL | 11099435 | 60 days | Liquid | Mix by gentle inversion at least five (5) times to ensure homogeneity prior to use. | 3 | 6 | Human serum albumin in buffered solution; sodium azide (0.09%) | 180 | 30 |  | No |  | Yes | N/A | 11099300 |  | 0.3 - 380.0 mg/dL | 3.0 - 3800 mg/L | No | N/A | N/A | 0.1 mg/dL | 1.0 mg/L | N/A | N/A | 0.1 mg/dL | 0.0 mg/dL | NA | 0.0 mg/L | 1.0 mg/L | N/A | 0.3 mg/dL | 30 mg/L | 1.4 | 3.6 | 29.2 mg/dL | 292 mg/L | 1 | 2 | 40.9 mg/dL | 409 mg/L | 1.2 | 1.4 | ADVIA Chemistry μALB_2 (urine) y = 1.05x + 0.1 mg/dL r = 0.999 n= 112 | ADVIA Chemistry μALB_2 (urine) y = 1.05x + 1.0 mg/L r = 0.999 n= 112 | ≤ 10% change up to 500 mg/dL of hemoglobin. | N/A | N/A | No Biotin in assay formulation | 20,000 mg/dL | 200,000 mg/L | N/A | N/A | Internal Standard | 11110173_EN Rev. 03, 2019-06 | IFU Document | 0.724 | 5.625 | 3.5 | 2.75 | 0.329 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097628 |  | sp02 | 00630414596228 | Global | Specific Proteins | No | Atellica CH Alpha-1-antitrypsin (AAT) | AATAAT | 1110 | 360 | 2 x 180 | N/A | N/A | 10 | 30 | Antihuman alpha-1-antitrypsin (goat) (polyclonal) | ;#Serum;#Lithium Heparin Plasma;# | 7 days at 4°C 3 months at -70°C. | 2.7 μL | No |  | PEG enhanced immunoturbidimetric | 5–500 mg/dL | 0.05–5.00 g/L | 1000 mg/dL | 10.0 g/L | 78–200 mg/dL for adults | 0.78–2.00 g/L for adults | N/A | N/A | N/A | Atellica CH Liquid Specific Protein Calibrators (LSP CAL) | 11099434 | 28 days | Liquid | No preparation required | 3 | 6 | Proteins derived from human serum; sodium azide | 60 | 7 |  | No |  | Yes | Atellica CH Diluent | 11099300 |  | 5-1000 mg/dL | 0.05-10.0 g/L | No | N/A | N/A | 3 mg/dL | 0.03 g/L | N/A | N/A | 3 mg/dL | 1.0 mg/dL | N/A | 0.01 g/L | 0.03 g/L | N/A | 76 mg/dL | 0.76 g/L | 0.9 | 2.1 | 177 mg/dL | 1.77 g/L | 0.7 | 2.7 | 352 mg/dL | 3.52 g/L | 1 | 1.8 | Serum: ADVIA Chemistry 1800 y = 0.99x - 4.0 mg/dL | Serum: ADVIA Chemistry 1800 y = 0.99x - 0.04 g/L | ≤ 10% change up to 1000 mg/dL of hemoglobin. | ≤ 10% change up to 25 mg/dL of bilirubin | ≤ 10% change up to 125 mg/dL of Intralipid. | No Biotin in assay formulation | 10,500 mg/dL | 105.0 g/L | 10,500 mg/dL | 105.0 g/L | Institute for Reference Materials and Measurements (IRMM) ERM‑DA 470 | 11110191_EN Rev. 03, 2019-04 | IFU Document | 0.319 | 2.875 | 3.5 | 2.75 | 0.145 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097625 |  | pl11 | 00630414596242 | Global | Specific Proteins | No | Atellica CH Apolipoprotein A-1 (APO A1) | APO A1APO A1 | 1120 | 300 | 2 x 150 | N/A | N/A | 10 | 14 | Antihuman Antihuman APO A1 (goat); Polyethylene glycol; Tris Buffer, sodium azide; | ;#Serum;#Lithium Heparin Plasma;# | 7 days at 2–8°C 3 months at -20°C. | 5 μL | No |  | PEG-enhanced immunoturbidimetric | 15 - 200 mg/dL | 0.15 - 2.00 g/L | 400 mg/dL | 4.00 g/L | Serum/Plasma: Adult Males: 79–169 mg/dL Adult Females: 76–214 mg/dL | Serum/Plasma: Adult Males: 0.79–1.69 g/L Adult Females:0.76–2.14 g/L | N/A | N/A | N/A | Atellica CH Apolipoprotein A‑1 & B Calibrator (APO A1 & B CAL) | 11099329 | 28 days | Lyophilized | Add 1.0 mL of distilled water into each vial. Replace the stopper. Let the vials stand for 30 min. at room temperature; Gently swirl and invert the vials. Users manually dilute the reconstituted calibrator 1:4 with isotonic saline | 3 | 5 | Lyophilized human serum | 60 | 14 |  | No |  | Yes | Atellica CH Diluent | 11099300 |  | 15-400 mg/dL | 0.15 - 4.00 g/L | No | N/A | N/A | 1.0 mg/dL | 0.01 g/L | N/A | N/A | 1.0 mg/dL | 0.0 mg/dL | N/A | 0.00 g/L | 0.01 g/L | N/A | 83 mg/dL | 0. 83 g/L | 1 | 2.4 | 121 mg/dL | 1.21 g/L | 0.8 | 2.2 | 166 mg/dL | 1.66 g/L | 0.7 | 1.7 | ADVIA Chemistry 1800 y =1.05x - 1.0 mg/dL r = 0.989 n = 50 | ADVIA Chemistry 1800 y =1.05x - 0.014 g/L r=0.989 | ≤ 10% change up to 525 mg/dL of hemoglobin. | ≤ 10% change up to 30 mg/dL of bilirubin | ≤ 10% change up to 650 mg/dL of Intralipid. | No Biotin in assay formulation | 2,205 mg/dL | 22.05 g/L | 2,205 mg/dL | 22.05 g/L | WHO-IFCC reference material SP1-01 | 11110188_EN Rev. 02, 2019-06 | IFU Document | 0.317 | 2.875 | 3.5 | 2.75 | 0.144 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097626 |  | pl01 | 00630414596235 | Global | Specific Proteins | No | Atellica CH Apolipoprotein B (APO B) | APO BAPO B | 1122 | 300 | 2 x 150 | N/A | N/A | 10 | 30 | Antihuman Antihuman APO B (goat); Polyethylene glycol; Tris Buffer, sodium azide; | ;#Serum;#Lithium Heparin Plasma;# | 7 days at 2–8°C 3 months at -20°C | 5 μL | No |  | PEG-enhanced immunoturbidimetric | 15 - 200 mg/dL | 0.15 - 2.00 g/L | 400 mg/dL | 4.00 g/L | Serum/ Plasma: Males: 46–174 mg/dL Females:46–142 mg/dL | Serum/ Plasma: Males: 0.46–1.74g/L Females: 0.46–1.42g/L | N/A | N/A | N/A | Atellica CH Apolipoprotein A‑1 & B Calibrator (APO A1 & B CAL) | 11099329 | 28 days | Lyophilized | Add 1.0 mL of distilled water; let the vials stand for 30 minutes at room temperature; gently swirl and invert the vials to ensure homogeneit. Reconstituted calibrator requires a 1:4 manual dilution (saline) prior to use. | 3 | 5 | Lyophilized human serum | 60 | 30 |  | No |  | Yes | Atellica CH Diluent | 11099300 |  | 15 - 400 mg/dL | 0.15 - 4.00 g/L | No | N/A | N/A | 2.0 mg/dL | 0.02 g/L | N/A | N/A | 2.0 mg/dL | 1.0 mg/dL | N/A | 0.01 g/L | 0.02 g/L | N/A | 61 mg/dL | 0. 61g/L | 1.6 | 6.1 | 120 mg/dL | 1.20 g/L | 0.9 | 3.5 | 158 mg/dL | 1.58 g/L | 0.6 | 2.8 | ADVIA Chemistry 1800 y = 0.99x - 4.0 mg/dL r = 0.999 | ADVIA Chemistry 1800 y = 0.99x - 0.04 g/L r = 0.999 | ≤ 10% change up to 525 mg/dL of hemoglobin. | ≤ 10% change up to 30 mg/dL of bilirubin | ≤ 10% change up to 650 mg/dL of Intralipid. | No Biotin in assay formulation | 5600 mg/dL | 56.0 g/L | 5600 mg/dL | 56.0 g/L | WHO-IFCC reference material SP3-07 | 11110189_EN Rev. 02, 2019-05 | IFU Document | 0.31 | 2.875 | 3.5 | 2.75 | 0.141 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097634 |  | sp03 | 00630414596167 | Global | Specific Proteins | No | Atellica CH Anti-Streptolysin-O_2 (ASO_2) | ASO_2ASO_2 | 0149 | 1200 | 3 x 400 | N/A | N/A | 10 | 30 | Latex particles coated with streptolysin-O antigen (microorganism), Glycerine buffer; | ;#Serum;#Lithium Heparin Plasma;#Potassium EDTA Plasma;# | 2 days at 2–8°C 6 months at -20°C | 7.5 μL | No |  | latex enhanced immunoturbidimetric | 25.0 - 1000 IU/mL | 25.0 - 1000 IU/mL | 2000 IU/mL | 2000 IU/mL | Children (preschool) ≤ 100.0 IU/mL SErum/ Plasma: Children (school age)≤ 250.0 IU/mL Adults≤ 200.0 IU/mL Siemens-determined adult reference range ≤ 194.0 IU/mL | Children (preschool) ≤ 100.0 IU/mL SErum/ Plasma: Children (school age)≤ 250.0 IU/mL Adults≤ 200.0 IU/mL Siemens-determined adult reference range ≤ 194.0 IU/mL | N/A | N/A | N/A | Atellica CH Liquid Specific Protein Calibrators (LSP CAL) | 11099434 | 28 days | Liquid | No preparation required | 3 | 2 | Proteins derived from human serum; sodium azide | 60 | 30 |  | No |  | Yes | Atellica CH Diluent | 11099300 |  | 25.0 - 2000 IU/mL | 25.0 - 2000 IU/mL | No | N/A | N/A | 12.7 IU/mL | 12.7 IU/mL | N/A | N/A | 12.7 IU/mL | 5.9 IU/mL | N/A | 5.9 IU/mL | 12.7 IU/mL | N/A | 177.5 IU/mL | 177.5 IU/mL | 2.7 | 3 | 402.1 IU/mL | 402.1 IU/mL | 2.1 | 2.3 | 607.1 IU/mL | 607.1 IU/mL | 1.1 | 2.4 | ADVIA Chemistry 1800 y = 1.01x - 1.0 IU/mL r= 0.999 n=100 | ADVIA Chemistry 1800 y = 1.01x - 1.0 IU/mL r= 0.999 | ≤ 10% change up to 1000 mg/dL of hemoglobin. | ≤ 10% change up to 60 mg/dL of bilirubin | ≤ 10% change up to 1000 mg/dL of Intralipid. | No Biotin in assay formulation | 8500 IU/mL. | 8500 IU/mL. | 8500 IU/mL. | 8500 IU/mL. | WHO 1st international reference material for ASO | 11110197_EN Rev. 03, 2019-11 | IFU Document | 0.777 | 4.187 | 3.5 | 2.75 | 0.353 | 10.636 | 8.89 | 6.985
Atellica® CH | 11097592 |  | ge29 | 00630414596488 | Global | General Chemistry | No | Atellica CH Glucose Hexokinase_3 (GluH_3) | GluH_3GluH_3 (serum/ plasma) | 0704 | 6240 | 4 x 1560 | N/A | N/A | 7 | 62 | hexokinase (microbial); G6PDH (microbial) | ;#Serum;#Lithium Heparin Plasma;#Potassium EDTA Plasma;#Sodium fluoride/potassium oxalate Plasma;# | 72 hours at 4°C variable stability is observed with longer storage conditions | 3.4 μL | No |  | Hexokinase | 4 -700 mg/dL | 0.2 - 38.9 mmol/L | 2100 mg/dL | 116.6 mmol/L | Adults; 74–106 mg/dL Newborn; 40–60 mg/dL Newborn; 1 day 50–80 mg/dL Child 60–100 mg/dL | Adults; 4.1–5.9 mmol/L Newborn 2.2–3.3 mmol/L Newborn; 1 day; 2.8–4.4 mmol/L Child; 3.3–5.6 mmol/L | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 days | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 182 | 30 |  | No |  | Yes | N/A | 11099300 |  | 4 - 2100 mg/dL | 0.2 - 116.6 mmol/L | No | N/A | N/A | 1.0 mg/dL | 0.1 mmol/L | N/A | N/A | 1.0 mg/dL | 0.0 mg/dL | N/A | 0.0 mmol/L | 0.1 mmol/L | N/A | 59 mg/dL | 3.3 mmol/L | 0.8 | 2 | 90 mg/dL | 5.0 mmol/L | 0.7 | 1.4 | 292 mg/dL | 16.2 mmol/L | 0.5 | 1.1 | ADVIA Chemistry 1800 y = 1.03x - 1.0 mg/dL y = 1.03x - 0.1 mmol/L r=1.000 n=105 | ADVIA Chemistry 1800 y = 1.03x - 1.0 mg/dL y = 1.03x - 0.1 mmol/L r=1.000 n=105 | ≤ 10% change up to 1000 mg/dL of hemoglobin. | ≤ 10% change up to 30 mg/dL of bilirubin | ≤ 10% change up to 1000 mg/dL of Intralipid. | No Biotin in assay formulation | N/A | N/A | N/A | N/A | NIST SRM 965a | 11110155_EN Rev. 04, 2020-08 | IFU Document | 1.188 | 5.512 | 3.543 | 2.756 | 0.54 | 14 | 9 | 7
Atellica® CH | 11097592 |  | ge29 | 00630414596488 | Global | General Chemistry | No | Atellica CH Glucose Hexokinase_3 (GluH_3) | GluH_3GluH_3 (urine) | 0704 | 6240 | 4 x 1560 | N/A | N/A | 7 | 62 | hexokinase (microbial); G6PDH (microbial) | ;#Urine;# | 4°C at collection | 3.4 μL | No |  | Hexokinase | 4 -700 mg/dL | 0.2 - 38.9 mmol/L | 2100 mg/dL | 116.6 mmol/L | < 0.5 g/day | 2.8 mmol/day | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 days | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 182 | 30 |  | No |  | Yes | N/A | 11099300 |  | 4 - 2100 mg/dL | 0.2 - 116.6 mmol/L | No | N/A | N/A | 0.0 mg/dL | 0.0 mmol/L | N/A | N/A | 0.0 mg/dL | 0 mg/dL | N/A | 0.0 mmol/L | 0.0 mmol/L | N/A | 38 mg/dL | 2.1 mmol/L | 1.9 | 3.8 | 285 mg/dL | 15.8 mmol/L | 0.4 | 1.3 | N/A | N/A |  |  | ADVIA Chemistry 1800 y = 1.04x - 1.0 mg/dL y = 1.04x - 0.1 mmol/L r=1.000 n=100 | ADVIA Chemistry 1800 y = 1.04x - 1.0 mg/dL y = 1.04x - 0.1 mmol/L r=1.000 n=100 | N/A | N/A | N/A | No Biotin in assay formulation | N/A | N/A | N/A | N/A | NIST SRM 965a | 11110155_EN Rev. 04, 2020-08 | IFU Document | 1.188 | 5.512 | 3.543 | 2.756 | 0.54 | 14 | 9 | 7
Atellica® CH | 11097592 |  | ge29 | 00630414596488 | Global | General Chemistry | No | Atellica CH Glucose Hexokinase_3 (GluH_3) | GluH_3GluH_3 (CSF) | 0704 | 6240 | 4 x 1560 | N/A | N/A | 7 | 62 | hexokinase (microbial); G6PDH (microbial) | ;#Cerebrospinal fluid (CSF);# | analyze immediately centrifuge and store at 4°C or -20°C if delay is unavoidable | 3.4 μL | No |  | Hexokinase | 4 -700 mg/dL | 0.2 - 38.9 mmol/L | 2100 mg/dL | 116.6 mmol/L | Adult: 40-70 mg/dL Infant/child 60-80 mg/dL | Adult 2.2-3.9 mmol/L Infant/child3.3-4.4 mmol/L | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 days | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 182 | 30 |  | No |  | Yes | N/A | 11099300 |  | 4 - 2100 mg/dL | 0.2 - 116.6 mmol/L | No | N/A | N/A | 2.0 mg/dL | 0.1 mmol/L | N/A | N/A | 2.0 mg/dL | 1 mg/dL | N/A | 0.1 mmol/L | 0.1 mmol/L | N/A | 42 mg/dL | 2.3 mmol/L | 2 | 2.3 | 96 mg/dL | 5.3 mmol/L | 0.7 | 1.2 | N/A | N/A |  |  | ADVIA Chemistry 1800 y = 1.03x + 2.0 mg/dL y = 1.03x + 0.1 mmol/L r= 1.000 n=110 | ADVIA Chemistry 1800 y = 1.03x + 2.0 mg/dL y = 1.03x + 0.1 mmol/L r= 1.000 n=110 | N/A | N/A | N/A | No Biotin in assay formulation | N/A | N/A | N/A | N/A | NIST SRM 965a | 11110155_EN Rev. 04, 2020-08 | IFU Document | 1.188 | 5.512 | 3.543 | 2.756 | 0.54 | 14 | 9 | 7
Atellica® CH | 11097621 |  | ge31 | 00630414595634 | Global | General Chemistry | No | Atellica CH Glucose Oxidase (GluO) | GluOGluO (serum/plasma) | 0703 | 5600 | 4 x 1400 | N/A | N/A | 10 | 51 | glucose oxidase (microbial); ; peroxidase (horseradish) | ;#Serum;#Lithium Heparin Plasma;# | 8 hours at 25°C 72 hours at 4°C3 days at room temperature by addition of sodium iodoacetate or sodium fluoride | 4 μL | No |  | Glucose Oxidase Trinder | 6 -750 mg/dL | 0.3 - 41.6 mmol/L | 1,500 mg/dL | 83.3 mmol/L | 74-106 mg/dL | 4.1-5.9 mmol/L | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 days | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 35 | 7 |  | No |  | Yes | N/A | 11099300 |  | 6 - 1,500 mg/dL | 0.3 - 83.3 mmol/L | No | N/A | N/A | 1 mg/dL | 0.1 mmol/L | N/A | N/A | 1 mg/dL | 0 mg/dL | N/A | 0.0 mmol/L | 0.1 mmol/L | N/A | 61 mg/dL | 3.4 mmol/L | 1 | 2.3 | 261 mg/dL | 14.5 mmol/L | 0.3 | 2 | 567 mg/dL | 31.5 mmol/L | 0.3 | 2.9 | ADVIA Chemistry 1800 GLUO y = 0.99x - 3.0 mg/dL r = 0.999 n= 101 | ADVIA Chemistry 1800 GLUO y = 0.99x - 0.2 mmol/L r = 0.999 n= 101 | ≤ 10% change up to 200 mg/dL of hemoglobin. | ≤ 10% change up to 7.5 mg/dL of bilirubin | ≤ 10% change up to 250 mg/dL of Intralipid. | No Biotin in assay formulation | N/A | N/A | N/A | N/A | NIST SRM 965a | 11110184_EN Rev. 03, 2020-08 | IFU Document | 0.535 | 2.875 | 3.5 | 2.75 | 0.243 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097621 |  | ge31 | 00630414595634 | Global | General Chemistry | No | Atellica CH Glucose Oxidase (GluO) | GluOGluO (urine) | 0703 | 5600 | 4 x 1400 | N/A | N/A | 10 | 51 | glucose oxidase (microbial); ; peroxidase (horseradish) | ;#Urine;# | 4°C at collection | 4 μL | No |  | Glucose Oxidase Trinder | 6 -750 mg/dL | 0.3 - 41.6 mmol/L | 1,500 mg/dL | 83.3 mmol/L | Adults: < 0.5 g/day | Adults: < 0.5 g/day | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 days | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 35 | 7 |  | No |  | Yes | N/A | 11099300 |  | 6 - 1,500 mg/dL | 0.3 - 83.3 mmol/L | No | N/A | N/A | 1 mg/dL | 0.1 mmol/L | N/A | N/A | 1 mg/dL | 0 mg/dL | N/A | 0.0 mmol/L | 0.1 mmol/L | N/A | 28 mg/dL | 1.6 mmol/L | 1.6 | 4.2 | 299 mg/dL | 16.6 mmol/L | 0.4 | 2 | N/A | N/A |  |  | ADVIA Chemistry 1800 GLUO y = 1.00x + 4.0 mg/dL r = 0.998 n=105 | ADVIA Chemistry 1800 GLUO y = 1.00x + 0.2 mmol/L r = 0.998 n=105 | N/A | N/A | N/A | No Biotin in assay formulation | N/A | N/A | N/A | N/A | NIST SRM 965a | 11110184_EN Rev. 03, 2020-08 | IFU Document | 0.535 | 2.875 | 3.5 | 2.75 | 0.243 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097621 |  | ge31 | 00630414595634 | Global | General Chemistry | No | Atellica CH Glucose Oxidase (GluO) | GluOGluO (CSF) | 0703 | 5600 | 4 x 1400 | N/A | N/A | 10 | 51 | glucose oxidase (microbial); ; peroxidase (horseradish) | ;#Cerebrospinal fluid (CSF);# | CSF should be analyzed immediately, centrifuged and stored at 4°C or -20°C if delay is unavoidable | 4 μL | No |  | Glucose Oxidase Trinder | 6 -750 mg/dL | 0.3 - 41.6 mmol/L | 1,500 mg/dL | 83.3 mmol/L | Adult: 40-70 mg/dL Infants/children: 60-80 mg/dL | Adult .2-3.9 mmol/L Infants/children: 3.3-4.4 mmol/L | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 days | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 35 | 7 |  | No |  | Yes | N/A | 11099300 |  | 6 - 1,500 mg/dL | 0.3 - 83.3 mmol/L | No | N/A | N/A | 1 mg/dL | 0.1 mmol/L | N/A | N/A | 1 mg/dL | 0 mg/dL | N/A | 0.0 mmol/L | 0.1 mmol/L | N/A | 53 mg/dL | 2.9 mmol/L | 0.8 | 2.6 | 104 mg/dL | 5.8 mmol/L | 0.5 | 3.5 | N/A | N/A |  |  | ADVIA Chemistry 1800 GLUO y = 0.98x + 3.0 mg/dL r = 1.000 n=122 | ADVIA Chemistry 1800 GLUO y = 0.98x + 0.2 mmol/L r = 1.000 n=122 | N/A | N/A | N/A | No Biotin in assay formulation | N/A | N/A | N/A | N/A | NIST SRM 965a | 11110184_EN Rev. 03, 2020-08 | IFU Document | 0.535 | 2.875 | 3.5 | 2.75 | 0.243 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097633 |  | pl06 | 00630414596174 | Global | Specific Proteins | No | Atellica CH High Sensitivity C-Reactive Protein (hsCRP) | hsCRPhsCRP | 0334 | 740 | 2 x 370 | N/A | N/A | 8 | 30 | CRP antibody synthetic latex (rabbit) | ;#Serum;#Lithium Heparin Plasma;#Potassium EDTA Plasma;# | 3 days at 4-8 °C 6 months at -20 °C | 10 μL | No |  | latex enhanced immunoturbidimetric | 0.16–10.0 mg/L | 0.16–10.0 mg/L | 200 mg/L | 200 mg/L | Healthy individuals ≤ 3 mg/L Low risk for cardiovascular disease prediction <1 mg/L Average risk for cardiovascular disease prediction >1 and <3 mg/L High risk for cardiovascular disease prediction >3 mg/L | Healthy individuals ≤ 3 mg/L Low risk for cardiovascular disease prediction <1 mg/L Average risk for cardiovascular disease prediction >1 and <3 mg/L High risk for cardiovascular disease prediction >3 mg/L | N/A | N/A | N/A | Atellica CH hsCRP CAL | 11099412 | 60 days | Liquid | Liquid | 3 | 6 | Recombinant human c‑reactive protein (CRP) in a stabilized protein matrix | 60 | 30 |  | No |  | Yes | NA | 11099300 |  | 0.16–200 mg/L | 0.16–200 mg/L | No | NA | NA | 0.07 mg/L | 0.05 mg/L | NA | NA | 0.07 mg/L | 0.05 mg/L | 0.16 mg/L | 0.05 mg/L | 0.07 mg/L | 0.16 mg/L | 0.24 mg/L | NA | 6.8 | 7.4 | 0.73 mg/L | NA | 2.2 | 2.7 | 1.52 mg/L | NA | 1.3 | 1.7 | ADVIA Chemistry 1800 y = 0.96x + 0.02 mg/L r = 0.990 | ADVIA Chemistry 1800 y = 0.96x + 0.02 mg/L r = 0.990 | ≤ 10% change up to 500 mg/dL of hemoglobin. | ≤ 10% change up to 30 mg/dL of bilirubin | ≤ 10% change up to 1000 mg/dL of Intralipid. | No Biotin in assay formulation | 2000 mg/L | 2000 mg/L | 2,000 mg/L | 2,000 mg/L | IRMM -CRM-470 | 11110196_EN Rev. 02, 2019-07 | IFU Document | 0.431 | 2.875 | 3.5 | 2.75 | 0.196 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097627 |  | sp04 | 00630414595559 | Global | Specific Proteins | No | Atellica CH Lipoprotein(a) (Lp(a)) | Lp(a)Lp(a) | 1210 | 400 | 2 x 200 | N/A | N/A | 10 | 30 | Anti-LPA antibody coupled to latex particles (rabbit), lot-specific | ;#Serum;#Lithium Heparin Plasma;# | 14 days at 2-8 °C 6 months at ≤ -20 °C | 9 μL | No |  | latex enhanced immunoturbidimetric | 10.00–85.00 mg/dL | 0.100–0.850 g/L | N/A | N/A | < 30.00 mg/dL | < 0.300 g/L | N/A | N/A | N/A | Atellica CH Lp(a) CAL | 11099433 | 14 days | Lyophilized | Add 1.0 mL of distilled water; Replace stopper; Let the vials stand for 10 minutes at room temperature; to mix gently inverting the vials at least 5 times | 3 | 6 | Lyophilized human serum | 60 | 30 |  | No |  | No | NA | 11099300 |  | N/A | N/A | No | NA | NA | 2.45 mg/dL | 1.86 mg/dL | NA | NA | 2.45 mg/dL | 1.86 mg/dL | 9.30 mg/dL | 0.019 g/L | 0.025 g/L | 0.093 mg/dL | 19.27 mg/dL | 0.193 g/L | 2 | 2 | 33.60 mg/dL | 0.336 g/L | 0.7 | 1 | 47.45 mg/dL | 0.475 g/L | 0.4 | 1.5 | ADVIA Chemistry 1800 y = 0.97x + 0.90 mg/dL r = 0.997 n = 80 | ADVIA Chemistry 1800 y = 0.97x + 0.009 g/L r = 0.997 n = 80 | ≤ 10% change up to 1000 mg/dL of hemoglobin. | ≤ 10% change up to 60 mg/dL of bilirubin | ≤ 10% change up to 1000 mg/dL of Intralipid. | No Biotin in assay formulation | 490.00 mg/dL | 4.900 g/L | 490.00 mg/dL | 4.900 g/L | Internal standard | 11110190_EN Rev. 02, 2019-05 | IFU Document | 0.334 | 2.875 | 3.5 | 2.75 | 0.152 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097617 |  | pl30 | 00630414596259 | Global | Specific Proteins | No | Atellica CH Prealbumin (PreAlb) | PreAlbPreAlb | 1608 | 400 | 2 x 200 | N/A | N/A | 10 | 30 | Antihuman prealbumin (goat) | ;#Serum;# | 8 hours at room temperature 2 days at 2-8 °C or frozen at -20 °C or colder | 7.5 μL | No |  | PEG-enhanced immunoturbidimetric | 5.0 - 70.0 mg/dL | 50 - 700 mg/L | 140.0 mg/dL | 1,400 mg/L | 10.0-40.0 mg/dL | 100-400 mg/L | N/A | N/A | N/A | Atellica CH LSP CAL | 11099434 | 28 days | Liquid | Ready to use | 3 | 6 | Proteins derived from human serum | 60 | 7 |  | No |  | Yes | N/A | 11099300 |  | 5.0-140.0 mg/dL | 50-1,400 mg/L | No | N/A | N/A | 1.8 mg/dL | 18 mg/L | NA | NA | 1.8 mg/dL | 0.0 mg/dL | N/A | 0 mg/L | 18 mg/L | NA | 20.8 mg/dL | 208 mg/L | 1.4 | 2.6 | 40.3 mg/dL | 403 mg/L | 1.2 | 1.7 | 60.3 mg/dL | 603 mg/L | 0.9 | 1.7 | ADVIA Chemistry 1800 y = 1.05x + 0.0 mg/dL r = 0.993 n = 80 | ADVIA Chemistry 1800 y = 1.05x + 0.0 mg/L r = 0.993 n = 80 | ≤ 10% change up to 1000 mg/dL of hemoglobin. | ≤ 10% change up to 25 mg/dL of bilirubin | ≤ 10% change up to 250 mg/dL of Intralipid. | No Biotin in assay formulation | 420.0 mg/dL | 4200 mg/L | 420.0 mg/dL | 4200 mg/L | IRMM CRM-470 from IFCC | 11110180_EN Rev. 04, 2020-06 | IFU Document | 0.317 | 2.875 | 3.5 | 2.75 | 0.144 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097618 |  | pl32 | 00630414595603 | Global | Specific Proteins | No | Atellica CH Rheumatoid Factor (RF) | RFRF | 1802 | 360 | 2 x 180 | N/A | N/A | 10 | 21 | BSA (bovine); IgG synthetic latex (human) | ;#Serum;# | 7 days at 2-8°C 3 months at -20°C | 11.5 μL | No |  | Latex-enhanced immunoturbidimetric assay | 3.5 - 90.0 IU/mL | 3.5 - 90.0 IU/mL | 900 IU/mL | 900 IU/mL | < 14 IU/mL | < 14 IU/mL | N/A | N/A | N/A | Atellica CH LSP CAL | 11099434 | 28 days | Liquid | Ready to Use | 3 | 6 | Proteins derived from human serum | 60 | 14 |  | No |  | Yes | NA | 11099300 |  | 3.5 - 900 IU/mL | 3.5 - 900 IU/mL | No | NA | NA | 3.2 IU/mL | 3.2 IU/mL | NA | NA | 3.2 IU/mL | 1.6 IU/mL | NA | 1.6 IU/mL | 3.2 IU/mL | NA | 20.8 IU/mL | 20.8 IU/mL | 2.5 | 3.3 | 50.9 IU/mL | 50.9 IU/mL | 0.5 | 1 | 71.1 IU/mL | 71.1 IU/mL | 0.6 | 1 | ADVIA Chemistry 1800 RF y = 0.99x + 0.9 IU/mL r = 0.993 n = 80 | ADVIA Chemistry 1800 RF y = 0.99x + 0.9 IU/mL r = 0.993 n = 80 | ≤ 10% change up to 1000 mg/dL of hemoglobin. | ≤ 10% change up to 25 mg/dL of bilirubin | ≤ 10% change up to 1000 mg/dL of Intralipid. | No Biotin in assay formulation | 450 IU/mL | 450 IU/mL | 450 IU/mL | 450 IU/mL | Internal standard | 11110181_EN Rev. 02, 2019-11 | IFU Document | 0.174 | 2.875 | 3.5 | 2.75 | 0.079 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097532 |  | ge06 | 00630414595801 | Global | General Chemistry | No | Atellica CH Direct Bilirubin 2 (DBil_2) | DBil_2DBil_2 | 0205 | 1792 | 4 x 448 | N/A | N/A | 10 | 30 | Tartrate buffer; Phosphate buffer: sodium metavanadate | ;#Serum;#Lithium Heparin Plasma;# | 5 days at 4°C; Frozen at-70°C up to 3 months | 14.3 μL | No |  | Vanadate oxidation | 0.1–15.0 mg/dL | 2–256 μmol/L | 0.1 - 22.5 mg/dL | 2 - 385 μmol/L | ≤ 0.3 mg/dL | ≤ 5 μmol/L | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 8 hours | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 60 | 30 |  | No |  | Yes | N/A | 11099300 |  | 0.1 - 22.5 mg/dL | 2 - 385 μmol/L | No | N/A | N/A | 0.0 mg/dL | 0.0 μmol/L | N/A | N/A | 0.0 mg/dL | 0.0 mg/dL | N/A | 0.0 μmol/L | 0.0 μmol/L | N/A | 0.4 mg/dL | 7 μmol/L | 1 | 1.6 | 3.5 mg/dL | 60 μmol/L | 1 | 1.6 | 11.9 mg/dL | 203 μmol/L | 0.2 | 0.5 | ADVIA Chemistry 1800 D_BIL_2 y = 1.03x + 0.0 mg/dL r=1.00 n=103 | ADVIA Chemistry 1800 D_BIL_2 y = 1.03x + 0 μmol/L r=1.00 n=103 | ≤ 10% change up to 1000 mg/dL of hemoglobin. | N/A | ≤ 10% change up to 750 mg/dL of Lipemia (Triglycerides Concentrate) | No Biotin in assay formulation | N/A | N/A | N/A | N/A | AACC reference method NIST SRM 916 reference materials | 11110150_EN Rev. 03, 2020-08 | IFU Document | 0.939 | 5.625 | 3.5 | 2.75 | 0.427 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097629 |  | sp01 | 00630414596211 | Global | Specific Proteins | No | Atellica CH Alpha-1-acid Glycoprotein (AAG) | AAGAAG | 1109 | 400 | 2 x 200 | N/A | N/A | 10 | 30 | Antihuman alpha-1-acid glycoprotein antibody (goat) (polyclonal) | ;#Serum;#Lithium Heparin Plasma;#Potassium EDTA Plasma;# | Up to 3 days at 4°C ; Frozen at -20°C for up to 6 months | 3.2 μL | No |  | PEG-enhaced immunoturbidimetric assay | 5.0–612.0 mg/dL | 0.05–6.12 g/L | 5.0 - 1836 mg/dL | 0.05 - 18.36 g/L | 58.2–154.8 mg/dL | 0.58–1.55 g/L | N/A | N/A | N/A | Atellica CH Liquid Specific Protein Calibrators (LSP CAL) | 11099434 | 28 days | Liquid | No preparation required | 3 | 6 | Proteins derived from human serum; sodium azide | 60 | 14 |  | No |  | Yes | Atellica CH Diluent | 11099300 |  | 5.0 - 1836 mg/dL | 0.05 - 18.36 g/L | No | N/A | N/A | 2.0 mg/dL | 0.02 g/L | N/A | N/A | 2.0 mg/dL | 0.9 mg/dL | N/A | 0.01 g/L | 0.02 g/L | N/A | 43.2 mg/dL | 0.43 g/dL | 1.4 | 2.3 | 94.4 mg/dL | 0.94 g/dL | 0.8 | 2.4 | 354.6 mg/dL | 3.55 g/dL | 0.8 | 1 | ADVIA Chemistry 1800 y = 1.01x + 0.6 mg/dL r=0.996 n = 101 | ADVIA Chemistry 1800 y = 1.01x + 0.01 g/dL r=0.996 n=101 | ≤ 10% at 1000 mg/dL of hemoglobin | ≤ 10% at 60 mg/dL of bilirubin | ≤ 10% at 1000 mg/dL of Intralipid | No Biotin in assay formulation | 2165.0 mg/dL | 21.65 g/L | 2165.0 mg/dL | 21.65 g/L | Institute for Reference Materials and Measurements (IRMM) ERM‑DA 470 | 11110192_EN Rev. 03, 2019-05 | IFU Document | 0.345 | 2.875 | 3.5 | 2.75 | 0.157 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097600 |  | ge01 | 00630414596419 | Global | General Chemistry | No | Atellica CH Alkaline Phosphatase, Concentrated (ALP_2c) | ALP_2cALP_2c | 1108 | 4800 | 4 x 1200 | N/A | N/A | 8 | 30 | R1: 2-amino-2-methyl-1-propanol (AMP) buffer; HEDTA; magnesium acetate;; zinc sulfate; R2: Paranitrophenyl-phosphate (p‑NPP) substrate; 6 mmol; sodium azide; ProClin 300 | ;#Serum;#Lithium Heparin Plasma;# | Up to 8 hours at 25°C; up to 7 days at 2–8°C Frozen at -20°C or colder for up to 6 months. | 10 μL | No |  | IFCC Standardization | 10 - 1000 U/L | 10 - 1000 U/L | 10 - 2300 U/L | 10 - 2300 U/L | 46–116 U/L | 46–116 U/L | N/A | N/A | N/A | Atellica CH ALP_2 CAL | 11099316 | 30 days | Liquid | No preparation required | 3 | 2 | Human serum albumin; alkaline phosphatase (porcine kidney) | 60 | 17 |  | No |  | Yes | N/A | 11099300 |  | 10 - 2300 U/L | 10 - 2300 U/L | No | N/A | N/A | 0.4 U/L | 0.4 U/L | N/A | N/A | 0.4 U/L | 0 U/L | 6.2 U/L | 0 U/L | 0.4 U/L | 6.2 U/L | 87 U/L | 87 U/L | 0.4 | 1.3 | 277 U/L | 277 U/L | 0.2 | 0.7 | 841 U/L | 841 U/L | 0.2 | 1.5 | Dimension ALPI y = 1.05x -4U/L r=1.000 n=104 | Dimension ALPI y = 1.05x -4U/L r=1.000 n=104 | ≤ 10% at 1000 mg/dL of hemoglobin | ≤ 10% at 80 mg/dL of bilirubin | ≤ 10% at 600 mg/dL of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | traceable to IFCC primary reference procedure for the measurement of catalytic activity of alkaline phosphatase at 37°C | 11110163_EN Rev. 03, 2019-07 | IFU Document | 1.063 | 5.625 | 3.5 | 2.75 | 0.483 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097529 |  | ge18 | 00630414595832 | Global | General Chemistry | No | Atellica CH Ammonia (Amm) | AmmAmm | 0116 | 480 | 4 x 120 | N/A | N/A | 4 | 30 | α-Ketoglutarate (10 mmol/L); glutamate dehydrogenase (24 kU/L); NADPH analog (0.2 mmol/L); | ;#Lithium Heparin Plasma;#Potassium EDTA Plasma;# | Up to 2 hours at 2–8°C; Samples should be analyzed within 30 minutes of centrifugation. | 25 μL | No |  | Enzymatic | 17 – 1277 μg/dL | 10-750 μmol/L | 17 - 2554 μg/dL | 10 - 1499 μmol/L | 19–54 μg/dL | 11–32 μmol/L | N/A | N/A | N/A | Atellica CH CHEM III CAL | 11099335 | 30 days | Liquid | No preparation required | 3 | 3 | Ammonium bicarbonate | 30 | 3 |  | No |  | Yes | N/A | 11099300 |  | 17 - 2554 μg/dL | 10 - 1499 μmol/L | No | N/A | N/A | 11 μg/dL | 6 μmol/L | N/A | N/A | 11 μg/dL | 3 μg/dL | 17 μg/dL | 2 μmol/L | 6 μmol/L | 10 μmol/L | 48 μg/dL | 28 μmol/L | 6.2 | 7.3 | 197 μg/dL | 115 μmol/L | 2.1 | 2.6 | 1050 μg/dL | 616 μmol/L | 0.4 | 1 | Dimension RxL y = 1.06x - 10 μg/dL r=0.999 n=102 | Dimension RxL; y = 1.06x - 6 μmol/L r=0.999 n=102 | 18% bias at 375 mg/dL of hemoglobin | 18% bias at 60 mg/dLof bilirubin | ≤ 10% at 200 mg/dL of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | ACS Grade Ammonium Sulfate | 11110147_EN Rev. 04, 2020-03 | IFU Document | 0.185 | 2.875 | 3.5 | 2.75 | 0.084 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097615 |  | ge20 | 00630414596273 | Global | General Chemistry | No | Atellica CH Cholinesterase (CHE) | CHECHE | 0339 | 270 | 2 x 135 | N/A | N/A | 9 | 30 per pack | Phosphate buffer; 5,5’-dithiobis (2-Nitrobenzoic acid); Butyrylthiocholine iodide | ;#Serum;#Lithium Heparin Plasma;# | Up to 6 hours at 25°C; up to 7 days at 2–8°C; Frozen at -70°C or colder for up to 6 months. | 2 μL | No |  | Butyrylthiocholine substrate | 1,500 - 30,000 U/L | 1,500 - 30,000 U/L | 1500 - 60,000 U/L | 1500 - 60,000 U/L | 7000–19,000 U/L | 7000–19,000 U/L | N/A | N/A | N/A | Atellica CH ENZ 1 CAL | 11099317 | 30 days | Liquid | No preparation required | 3 | 2 | Bovine Serum; CHE (equine serum) | 120 | 7 |  | No |  | Yes | N/A | 11099300 |  | 1500 - 60,000 U/L | 1500 - 60,000 U/L | No | N/A | N/A | 100 U/L | 100 U/L | N/A | N/A | 100 U/L | 42 U/L | N/A | 42 U/L | 100 U/L | N/A | 4270 U/L | 4270 U/L | 0.3 | 2.9 | 12,499 U/L | 12,499 U/L | 0.3 | 1.4 | 27,050 U/L | 27,050 U/L | 0.3 | 1 | ADVIA Chemistry 1800 y = 1.00x - 56 U/L r=0.997 n=100 | ADVIA Chemistry 1800 y = 1.00x - 56 U/L r=0.997 n=100 | ≤ 10% at 500 mg/dL of hemoglobin | ≤ 10% at 25 mg/dL of bilirubin | ≤ 10% at 500 mg/dL of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | Traceable to the molar extinction coefficient of the reactionproduct | 11110178_EN Rev. 01, 2017-05 | IFU Document | 0.442 | 5.625 | 4.125 | 2.75 | 0.201 | 14.287 | 10.477 | 6.985
Atellica® CH | 11097521 |  | ge24 | 00630414595900 | Global | General Chemistry | No | Atellica CH Carbon Dioxide, Concentrated (CO2_c) | CO2_cCO2_c | 2015 | 7600 | 4 x 1900 | N/A | N/A | 2 | 27 | malate dehydrogenase (MDH) (mammalian); PEP: microorganism (PEPC); | ;#Serum;#Lithium Heparin Plasma;# | Up to 3 days at 2–8°C; Frozen at-20°C for up to 60 days | 5.7 μL | No |  | Enzymatic | 10.0 - 40.0 mEq/L | 10.0 - 40.0 mmol/L | N/A | N/A | 20.0–31.0 mEq/L | 20.0–31.0 mmol/L | N/A | N/A | N/A | Atellica CH CO2 CAL | 11099401 | 30 days | Liquid | No preparation required | 3 | 2 | Not provided | 30 | 6 |  | No |  | No | N/A | 11099300 |  | N/A | N/A | No | N/A | N/A | 0.9 mEq/L | 0.9 mmol/L | N/A | N/A | 0.9 mEq/L | ≤ 0.4 mEq/L | N/A | ≤ 0.4mmol/L | 0.9 mmol/L | N/A | 17.3 mEq/L | 17.3 mmol/L | 1.11 | 3.45 | 24.4 mEq/L | 24.4 mmol/L | 1.73 | 4.13 | 31.8 mEq/L | 31.8 mmol/L | 1.24 | 3.57 | ADVIA Chemistry 1800 CO2_L y = 0.97x - 0.2 mEq/L r=0.990 n=59 | ADVIA Chemistry 1800 CO2_L y = 0.97x- 0.2 mmol/L r=0.990 n=59 | ≤ 10% at 600 mg/dL of hemoglobin | ≤ 10% at 30 mg/dL of bilirubin | ≤ 10% at 750 mg/dL of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | Corning 965 method / NIST SRM 192 | 11110139_EN Rev. 02, 2019-04 | IFU Document | 0.554 | 2.875 | 3.5 | 2.75 | 0.252 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097632 |  | ge16 | 00630414596181 | Global | General Chemistry | No | Atellica CH LDL Cholesterol Direct (DLDL) | DLDLDLDL | 0335 | 1600 | 4 x 400 | N/A | N/A | 10 | 15 | ; N‑ethyl‑N‑(2‑hydroxy‑3‑sulfopropyl)‑3‑methylaniline (TOOS); cholesterol esterase (microbial); cholesterol oxidase (microbial): catalase (microbial); ; 4-aminoantipyrine; peroxidase (horseradish) | ;#Serum;#Lithium Heparin Plasma;# | Up to 3 days at 2–8°C; Frozen at-20°C for up to several weeks; Frozen at -70°C or lower for longer time | 5 μL | No |  | Elimination/catalase | 5.0 – 1000.0 mg/dL | 0.13 – 25.90 mmol/L | N/A | N/A | Adults Optimal; <100.0 mg/dL Near optimal: 100.0–129.0 mg/dL Borderline high: 130.0–159.0 mg/dL High: 160.0–189.0 mg/dL Very high: ≥ 190.0 mg/dL | Adults Optimal: <2.60 mmol/L Near optimall: 2.60–3.30 mmol/L Borderline high: 3.40–4.10 mmol/L High: 4.10–4.90 mmol/L Very high: ≥ 4.90 mmol/L | N/A | N/A | N/A | Atellica CH HDL/LDL CAL | 11099402 | 3 days | Lyophilized | 1.0 mL of distilled water; Let stand 30 minutes | 3 | 2 | Human serum | 121 | 15 |  | No |  | No | N/A | 11099300 |  | N/A | N/A | No | N/A | N/A | 3.0 mg/dL | 0.08 mmol/L | N/A | N/A | 3.0 mg/dL | 0.8 mg/dL | N/A | 0.02 mmol/L | 0.08 mmol/L | N/A | 104.5 mg/dL | 2.71 mmol/L | 1.3 | 2.3 | 183.0 mg/dL | 4.74 mmol/L | 0.7 | 1.6 | 953.8 mg/dL | 24.7 mmol/L | 0.4 | 1 | ADVIA Chemistry 1800 DLDL y = 0.99x - 0.8 mg/dL r=1.000 n=143 | ADVIA Chemistry 1800 DLDL y = 0.99x - 0.02 mmol/L r=1.000 n=143 | ≤ 10% at 810 mg/dL of hemoglobin | ≤ 10% at 20 mg/dL of conjugated bilirubin ≤ 10% at 30 mg/dL of unconjugated bilirubin | ≤ 10% at 505 mg/dL of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | CDC Reference Method via patient sample correlation | 11110195_EN Rev. 03, 2019-11 | IFU Document | 0.906 | 5.625 | 3.5 | 2.75 | 0.412 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097533 |  | ge27 | 00630414595795 | Global | General Chemistry | No | Atellica CH Enzymatic Creatinine_2 (ECre_2) | ECre_2 (serum/plasma)ECre_2 (serum/plasma) | 0340 | 1400 | 4 x 350 | N/A | N/A | 10 | 30 | Creatinase; sarcosine oxidase; N-(3 sulfopropyl)-3 methoxy-5 methylaniline; ; 4-aminoantipyrine | ;#Serum;#Lithium Heparin Plasma;#Potassium EDTA Plasma;#Urine;# | Up to 2 days at 2–8°C; Frozen at -20° or below | 13 μL | No |  | Enzymatic | 0.1 - 30.00 mg/dL | 9 - 2652 µmol/L | 0.10 - 150 mg/dL | 9 - 13,260 μmol/L | Adult males: Serum/Plasma 0.6–1.1 mg/dL Adult females: Serum/Plasma 0.5–0.8 mg/dL | Adult males Serum/Plasma: 53–97 μmol/L Adult females Serum/Plasma: 44–71 μmol/L | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 180 | 30 |  | No |  | Yes | N/A | 11099300 |  | 0.10 - 150 mg/dL | 9 - 13,260 μmol/L | No | N/A | N/A | 0.03 mg/dL | 3 μmol/L | N/A | N/A | 0.03 mg/dL | 0.02 mg/dL | N/A | 2 μmol/L | 3 μmol/L | N/A | 0.70 mg/dL | 62 µmol/L | 2.5 | 2.7 | 6.90 mg/dL | 610 µmol/L | 0.5 | 2.1 | 17.74 mg/dL | 1568 µmol/L | 0.3 | 2.2 | ADVIA Chemistry 1800 y = 1.00x + 0.04 mg/dL r=1.000 n=103 | ADVIA Chemistry 1800 y = 1.00x + 4 µmol/L r=1.000 n=103 | < 10% at 750 mg/dL of hemoglobin | < 10% at 30 mg/dL of bilirubin | < 10% at 1000 mg/dL of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | NIST SRM 967 | 11110151_EN Rev. 05, 2020-03 | IFU Document | 0.856 | 5.625 | 3.5 | 2.75 | 0.389 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097533 |  | ge27 | 00630414595795 | Global | General Chemistry | No | Atellica CH Enzymatic Creatinine_2 (ECre_2) | ECre_2 (urine)ECre_2 (urine) | 0340 | 1400 | 4 x 350 | N/A | N/A | 10 | 30 | Creatinase ; sarcosine oxidase; N-(3 sulfopropyl)-3 methoxy-5 methylaniline; 4-aminoantipyrinE; peroxidase | ;#Urine;# | Up to 4 days at 2–8°C; frozen at -20°C or below | 13 μL | No |  | Enzymatic | 1.0 - 245.00 mg/dL | Urine 88 - 21658 µmol/L | 1.00 - 1225 mg/dL | 88 - 108,290 μmol/L | Adult males Urine 800–2000 mg/day Adult females Urine 600–1800 mg/day | Adult males Urine 800–2000 mg/day Adult females Urine 600–1800 mg/day | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 180 | 30 |  | No |  | Yes | N/A | 11099300 |  | 1.00 - 1225 mg/dL | 88 - 108,290 μmol/L | No | N/A | N/A | 0.22 mg/dL | 19 μmol/L | NA | NA | 0.22 mg/dL | 0.02 mg/dL | N/A | 2 μmol/L | 19 μmol/L | N/A | 57.02 mg/dL | 5041 µmol/L | 0.6 | 2.5 | 123.18 mg/dL | 10,889 µmol/L | 0.3 | 2.3 | 199.78 mg/dL | 17,661 µmol/L | 0.7 | 3.2 | ADVIA Chemistry 1800 y = 0.97x + 0.00 mg/dL r=1.000 n=109 | ADVIA Chemistry 1800 y = 0.97x + 0 µmol/L r=1.000 n=109 | N/A | N/A | N/A | No Biotin in assay formulation | N/A | N/A | N/A | N/A | NIST SRM 967 | 11110151_EN Rev. 05, 2020-03 | IFU Document | 0.856 | 5.625 | 3.5 | 2.75 | 0.389 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097597 |  | ge28 | 00630414596440 | Global | General Chemistry | No | Atellica CH Gamma-Glutamyl Transferase (GGT) | GGTGGT | 0701 | 1792 | 4 x 448 | N/A | N/A | 10 | 22 | synthetic | ;#Serum;#Lithium Heparin Plasma;# | Up to 7 days at 25°C or at 2–8°C; Frozen at -20°C for up to 6 months | 10 μL | No |  | Modified IFCC | 7 - 1200 U/L | 7 - 1200 U/L | 7 - 3600 U/L | 7 - 3600 U/L | Adult males Serum/plasma: < 73 U/L Adult females Serum/plasma; < 38 U/L | Adult males Serum/plasma; 73 U/L Adult females Serum/plasma; 38 U/L | N/A | N/A | N/A | Atellica CH ENZ 1 CAL | 11099317 | 30 days | Liquid | No preparation required | 3 | 2 | Bovine serum; GGT (bovine kidney) | 60 | 22 |  | No |  | Yes | N/A | 11099300 |  | 7 - 3600 U/L | 7 - 3600 U/L | No | N/A | N/A | 2 U/L | 2 U/L | N/A | N/A | 2 U/L | 1 U/L. | N/A | 1 U/L. | 2 U/L | N/A | 42 U/L | 42 U/L | 1.9 | 2.6 | 81 U/L | 81 U/L | 1 | 1.5 | 1044 U/L | 1044 U/L | 0.2 | 0.9 | ADVIA Chemistry 1800 GGT y = 1.01x - 3 U/L r=0.999 n=100 | ADVIA Chemistry 1800 GGT y = 1.01x - 3 U/L r=0.999 n=100 | ≤ 10% at 1000 mg/dL of hemoglobin | ≤ 10% at 10 mg/dL of bilirubin | ≤ 10% at 500 mg/dL of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | IFCC-452 reference material | 11110160_EN Rev. 02, 2019-07 | IFU Document | 1.217 | 5.625 | 3.5 | 2.75 | 0.553 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097611 |  | ge32 | 00630414596303 | Global | General Chemistry | No | Atellica CH Inorganic Phosphorus (IP) | IPIP (serum/plasma) | 1605 | 5100 | 3 x 1700 | N/A | N/A | 10 | 30 | Ammonium molybdate; Sulfuric acid | ;#Serum;#Lithium Heparin Plasma;#Urine;# | Up to several days at 2-8 °C; Frozen at -70 °C for up to several months | 8 μL | No |  | Phosphomolybdate | 0.3 -20 mg/dL | 0.10 - 6.46 mmol/L | 0.3 - 40.0 mg/dL | 0.1 - 12.92 mmol/L | Adult 2.4-5.1 mg/dL | Adult 0.78-1.65 mmol/L | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 days | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 180 | 7 |  | No |  | Yes | N/A | 11099300 |  | 0.3 - 40.0 mg/dL | 0.1 - 12.92 mmol/L | No | N/A | N/A | 0.1 mg/dL | 0.03 mmol/L | N/A | N/A | 0.1 mg/dL | 0.00 mg/dL | N/A | 0.00 mmol/L | 0.03 mmol/L | N/A | 3.0 mg/dL | 0.97 mmol/L | 2 | 2.7 | 4.6 mg/dL | 1.49 mmol/L | 1.1 | 4.6 | 11.8 mg/dL | 3.81 mmol/L | 0.7 | 0.9 | ADVIA Chemistry 1800 IP y = 1.03x - 0.2 mg/dL r=1.000 n= 106 | ADVIA Chemistry 1800 IP y = 1.03x - 0.06 mmol/L r=1.000 n= 106 | ≤ 10% at 500 mg/dL of hemoglobin | ≤ 10% at 30 mg/dL of bilirubin | ≤ 10% at 650 mg/dL of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | NIST SRM 200 | 11110174_EN Rev. 04, 2020-07 | IFU Document | 0.785 | 4.187 | 3.5 | 2.75 | 0.357 | 10.636 | 8.89 | 6.985
Atellica® CH | 11097611 |  | ge32 | 00630414596303 | Global | General Chemistry | No | Atellica CH Inorganic Phosphorus (IP) | IPIP (urine) | 1605 | 5100 | 3 x 1700 | N/A | N/A | 10 | 30 | Ammonium molybdate; Sulfuric acid | ;#Serum;#Lithium Heparin Plasma;#Urine;# | Up to 6 months at 2-8 °C for refrigerated acidifed urine specimens | 8 μL | No |  | Phosphomolybdate | 4 -100 mg/dL | 1.29 - 32.30 mmol/L | 4.0 - 260 mg/dL | 1.29 - 83.98 mmol/L | 0.4-1.3 g/day for adults on an unrestricted diet | 12.9-42.0 mmol/day for adults on an unrestricted diet | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 days | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 180 | 7 |  | No |  | Yes | N/A | 11099300 |  | 4.0 - 260 mg/dL | 1.29 - 83.98 mmol/L | No | N/A | N/A | 1.4 mg/dL | 0.45 mmol/L | N/A | N/A | 1.4 mg/dL | 0.5 mg/dL | N/A | 0.16 mmol/L | 0.45 mmol/L | N/A | 17.3 mg/dL | 5.59 mmol/L | 1.5 | 4 | 51.8 mg/dL | 16.73 mmol/L | 1.1 | 2.1 | 84.1 mg/dL | 27.16 mmol/L | 0.5 | 2.3 | ADVIA Chemistry 1800 IP y = 1.01x + 0.7 mg/dL r=0.999 n=104 | ADVIA Chemistry 1800 IP y = 1.01x + 0.23 mmol/L r=0.999 n=104 | N/A | N/A | N/A | No Biotin in assay formulation | N/A | N/A | N/A | N/A | NIST SRM 200 | 11110174_EN Rev. 04, 2020-07 | IFU Document | 0.785 | 4.187 | 3.5 | 2.75 | 0.357 | 10.636 | 8.89 | 6.985
Atellica® CH | 11097594 |  | ge08 | 00630414596464 | Global | General Chemistry | No | Atellica CH Lactate Dehydrogenase L-P (LDLP) | LDLPLDLP | 1202 | 1792 | 4 x 448 | N/A | N/A | 9 | 28 | L-lactic acid; NAD; 2‑chloracetamide | ;#Serum;#Lithium Heparin Plasma;# | 7 days at 20–25°C; 4 days at 4–8°C; 42 days at -20°C | 14 μL | No |  | Lactate / NAD | 14-750 U/L | 14-750 U/L | 14 - 4500 U/L | 14 - 4500 U/L | 120 - 246 U/L | 120 - 246 U/L | N/A | N/A | N/A | Atellica CH ENZ 1 CAL | 11099317 | 30 days | Liquid | No preparation required | 3 | 2 | Bovine serum albumin; LDHL (chicken heart) | 60 | 28 |  | No |  | Yes | N/A | 11099300 |  | 14 - 4500 U/L | 14 - 4500 U/L | No | N/A | N/A | 3 U/L | 3 U/L | N/A | N/A | 3 U/L | 2 U/L | 11 U/L | 2 U/L | 3 U/L | 11 U/L | 189 U/L | 189 U/L | 0.9 | 1 | 403 U/L | 403 U/L | 0.7 | 1 | 557 U/L | 557 U/L | 0.3 | 0.7 | ADVIA Chemistry 1800 y = 0.96x + 6 U/L r=0.999 n=100 | ADVIA Chemistry 1800 y = 0.96x + 6 U/L r=0.999 n=100 | N/A | ≤ 10% at 30 mg/dL of bilirubin | ≤ 10% at 650 mg/dL of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | IFCC-453 reference material | 11110157_EN Rev. 04, 2019-11 | IFU Document | 0.988 | 5.625 | 3.5 | 2.75 | 0.449 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097524 |  | ge44 | 00630414595870 | Global | General Chemistry | No | Atellica CH Total Protein_2 (Urine) (UPro) | UProUpro | 2038 | 1600 | 4 x 400 | N/A | N/A | 10 | 30 | Pyrogallol red; sodium molybdate; | ;#Urine;#Cerebrospinal fluid (CSF);# | Up to 3 days at 4 °C; frozen at -20 °C for up to 6 months | 13.3 μL | No |  | Dye binding/ Pyrogallol red | 1.0 - 125.0 mg/dL | 10 - 1250 mg/L | 1.0 - 500.0 mg/dL | 10 - 5000 mg/L | Urine: 24-hour: 1-14 mg/dL; At rest: 50-80 mg/day; After intense exercise 250 mg/day CSF: 8-32 mg/dL | Urine: 24-hour: 10-140 mg/L; At rest: 50-80 mg/day; After intense exercise 250 mg/day CSF: 80-320 mg/L | N/A | N/A | N/A | Atellica CH UPro CAL | 11099441 | 21 days | Liquid | No preparation required | 3 | 2 | Human serum albumin | 60 | 30 |  | No |  | Yes | N/A | 11099300 |  | 1.0 - 500.0 mg/dL | 10 - 5000 mg/L | No | N/A | N/A | 0.9 mg/dL | 9 mg/L | N/A | N/A | 0.9 mg/dL | 0.3 mg/dL | N/A | 3 mg/L | 9 mg/L | N/A | 12.1 mg/dL | 121 mg/L | 4.9 | 5.3 | 54.9 mg/dL | 549 mg/L | 1.2 | 1.9 | 113.1 mg/dL | 1131 mg/L | 0.6 | 0.9 | ADVIA Chemistry 1800 Urine; y = 0.99x + 0.5 mg/dL r=0.994 n=100 CSF: y=1.00x + 0.2mg/dL r=0.999 n=101 | ADVIA Chemistry 1800 Urine; y = 0.99x + 5 mg/L r=0.994 n=100 CSF: y=1.00x + 2mg/L r=0.999 n=101 | 38% bias at 16 mg/dL of hemoglobin | N/A | N/A | No Biotin in assay formulation | N/A | N/A | N/A | N/A | Traceable to an internal standard | 11110142_EN Rev. 02, 2019-07 | IFU Document | 0.585 | 2.875 | 3.5 | 2.75 | 0.266 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097630 |  | ge12 | 00630414596204 | Global | General Chemistry | No | Atellica CH Direct HDL Cholesterol (D-HDL) | D-HDLD-HDL | 0313 | 1792 | 4 x 448 | N/A | N/A | 8 | 89 | N (2-hydroxy-3-sulfopropyl)-3,5-dimethoxyaniline;cholesterol esterase (microbial); cholesterol oxidase (microbial): catalase (microbial); ; 4-aminoantipyrine; peroxidase (plant) | ;#Serum;#Lithium Heparin Plasma;# | 7 days at 2–8°C; Frozen at -70°C or colder for up to 3 months | 5 μL | No |  | Elimination/catalase | 20.0 – 129.0 mg/dL | 0.52 - 3.34 mmol/L | 258.0 mg/dL | 6.68 mmol/L for | Adults Serum/plasma: Low (undesirable, high risk): 40.0 mg/dL High (desirable, low risk): ≥ 60.0 mg/dL | Adults Serum/plasma: Low (undesirable, high risk): 1.00 mmol/L High (desirable, low risk): ≥ 1.60 mmol/L | N/A | N/A | N/A | Atellica CH HDL/LDL CAL | 11099402 | 3 days | Lyophilized | 1.0 mL of distilled water; Let stand 30 minutes | 3 | 2 | Human serum | 60 | 30 |  | No |  | Yes | N/A | 11099300 |  | 20.0 - 258.0 mg/dL | 0.52 - 6.68 mmol/L for | No | N/A | N/A | 2.1 mg/dL | 0.05 mmol/L | N/A | N/A | 2.1 mg/dL | 0.5 mg/dL | 2.7 mg/dL | 0.01 mmol/L | 0.05 mmol/L | 0.07 mmol/L | 37.5 mg/dL | 0.97 mmol/L | 1.8 | 2 | 64.6 mg/dL | 1.67 mmol/L | 0.6 | 1.6 | 101.5 mg/dL | 2.63 mmol/L | 0.7 | 0.7 | ADVIA Chemistry 1800 D-HDL y = 0.97x + 1.4 mg/dL r=0.999 n= 135 | ADVIA Chemistry 1800 D-HDL y = 0.97x + 0.04 mmol/L r=0.999 n= 135 | ≤ 10% at 500 mg/dL of hemoglobin | ≤ 10% at 30 mg/dL of bilirubin | ≤ 10% at 1000 mg/dL of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | NCEP Designated Comparison Method via patient sample correlation. | 11110193_EN Rev. 04, 2020-08 | IFU Document | 0.964 | 5.625 | 3.5 | 2.75 | 0.438 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097522 |  | dr22 | 00630414595894 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Acetaminophen (Acet) | Acet | 0102 | 1200 | 4 x 300 |  |  | 10 | 14 |  | ;#Serum;#Lithium Heparin Plasma;# | 8 hours at 20–25°C; frozen at -20°C for up to 45 days | 6 μL |  |  | Colorimetric: acyl amidohydrolase | 0.2 - 20.0 mg/dL | 13.2 - 1322.0 μmol/L | 0.2 - 60.0 mg/dL | 13.2 - 3966.0 μmol/L | 1.0–2.0 mg/dL | 66.0 - 132.0 µmol/L |  |  |  | Atellica CH TOX CAL | 11099440 | 3 days | Liquid |  | 3 | 2 |  | 62 | 1 |  | No |  | Yes | [5 empty] | No | [6 empty] | ≤0.2 mg/dL | 0.1 mg/dL | n/a | 6.6 μmol/L | ≤13.2 μmol/L | n/a | 4.6 mg/dL | 304.1 µmol/L | 1.2 | 3.7 | 9.0 mg/dL | 594.9 µmol/L | 0.6 | 1.5 | 16.7 mg/dL | 1103.9 µmol/L | 1.3 | 1.9 | ADVIA Chemistry 1800 y = 0.97x - 0.1 mg/dL (y = 0.97x - 6.6 μmol/L) r = 0.998 |  | ≤ 6% change up to 500 mg/dL of hemoglobin | ≤ 12% change up to 10 mg/dL of bilirubin | ≤ -14% change up to 500 mg/dL of Intralipid | No Biotin in assay formulation |  |  | n/a | n/a | Traceable to an internal standard manufactured using highly purified material. | 11110140_EN Rev. 04, 2019-04 | IFU Document | 0.442 | 5.625 | 3.5 | 2.75 | 0.201 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097506 |  | dr01 | 00630414596068 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Amphetamines (Amp) | Amp300 | 0140 | 1520 | 4 x 380 |  |  | 7 | 30 |  | ;#Urine;# | 7 days at 25°C; 30 days at 2–8°C; 12 months at -20°C | 13.5 μL |  |  | EMIT | 300 ng/mL;Semiquantitative 100 - 1000 ng/mL | n/a | n/a | n/a | qualitative/semi-quantitative assay | n/a |  |  |  | Emit Calibrator 0, 1, 2, 3 | 10445406 10445407 10445408 10445409 | 5 weeks | Liquid |  | 3 | 5 |  | 60 | 20 |  | No |  | No | [12 empty] | 53 ng/mL | 41 ng/mL | n/a | n/a | n/a | n/a | 307.8 mAU/min | N/A | 0.2 | 0.4 | 326.2 mAU/min |  | 0.2 | 0.5 | 346.0 mAU/min |  | 0.2 | 0.5 | Viva-E Analyzer ± 50% of 300 ng/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | Traceable to Syva Emit Cal/Cont which are gravimetrically prepared standards- qualified by GC/MS within ± 10% of nominal. | 11110124_EN Rev. 03, 2019-07 | IFU Document | 0.915 | 5.625 | 3.5 | 2.75 | 0.416 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097506 |  | dr01 | 00630414596068 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Amphetamines (Amp) | Amp500 | 0137 | 1520 | 4 x 380 |  |  | 7 | 30 |  | ;#Urine;# | 7 days at 25°C; 30 days at 2–8°C; 12 months at -20°C | 8 μL |  |  | EMIT | 500 ng/mL;Semiquantitative 150 - 1800 ng/mL | n/a | n/a | n/a | qualitative/semi-quantitative assay | n/a |  |  |  | Emit Calibrator 0, 1, 2, 3, 5 | 10445406 10445407 10445408 10445409 10445411 | 5 weeks | Liquid |  | 3 | 5 |  | 60 | 20 |  | No |  | No | [12 empty] | 63 ng/mL | 44 ng/mL | n/a | n/a | n/a | n/a | 323.7 mAU/min | N/A | 0.2 | 0.4 | 345.3 mAU/min |  | 0.2 | 0.4 | 365,9 |  | 0.3 | 0.5 | Viva-E Analyzer ± 50% of 500 ng/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | Traceable to Syva Emit Cal/Cont which are gravimetrically prepared standards- qualified by GC/MS within ± 10% of nominal. | 11110124_EN Rev. 03, 2019-07 | IFU Document | 0.915 | 5.625 | 3.5 | 2.75 | 0.416 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097506 |  | dr01 | 00630414596068 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Amphetamines (Amp) | Amp1K | 0138 | 1520 | 4 x 380 |  |  | 7 | 30 |  | ;#Urine;# | 7 days at 25°C; 30 days at 2–8°C; 12 months at -20°C | 5 μL |  |  | EMIT | 1000 ng/mL;Semiquantitative 150 - 1800 ng/mL | n/a | n/a | n/a | qualitative/semi-quantitative assay | n/a |  |  |  | Emit Calibrator 0, 1, 2, 3, 5 | 10445406 10445407 10445408 10445409 10445411 | 5 weeks | Liquid |  | 3 | 5 |  | 60 | 20 |  | No |  | No | [12 empty] | 121 ng/mL | 83 ng/mL | n/a | n/a | n/a | n/a | 345.1 mAU/min | N/A | 0.2 | 0.6 | 373.8 mAU/min |  | 0.3 | 0.5 | 401.1 mAU/min |  | 0.2 | 0.5 | Viva-E Analyzer ± 50% of 1000 ng/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | Traceable to Syva Emit Cal/Cont which are gravimetrically prepared standards- qualified by GC/MS within ± 10% of nominal. | 11110124_EN Rev. 03, 2019-07 | IFU Document | 0.915 | 5.625 | 3.5 | 2.75 | 0.416 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097505 |  | dr03 | 00630414596075 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Benzodiazepines (Bnz) | Bnz200 | 0202 | 1520 | 4 x 380 |  |  | 7 | 30 |  | ;#Urine;# | 7 days at 25°C; 2–8°C or frozen at -20°C for up to 30 days | 17.5 μL |  |  | EMIT | 200 ng/mL;Semi-quant 50-900 ng/mL | n/a | n/a | n/a | qualitative/semi-quantitative assay | n/a |  |  |  | Emit Calibrator/Control Levels 0, 2, 3, 4, 5 | 10445406 10445408 10445409 10445410 10445411 | 5 weeks | Liquid |  | 3 | 5 |  | 60 | 20 |  | No |  | No | [12 empty] | 27 ng/mL | 15 ng/mL | n/a | n/a | n/a | n/a | 187.3 mAU/min | N/A | 0.3 | 0.7 | 236.7 mAU/min |  | 0.6 | 0.8 | 256.4 mAU/min |  | 0.3 | 0.6 | Viva-E Analyzer ± 50% of 200 ng/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | Traceable to Syva Emit Cal/Cont which are gravimetrically prepared standards- qualified by GC/MS within ± 10% of nominal. | 11110123_EN Rev. 04, 2019-11 | IFU Document | 0.913 | 5.625 | 3.5 | 2.75 | 0.415 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097505 |  | dr03 | 00630414596075 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Benzodiazepines (Bnz) | Bnz300 | 0202 | 1520 | 4 x 380 |  |  | 7 | 30 |  | ;#Urine;# | 7 days at 25°C; 2–8°C or frozen at -20°C for up to 30 days | 17.5 μL |  |  | EMIT | 300 ng/mLSemi-quant 50-900 ng/mL | n/a | n/a | n/a | qualitative/semi-quantitative assay | n/a |  |  |  | Emit Calibrator/Control Levels 0, 2, 3, 4, 5 | 10445406 10445408 10445409 10445410 10445411 | 5 weeks | Liquid |  | 3 | 5 |  | 60 | 20 |  | No |  | No | [12 empty] | 27 ng/mL | 15 ng/mL | n/a | n/a | n/a | n/a | 187.3 mAU/min | N/A | 0.3 | 0.7 | 249.8 mAU/min |  | 0.2 | 0.6 | 256.4 mAU/min |  | 0.3 | 0.6 | Viva-E Analyzer ± 50% of 300 ng/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | Traceable to Syva Emit Cal/Cont which are gravimetrically prepared standards- qualified by GC/MS within ± 10% of nominal. | 11110123_EN Rev. 04, 2019-11 | IFU Document | 0.913 | 5.625 | 3.5 | 2.75 | 0.415 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097507 |  | dr02 | 00630414596051 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Barbiturates (Brb) | Brb200 | 0203 | 1520 | 4 x 380 |  |  | 7 | 30 |  | ;#Urine;# | 7 days at 25°C; 2–8°C or frozen at -20°C for up to 30 days | 17.5 μL |  |  | EMIT | 200 ng/mL;Semi-quant 40 – 750 ng/mL | n/a | n/a | n/a | qualitative/semi-quantitative assay | n/a |  |  |  | Emit Calibrator/Control Levels 0, 2, 3, 4, 5 | 10445406 10445408 10445409 10445410 10445411 | 5 weeks | Liquid |  | 3 | 5 |  | 60 | 30 |  | No |  | No | [12 empty] | 9 ng/mL | 5 ng/mL | n/a | n/a | n/a | n/a | 123.2 mAU/min | N/A | 0.3 | 0.6 | 171.9 mAU/min |  | 0.3 | 0.5 | 214.6 mAU/min |  | 0.3 | 0.5 | Viva-E Analyzer ± 50% of 200 ng/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | Traceable to Syva Emit Cal/Cont which are gravimetrically prepared standards- qualified by GC/MS within ± 10% of nominal. | 11110125_EN Rev. 03, 2019-07 | IFU Document | 0.917 | 5.625 | 3.5 | 2.75 | 0.417 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097507 |  | dr02 | 00630414596051 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Barbiturates (Brb) | Brb300 | 0203 | 1520 | 4 x 380 |  |  | 7 | 30 |  | ;#Urine;# | 7 days at 25°C; 2–8°C or frozen at -20°C for up to 30 days | 17.5 μL |  |  | EMIT | 300 ng/mL;Semi-quant 40 – 750 ng/mL | n/a | n/a | n/a | qualitative/semi-quantitative assay | n/a |  |  |  | Emit Calibrator/Control Levels 0, 2, 3, 4, 5 | 10445406 10445408 10445409 10445410 10445411 | 5 weeks | Liquid |  | 3 | 5 |  | 60 | 30 |  | No |  | No | [12 empty] | 9 ng/mL | 5 ng/mL | n/a | n/a | n/a | n/a | 123.2 mAU/min | N/A | 0.3 | 0.6 | 197.9 mAU/min |  | 0.4 | 0.6 | 214.6 mAU/min |  | 0.3 | 0.5 | Viva-E Analyzer ± 50% of 300 ng/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | Traceable to Syva Emit Cal/Cont which are gravimetrically prepared standards- qualified by GC/MS within ± 10% of nominal. | 11110125_EN Rev. 03, 2019-07 | IFU Document | 0.917 | 5.625 | 3.5 | 2.75 | 0.417 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097515 |  | th02 | 00630414595986 | Global | Therapeutic Drug Monitoring (TDM) | No | Atellica CH Carbamazepine (Carb) | Carb | 0309 | 400 | 4 x 100 | Atellica CH DRUG CAL II |  | 8 | 30 |  | ;#Serum;#Lithium Heparin Plasma;# | 8 hours at 25°C; 2 days at 2-8°C; frozen at -20°C for up to 30 days | 3.8 μL | No |  | PETINIA | 0.4 - 20.0 µg/mL | 1.7 - 84.7 µmol/L | 0.4 - 40.0 μg/mL | 1.7 - 169.3 μmol/L | 4.0–12.0 μg/mL | 16.9–50.8 μmol/L |  |  |  | Atellica CH DRUG CAL II | 11099405 | 30 days | Liquid |  | 3 | 5 |  | 30 | 7 |  | No |  | Yes | [5 empty] | Not Specified | [6 empty] | 0.4 μg/mL | 0.3 μg/mL | n/a | 1.3 μmol/L | 1.7 μmol/L | n/a | 3.2 μg/mL | 13.4 μmol/L | 1.5 | 2.4 | 6.4 μg/mL | 27.2 μmol/L | 1 | 1.3 | 16.5 μg/mL | 69.8 μmol/L | 0.8 | 1.2 | Dimension y = 0.98x + 0.01 μg/mL (y = 0.98x + 0.04 μmol/L) r = 0.998 |  | ≤ 2% change up to 500 mg/dL of hemoglobin | 0% change up to 20 mg/dL of bilirubin | ≤ 8% change up to 800 mg/dL of Intralipid | No Biotin in assay formulation |  |  | N/A | N/A | USP Carbamazepine Standard | 11110133_EN Rev. 01, 2017-05 | IFU Document | 0.629 | 5.625 | 3.5 | 2.75 | 0.286 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097504 |  | dr04 | 00630414596082 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Cocaine Metabolite (Coc) | Coc150 | 0336 | 1520 | 4 x 380 |  |  | 7 | 30 |  | ;#Urine;# | 7 days at 25°C; 30 days at 2–8°C or frozen at -20°C. | 25 μL |  |  | EMIT | 150 ng/mL; Semi-quant 36-900 ng/mL | n/a | N/A | N/A | qualitative/semi-quantitative assay | N/A |  |  |  | Emit Calibrator 0, 2, 3, 4, 5 | 10445406 10445408 10445409 10445410 10445411 | 5 weeks | Liquid |  | 3 | 5 |  | 40 | 20 |  | No |  | No | [12 empty] | 7 ng/mL | 4 ng/mL | n/a | N/A | N/A | n/a | 129.6 mAU/min | N/A | 0.3 | 0.6 | 163.9 mAU/min |  | 0.3 | 0.6 | 183.0 mAU/min |  | 0.2 | 0.6 | Viva-E Analyzer ± 50% of 150 ng/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | Traceable to Syva Emit Cal/Cont which are gravimetrically prepared standards- qualified by GC/MS within ± 10% of nominal. | 11110122_EN Rev. 03, 2019-10 | IFU Document | 0.884 | 5.625 | 3.5 | 2.75 | 0.402 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097504 |  | dr04 | 00630414596082 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Cocaine Metabolite (Coc) | Coc300 | 0337 | 1520 | 4 x 380 |  |  | 7 | 30 |  | ;#Urine;# | 7 days at 25°C; 30 days at 2–8°C or frozen at -20°C. | 25 μL |  |  | EMIT | 300 ng/mL; Semi-quant 36-900 ng/mL | n/a | N/A | N/A | qualitative/semi-quantitative assay | N/A |  |  |  | Emit Calibrator 0, 2, 3, 4, 5 | 10445406 10445408 10445409 10445410 10445411 | 5 weeks | Liquid |  | 3 | 5 |  | 40 | 20 |  | No |  | No | [12 empty] | 7 ng/mL | 4 ng/mL | n/a | N/A | N/A | n/a | 129.6 mAU/min | N/A | 0.3 | 0.6 | 163.9 mAU/min |  | 0.3 | 0.6 | 183.0 mAU/min |  | 0.2 | 0.6 | Viva-E Analyzer ± 50% of 300 ng/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | Traceable to Syva Emit Cal/Cont which are gravimetrically prepared standards- qualified by GC/MS within ± 10% of nominal. | 11110122_EN Rev. 03, 2019-10 | IFU Document | 0.884 | 5.625 | 3.5 | 2.75 | 0.402 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097526 |  | th03 | 00630414595863 | Global | Therapeutic Drug Monitoring (TDM) | No | Atellica CH Digoxin (Dgn) | Dgn | 0405 | 1600 | 4 x 400 |  |  | 10 | 30 |  | ;#Serum;#Lithium Heparin Plasma;# | 8 hours at 20–25°C; up to 7 days at 2–8°C; frozen at -20°C for up to 6 months | 15.6 μL |  |  | latex-enhanced immunoturbidimetric | 0.14 - 5.00 ng/mL | 0.18 - 6.40 nmol/L | 0.14 - 10.00 ng/mL | 0.18 - 12.80 nmol/L | 0.80–2.00 ng/mL | 1.02–2.56 nmol/L |  |  |  | Atellica CH TDM CAL | 11099439 | 7 days when protected fromlight | Lyophilized |  | 3 | 6 |  | 60 | 7 |  | No |  | Yes | [12 empty] | 0.14 ng/mL | 0.08 ng/mL | n/a | 0.10 nmol/L | 0.18 nmol/L | n/a | 1.51 ng/mL | 1.93 nmol/L | 2.4 | 2.9 | 2.68 ng/mL | 3.43 nmol/L | 3.2 | 3.5 | 4.20 ng/mL | 5.38 nmol/L | 1.8 | 2 | ADVIA Chemistry 1800 y = 0.97x + 0.18 ng/mL (y = 0.97x + 0.23 nmol/L) r = 0.994 |  | ≤ 8% change up to 500 mg/dL of hemoglobin | ≤ 7% change up to 60 mg/dL of bilirubin | ≤ 9% change up to 399 mg/dL of Intralipid | No Biotin in assay formulation |  |  | N/A | N/A | Traceable to an internal standard manufactured using highly purified material traceable to USP. | 11110144_EN Rev. 01, 2017-05 | IFU Document | 0.845 | 5.625 | 3.5 | 2.75 | 0.384 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097501 |  | dr14 | 00630414596112 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Ethyl Alcohol (ETOH) | ETOH | 0141 | 1200 | 4 x 300 |  |  | 8 | 30 |  | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#Potassium EDTA Plasma;#Sodium fluoride/potassium oxalate Plasma;#Urine;# | 2 days at room temperature; 2 weeks at 2–8°C; or indefinitely frozen at ≤ -20°C | 8 μL |  |  | Enzymatic reaction | 3.0 - 300.0 mg/dL | 0.7 - 65.1 mmol/L | 3.0 - 900.0 mg/dL | 0.7 - 195.3 mmol/L | Varied. Reportedly fatal 400.0 mg/mL | Varied. Reportedly fatal 86.8 mmol/L |  |  |  | Atellica CH TOX CAL | 11099440 | 3 days | Liquid |  | 3 | 2 | Human serum albumin | 60 | 10 |  | No |  | Yes | [12 empty] | 2.8 mg/dL | 1.3 mg/dL | n/a | 0.3 mmol/L | 0.6 mmol/L | n/a | 50.7 mg/dL | 11.0 mmol/L | 1.5 | 5.4 | 106.6 mg/dL | 23.1 mmol/L | 0.9 | 2.9 | 266.7 mg/dL | 57.9 mmol/L | 0.5 | 1.7 | Dimension RxL y = 1.05x + 1.5 mg/dL (y = 1.05x + 0.3 mmol/L) r = 0.995 |  | ≤ 6% change up to 1000 mg/dL of hemoglobin | ≤ 4% change up to 80 mg/dL of bilirubin | ≤ 1% change up to 3000 mg/dL of Intralipid | No Biotin in assay formulation |  |  | N/A | N/A | USP grade ethyl alcohol | 11110091_EN Rev. 03, 2019-07 | IFU Document | 0.779 | 5.625 | 3.5 | 2.75 | 0.354 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097516 |  | th05 | 00630414595979 | Global | Therapeutic Drug Monitoring (TDM) | No | Atellica CH Gentamicin (Gent) | Gent | 0702 | 400 | 4 x 100 | Atellica CH DRUG CAL II |  | 8 | 30 per pack |  | ;#Serum;#Lithium Heparin Plasma;# | 8 hours at room temperature; 2 days at 2–8°C; 30 days at -20°C | 3.8 μL |  |  | PETINIA | 0.5 - 12.0 µg/mL | 1.1 - 25.9 µmol/L | 0.5 - 24.0 μg/mL | 1.1 - 51.8 μmol/L | 4.0 - 10.0 μg/mL | 8.64 - 21.6 μmol/L |  |  |  | Atellica CH DRUG II CAL | 11099405 | 30 days | Liquid |  | 3 | 5 |  | 30 | 7 |  | No |  | Yes | [12 empty] | 0.4 μg/mL | 0.3 μg/mL | n/a | 0.6 μmol/L | 0.9 μmol/L | n/a | 1.5 μg/mL | 3.2 µmol/L | 5.2 | 6.5 | 4.2 μg/mL | 9.0 µmol/L | 1.7 | 2.4 | 7.8 μg/mL | 16.8 µmol/L | 1.7 | 1.8 | Dimensiony = 0.98x + 0.02 µg/mL (y = 0.98x + 0.04 µmol/L) r = 0.984 |  | ≤ 2% change up to 500 mg/dL of hemoglobin | ≤ 2% change up to 20 mg/dL of bilirubin | ≤ 6% change up to 1000 mg/dL of Intralipid | No Biotin in assay formulation |  |  | N/A | N/A | USP Gentamicin Standard | 11110134_EN Rev. 01, 2017-05 | IFU Document | 0.629 | 5.625 | 3.5 | 2.75 | 0.286 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097535 |  | th06 | 00630414006789 | Global | Therapeutic Drug Monitoring (TDM) | No | Atellica CH Lithium (Li) | Li | 1208 | 400 | 2 x 200 |  |  | 10 | 13 |  | ;#Serum;#Sodium Heparin Plasma;#Potassium EDTA Plasma;# | 24 hours at room temperature; 7 days at 4°C; 180 days at -18°C | 13 μL |  |  | Colorimetric endpoint chemistry | 0.10 - 3.00 mmol/L | n/a | 0.10 - 6.00 mmol/L | n/a | 1.00 - 1.20 mmol/L | n/a |  |  |  | Atellica CH CHEM CAL | 11099411 | 2 days | Lyophilized | Reconstituted to 3.0 mL/vial | 3 | 2 | bovine serum | 63 | 4 |  | No |  | Yes | [12 empty] | 0.07 mmol/L | 0.05 mmol/L | n/a | n/a | n/a | n/a | 0.91 mmol/L | N/A | 1.2 | 2.1 | 1.99 mmol/L |  | 0.6 | 1.3 | 2.76 mmol/L |  | 0.5 | 1 | ADVIA Chemistry 1800 y = 0.98x + 0.02 mmol/L r = 0.997 |  | ≤ 7% change up to 750 mg/dL of hemoglobin | ≤ -8% change up to 40 mg/dL of bilirubin | ≤ -4% change up to 500 mg/dL of Intralipid | No Biotin in assay formulation |  |  | n/a | n/a | NIST SRM 908b1 and 909b2 | 11110126_EN Rev. 02, 2018-05 | IFU Document | 0.455 | 2.875 | 3.5 | 2.75 | 0.207 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097520 |  | dr06 | 00630414595917 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Methadone (Mdn) | Mdn150 | 1303 | 1520 | 4 x 380 |  |  | 7 | 30 |  | ;#Urine;# | 7 days at 25°C; 30 days at 2–8°C or storedfrozen at -20°C | 23 μL |  |  | EMIT | 150 ng/mL; Semi-quant 40 - 450 ng/mL | n/a | n/a | n/a | qualitative/semi-quantitative assay | n/a |  |  |  | Emit Calibrator 0, 2, 3, 4 | 10445406 10445408 10445409 10445410 | 5 weeks | Liquid |  | 3 | 4 |  | 60 | 20 |  | No |  | No | [12 empty] | 15 ng/mL | 12 ng/mL | n/a | n/a | n/a | n/a | 106.0 mAU/min | N/A | 0.3 | 0.4 | 198.6 mAU/min | N/A | 0.3 | 0.4 | 221.6 mAU/min | N/A | 0.4 | 0.4 | Viva-E Analyzer ± 50% of 150 ng/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | Traceable to Syva Emit Cal/Cont which are gravimetrically prepared standards- qualified by GC/MS within ± 10% of nominal. | 11110138_EN Rev. 02, 2019-11 | IFU Document | 0.618 | 5.625 | 3.5 | 2.75 | 0.281 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097520 |  | dr06 | 00630414595917 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Methadone (Mdn) | Mdn300 | 1314 | 1520 | 4 x 380 |  |  | 7 | 30 |  | ;#Urine;# | 7 days at 25°C; 30 days at 2–8°C or storedfrozen at -20°C | 11.5 μL |  |  | EMIT | 300 ng/mL; Semi-quant 90 - 900 ng/mL | n/a | n/a | n/a | qualitative/semi-quantitative assay | n/a |  |  |  | Emit Calibrator 0, 3, 4, 5 | 10445406 10445409 10445410 10445411 | 5 weeks | Liquid |  | 3 | 4 |  | 60 | 20 |  | No |  | No | [12 empty] | 22 ng/mL | 16 ng/mL | n/a | n/a | n/a | n/a | 119.0 mAU/min | N/A | 0.3 | 0.4 | 225.7 mAU/min | N/A | 0.3 | 0.5 | 248.8 mAU/min | N/A | 0.3 | 0.5 | Viva-E Analyzer ± 50% of 150 ng/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | Traceable to Syva Emit Cal/Cont which are gravimetrically prepared standards- qualified by GC/MS within ± 10% of nominal. | 11110138_EN Rev. 02, 2019-11 | IFU Document | 0.618 | 5.625 | 3.5 | 2.75 | 0.281 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097527 |  | dr07 | 00630414595856 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Methadone Metabolite (EDDP) (MetMtb) | MetMtb | 1315 | 2000 | 4 x 500 |  |  | 7 | 30 |  | ;#Urine;# | 5 days at 2–8°C or ≤ -20°C | 25 μL |  |  | Enzyme Immunoassay | 1000 ng/mL | n/a | n/a | n/a | qualitative assay | n/a |  |  |  | Atellica CH MetMtb CAL | 11099437 | 30 days | Liquid |  | 3 | 1 |  | 60 | 14 |  | No |  | No | [12 empty] | n/a | n/a | n/a | n/a | n/a | n/a | 193.5 mAU/min | N/A | 0.3 | 1.3 | 221.0 mAU/min |  | 0.5 | 1.3 | n/a |  |  |  | ADVIA 1800 @ 1000 ng/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | Traceable to Syva Emit Cal/Cont which are gravimetrically prepared standards- qualified by GC/MS within ± 10% of nominal. | 11110145_EN Rev. 02, 2019-07 | IFU Document | 0.609 | 5.625 | 3.5 | 2.75 | 0.277 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097502 |  | dr08 | 00630414596105 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Opiates (Op) | Op300 | 1504 | 1520 | 4 x 380 |  |  | 7 | 30 |  | ;#Urine;# | 7 days at 25°C; 30 days at 2–8°C or storedfrozen at -20°C | 24 μL |  |  | EMIT | 300 ng/mL; Semi-quant 75-1800 ng/mL | n/a | n/a | n/a | qualitative/semi-quantitative assay | n/a |  |  |  | Emit Calibrator 0, 1, 2, 3 | 10445406 10445407 10445408 10445409 | 5 weeks | Liquid |  | 3 | 4 |  | 60 | 30 |  | No |  | No | [12 empty] | 27 ng/mL | 17 ng/mL | n/a | n/a | n/a | n/a | 105 mAU/min | N/A | 0.5 | 0.6 | 142 mAU/min |  | 0.5 | 0.6 | 154 mAU/min |  | 0.4 | 0.6 | Viva-E Analyzer ± 50% of 300 ng/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | N/A | N/A | Traceable to Syva Emit Cal/Cont which are gravimetrically prepared standards- qualified by GC/MS within ± 10% of nominal. | 11110106_EN Rev. 03, 2019-10 | IFU Document | 0.904 | 5.625 | 3.5 | 2.75 | 0.411 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097502 |  | dr08 | 00630414596105 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Opiates (Op) | Op2000 | 1505 | 1520 | 4 x 380 |  |  | 7 | 30 |  | ;#Urine;# | 7 days at 25°C; 30 days at 2–8°C or storedfrozen at -20°C | 6 μL |  |  | EMIT | 2000 ng/mL; Semi-quant 320-3800 ng/mL | n/a | n/a | n/a | qualitative/semi-quantitative assay | n/a |  |  |  | Emit Calibrator 0, 2, 3, 5 | 10445406 10445408 10445409 10445411 | 5 weeks | Liquid |  | 3 | 4 |  | 60 | 30 |  | No |  | No | [12 empty] | 179 ng/mL | 122 ng/mL | n/a | n/a | n/a | n/a | 127 mAU/min | N/A | 0.5 | 0.6 | 211 mAU/min |  | 0.5 | 0.7 | 222 mAU/min |  | 0.5 | 0.6 | Viva-E Analyzer ± 50% of 2000 ng/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | N/A | N/A | Traceable to Syva Emit Cal/Cont which are gravimetrically prepared standards- qualified by GC/MS within ± 10% of nominal. | 11110106_EN Rev. 03, 2019-10 | IFU Document | 0.904 | 5.625 | 3.5 | 2.75 | 0.411 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097509 |  | dr09 | 00630414596044 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Phencyclidine (Pcp) | Pcp25 | 1602 | 1520 | 4 x 380 |  |  | 7 | 30 |  | ;#Urine;# | 7 days unrefrigerated; 30 days at 2 – 8°C; 12 months at ≤ -20°C. | 23.5 μL |  |  | EMIT | 25 ng/mL; Semi-quant 5 - 75 ng/mL | n/a | n/a | n/a | qualitative/semi-quantitative assay | n/a |  |  |  | Emit Calibrator 0, 2, 3, 4 | 10445406 10445408 10445409 10445410 | 5 weeks | Liquid |  | 3 | 4 |  | 60 | 19 |  | No |  | No | [12 empty] | n/a | n/a | n/a | n/a | n/a | n/a | 18 ng/mL | N/A | 1.8 | 3.3 | 25 ng/mL | N/A | 1.6 | 3.4 | 30 ng/mL | N/A | 2.2 | 2.7 | GC/MS ± 50% of the 25 ng/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | Traceable to Syva Emit Cal/Cont which are gravimetrically prepared standards- qualified by GC/MS within ± 10% of nominal. | 11110127_EN Rev. 02, 2019-07 | IFU Document | 0.913 | 5.625 | 3.5 | 2.75 | 0.415 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097514 |  | th09 | 00630414595993 | Global | Therapeutic Drug Monitoring (TDM) | No | Atellica CH Phenobarbital (Phnb) | Phnb | 1603 | 400 | 4 x 100 |  |  | 8 | 30 per pack |  | ;#Serum;#Lithium Heparin Plasma;# | 8 hours at 25°C; 2 days at 2-8°C; 30 days at -20°C | 5 μL |  |  | PETINIA | 3.0 - 80.0 µg/mL | 12.9 - 344.8 µmol/L | 3.0 - 160.0 μg/mL | 12.9 - 689.6 μmol/L | 15.0-40.0 μg/mL | 64.7-172.4 μmol/L |  |  |  | Atellica CH DRUG CAL | 11099336 | 90 days | Liquid |  | 3 | 5 |  | 30 | 7 |  | No |  | Yes | [12 empty] | 2.0 μg/mL | 1.5 μg/mL | n/a | 6.5 μmol/L | 8.6 μmol/L | n/a | 14.1 μg/mL | 60.8 µmol/L | 1.4 | 2.8 | 47.0 μg/mL | 202.6 µmol/L | 1 | 1.6 | 67.6 μg/mL | 291.4 µmol/L | 0.9 | 2.3 | Dimensiony = 0.92x + 0.7 µg/mL (y = 0.92x + 3.0 µmol/L) r = 0.994 |  | ≤ 5% change up to 500 mg/dL of hemoglobin | ≤ 6% change up to 80 mg/dL of bilirubin | ≤ 10% change up to 200 mg/dL of Intralipid | No Biotin in assay formulation |  |  | n/a | n/a | USP Phenobarbital Standard | 11110132_EN Rev. 01, 2017-05 | IFU Document | 0.625 | 5.625 | 3.5 | 2.75 | 0.284 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097510 |  | th11 | 00630414596037 | Global | Therapeutic Drug Monitoring (TDM) | No | Atellica CH Phenytoin (Phny) | Phny | 1604 | 400 | 4 x 100 |  |  | 8 | 30 per pack |  | ;#Serum;#Lithium Heparin Plasma;# | 24 hours at ambient temperature; 48 hours at 2-8°C; 5 months at -20°C | 5 μL |  |  | PETINIA | 2.0 - 40.0 µg/mL | 7.9 - 158.4 µmol/L | 2.0 - 80.0 μg/mL | 7.9 - 316.8 μmol/L | 10.0-20.0 μg/mL | 39.6-79.2 μmol/L |  |  |  | Atellica CH DRUG CAL | 11099336 | 90 days | Liquid |  | 3 | 5 |  | 28 | 7 |  | No |  | Yes | [12 empty] | 0.8 μg/mL | 0.4 μg/mL | n/a | 1.6 μmol/L | 3.2 μmol/L | n/a | 8.2 μg/mL | 32.5 µmol/L | 2.2 | 3.1 | 16.0 μg/mL | 63.4 µmol/L | 1.3 | 1.8 | 35.6 μg/mL | 141.0 µmol/L | 1.1 | 1.9 | Dimension RxL y = 0.95x - 0.1 µg/mL (y = 0.95x - 0.4 µmol/L) r = 0.995 |  | ≤ 9% change up to 300 mg/dL of hemoglobin | ≤ 4% change up to 20 mg/dL of bilirubin | ≤ 7% change up to 250 mg/dL of Intralipid | No Biotin in assay formulation |  |  | n/a | n/a | USP Phenytoin Standard | 11110128_EN Rev. 01, 2017-05 | IFU Document | 0.623 | 5.625 | 3.5 | 2.75 | 0.283 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097519 |  | dr10 | 00630414595948 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Propoxyphene (Ppx) | Ppx | 2702 | 1520 | 4 x 380 |  |  | 7 | 30 |  | ;#Urine;# | 7 days at 25°C; 30 days at 2–8°C or storedfrozen at -20°C | 20 μL |  |  | EMIT | 300 ng/mL; Semi-quant 103 - 1000 ng/mL | n/a | n/a | n/a | qualitative/semi-quantitative assay | n/a |  |  |  | Emit Calibrator 0, 2, 3, 4, 5 | 10445406 10445408 10445409 10445410 10445411 | 5 weeks | Liquid |  | 3 | 5 |  | 60 | 30 |  | No |  | No | [12 empty] | 18 ng/mL | 6 ng/mL | n/a | n/a | n/a | n/a | 166 mAU/min | N/A | 0.7 | 0.9 | 215 mAU/min |  | 0.6 | 0.7 | 226 mAU/min |  | 0.5 | 0.5 | Viva-E Analyzer ± 50% of 300 ng/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | Traceable to Syva Emit Cal/Cont which are gravimetrically prepared standards- qualified by GC/MS within ± 10% of nominal. | 11110137_EN Rev. 02, 2019-10 | IFU Document | 0.911 | 5.625 | 3.5 | 2.75 | 0.414 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097523 |  | dr15 | 00630414595887 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Salicylate (Sal) | Sal | 1901 | 1200 | 4 x 300 |  |  | 10 | 30 |  | ;#Serum;#Lithium Heparin Plasma;# | 7 days at 25°C; 14 days at 2-8 °C; 6 months at ≤ -20 °C | 8 μL |  |  | Colorimetric | 3.0 - 100.0 mg/dL | 0.22 - 7.20 mmol/L | 3.0 - 400.0 mg/dL | 0.22 - 28.8 mmol/L | Toxic >30 mg/dL | n/a |  |  |  | Atellica CH TOX CAL | 11099440 | 3 days | Liquid |  | 3 | 2 |  | 180 | 21 |  | No |  | Yes | [12 empty] | 0.9 mg/dL | 0.5 mg/dL | n/a | 0.04 mmol/L | 0.06 mmol/L | n/a | 13.2 mg/dL | 0.95 mmol/L | 2.4 | 2.8 | 30.9 mg/dL | 2.2 mmol/L | 1.2 | 1.4 | 85.8 mg/dL | 6.2 mmol/L | 0.4 | 0.5 | ADVIA Chemistry 1800 Sal y = 1.01x - 0.9 mg/dL (y = 1.01x - 0.1 mmol/L) r = 0.997 r = 0.997 |  | ≤ -9% change up to 1000 mg/dL of hemoglobin | ≤ -9% change up to 30 mg/dL of bilirubin | ≤ -5% change up to 1000 mg/dL of Intralipid | No Biotin in assay formulation |  |  | n/a | n/a | Internal HPLC Standard | 11110141_EN Rev. 03, 2018-05 | IFU Document | 0.702 | 5.625 | 3.5 | 2.75 | 0.319 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097503 |  | dr11 | 00630414596099 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Cannabinoids THC (Thc) | Thc 20 | 0346 | 1520 | 4 x 380 |  |  | 7 | 30 |  | ;#Urine;# | 7 days at 25°C; 30 days at 2–8°C; 12 months at -20°C. | 25 μL |  |  | EMIT | 20 ng/mL; Semi-quant 10 - 55 ng/mL | n/a | n/a | n/a | qualitative/semi-quantitative assay | n/a |  |  |  | Emit Calibrator 0, 2, 3, 4 | 10445406 10445408 10445409 10445410 | 5 weeks | Liquid |  | 3 | 5 |  | 60 | 5 |  | No |  | No | [12 empty] | 7 ng/mL | 6 ng/mL | n/a | n/a | n/a | n/a | 19 ng/mL | N/A | 2.8 | 4.8 | 19 ng/mL |  | 6.1 | 6.5 | 23 ng/mL |  | 1.6 | 2.3 | Viva-E Analyzer ± 50% of 20 ng/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | Traceable to Syva Emit Cal/Cont which are gravimetrically prepared standards- qualified by GC/MS within ± 10% of nominal. | 11110121_EN Rev. 03, 2019-05 | IFU Document | 0.911 | 5.625 | 3.5 | 2.75 | 0.414 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097503 |  | dr11 | 00630414596099 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Cannabinoids THC (Thc) | Thc 50 | 0308 | 1520 | 4 x 380 |  |  | 7 | 30 |  | ;#Urine;# | 7 days at 25°C; 30 days at 2–8°C; 12 months at -20°C. | 13 μL |  |  | EMIT | 50 ng/mL; Semi-quant 15 - 180 ng/mL | n/a | n/a | n/a | qualitative/semi-quantitative assay | n/a |  |  |  | Emit Calibrator 0, 2, 3, 4, 5 | 10445406 10445408 10445409 10445410 10445411 | 5 weeks | Liquid |  | 3 | 5 |  | 60 | 5 |  | No |  | No | [12 empty] | 7 ng/mL | 6 ng/mL | n/a | n/a | n/a | n/a | 23 ng/mL | N/A | 3.7 | 8 | 47 ng/mL | N/A | 2 | 3.2 | 62 ng/mL | N/A | 1.2 | 4 | Viva-E Analyzer ± 50% of 50 ng/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | Traceable to Syva Emit Cal/Cont which are gravimetrically prepared standards- qualified by GC/MS within ± 10% of nominal. | 11110121_EN Rev. 03, 2019-05 | IFU Document | 0.911 | 5.625 | 3.5 | 2.75 | 0.414 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097503 |  | dr11 | 00630414596099 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Cannabinoids THC (Thc) | Thc 100 | 0345 | 1520 | 4 x 380 |  |  | 7 | 30 |  | ;#Urine;# | 7 days at 25°C; 30 days at 2–8°C; 12 months at -20°C. | 9.5 μL |  |  | EMIT | 100 ng/mL; Semi-quant 15 - 180 ng/mL | n/a | n/a | n/a | qualitative/semi-quantitative assay | n/a |  |  |  | Emit Calibrator 0, 3, 4, 5 | 10445406 10445409 10445410 10445411 | 5 weeks | Liquid |  | 3 | 5 |  | 60 | 5 |  | No |  | No | [12 empty] | 9 ng/mL | 7 ng/mL | n/a | n/a | n/a | n/a | 73 ng/mL | N/A | 2.2 | 4.5 | 98 ng/mL | N/A | 1.2 | 2.6 | 121 ng/mL | N/A | 0.7 | 2 | Viva-E Analyzer ± 50% of 100 ng/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | Traceable to Syva Emit Cal/Cont which are gravimetrically prepared standards- qualified by GC/MS within ± 10% of nominal. | 11110121_EN Rev. 03, 2019-05 | IFU Document | 0.911 | 5.625 | 3.5 | 2.75 | 0.414 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097513 |  | th12 | 00630414596006 | Global | Therapeutic Drug Monitoring (TDM) | No | Atellica CH Theophylline (Theo) | Theo | 2008 | 400 | 4 x 100 |  |  | 6 | 30 per pack |  | ;#Serum;#Lithium Heparin Plasma;# | 8 hours at 25 °C; 7 days at 2-8 °C; 90 days at -20 °C | 5.0 μL |  |  | PETINIA | 2.0 - 40.0 µg/mL | 11.1 - 222.0 µmol/L | 2.0 - 80.0 μg/mL | 11.1 - 444.0 μmol/L | 10-20 μg/mL | 56-111 μmol/L |  |  |  | Atellica CH DRUG CAL | 11099336 | 90 days | Liquid |  | 3 | 5 |  | 30 | 7 |  | No |  | Yes | [12 empty] | 0.4 μg/mL | 0.3 μg/mL | n/a | 1.7 μmol/L | 2.2 μmol/L | n/a | 7.0 μg/mL | 38.9 µmol/L | 1.4 | 2.2 | 14.8 μg/mL | 82.1 µmol/L | 0.9 | 2.2 | 27.4 μg/mL | 152.1 µmol/L | 1.5 | 2 | Dimension RxL y = 0.97x - 0.1 µg/mL (y = 0.97x - 0.6 µmol/L) r = 0.993 |  | ≤ -3% change up to 1000 mg/dL of hemoglobin | ≤ 4% change up to 80 mg/dL of bilirubin | ≤ 10% change up to 1000 mg/dL of Intralipid | No Biotin in assay formulation |  |  | n/a | n/a | USP Theophylline Standard | 11110131_EN Rev. 02, 2017-10 | IFU Document | 0.634 | 5.625 | 3.5 | 2.75 | 0.288 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097517 |  | th13 | 00630414595962 | Global | Therapeutic Drug Monitoring (TDM) | No | Atellica CH Tobramycin (Tob) | Tob | 2014 | 400 | 4 x 100 | Atellica CH DRUG CAL II |  | 9 | 30 per pack |  | ;#Serum;#Lithium Heparin Plasma;# | 8 hours at 25°C; 72 hours at 2–8°C; 30 days at -20°C. | 3.8 μL |  |  | PETINIA | 0.3 - 12.0 µg/mL | 0.6 - 25.7 µmol/L | 0.3 - 24.0 μg/mL | 0.6 - 51.4 μmol/L | 4–10 μg/mL | 8.6–21.4 µmol/L |  |  |  | Atellica CH DRUG CAL II | 11099405 | 30 days | Liquid |  | 3 | 5 |  | 30 | 7 |  | No |  | Yes | [12 empty] | 0.2 μg/mL | 0.1 μg/mL | n/a | 0.2 μmol/L | 0.4 μmol/L | n/a | 1.3 μg/mL | 2.8 µmol/L | 5 | 7.8 | 2.8 μg/mL | 6.0 µmol/L | 3.2 | 4.6 | 9.8 μg/mL | 21.0 µmol/L | 1.2 | 1.5 | Dimension RxL y = 1.01x + 0.07 µg/mL (y = 1.01x + 0.15 µmol/L) r = 0.996 |  | ≤ 2% change up to 500 mg/dL of hemoglobin | ≤ 2% change up to 20 mg/dL of bilirubin | ≤ 4% change up to 800 mg/dL of Intralipid | No Biotin in assay formulation |  |  | n/a | n/a | USP TobramycinStandard | 11110135_EN Rev. 01, 2017-05 | IFU Document | 0.638 | 5.625 | 3.5 | 2.75 | 0.29 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097511 |  | th01 | 00630414596020 | Global | Therapeutic Drug Monitoring (TDM) | No | Atellica CH Vancomycin (Vanc) | Vanc | 2202 | 400 | 4 x 100 | Atellica CH DRUG CAL II |  | 9 | 30 per pack |  | ;#Serum;#Lithium Heparin Plasma;# | 8 hours at room temperature; 2 days at 2-8°C; 30 days at -20°C | 2.5 μL |  |  | PETINIA | 3.0 - 50.0 µg/mL | 2.1 - 34.5 µmol/L | 3.0 - 100.0 μg/mL | 2.1 - 69.0 μmol/L | Peak 18 - 26 µg/mLTrough 5 - 10 µg/mL | Peak 12.4 - 17.9 µmol/LTrough 3.5 - 6.9 µmol/L |  |  |  | Atellica CH DRUG II CAL | 11099405 | 30 days | Liquid |  | 3 | 5 |  | 30 | 7 |  | No |  | Yes | [12 empty] | 0.2 μg/mL | 0.1 μg/mL | 3.0 μg/mL | 0.1 μmol/L | 0.1 μmol/L | 2.1 μmol/L | 11.2 μg/mL | 7.7 μmol/L | 1.8 | 2 | 32.7 μg/mL | 22.6 μmol/L | 1.4 | 2 | 45.8 μg/mL | 31.6 μmol/L | 1.7 | 1.8 | Dimension y = 1.04x - 1.04 µg/mL (y = 1.04x - 0.7 µmol/L) r = 0.997 |  | ≤ 2.1% change up to 600 mg/dL of hemoglobin | ≤ 0.9% change up to 20 mg/dL of bilirubin | ≤ 2.7% change up to 1000 mg/dL of Intralipid | No Biotin in assay formulation |  |  | n/a | n/a | USP material | 11110129_EN Rev. 01, 2017-05 | IFU Document | 0.638 | 5.625 | 3.5 | 2.75 | 0.29 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097512 |  | th15 | 00630414596013 | Global | Therapeutic Drug Monitoring (TDM) | No | Atellica CH Valproic Acid (VPA) | VPA | 2201 | 400 | 4 x 100 | Atellica CH DRUG CAL II |  | 8 | 30 per pack |  | ;#Serum;#Lithium Heparin Plasma;# | 8 hours at 25°C; 2 days at 2-8°C; 30 days at -20°C | 2.5 μL |  |  | PETINIA | 3.0 - 150.0 µg/mL | 20.8 - 1040.1 µmol/L | 3.0 - 300.0 μg/mL | 20.8 - 2080.2 μmol/L | 50-100 μg/mL | 346-693 μmol/L |  |  |  | Atellica CH DRUG II CAL | 11099405 | 30 days | Liquid |  | 3 | 5 |  | 30 | 7 |  | No |  | Yes | [12 empty] | 1.5 μg/mL | 0.9 μg/mL | n/a | 6.2 μmol/L | 10.4 μmol/L | n/a | 24.0 μg/mL | 166.4 μmol/L | 1.7 | 2.9 | 53.3 μg/mL | 369.6 µmol/L | 1.2 | 1.7 | 97.5 μg/mL | 676.1 μmol/L | 1.3 | 1.7 | Dimension RxL y = 1.01x - 1.4 µg/mL (y = 1.01x - 9.7 µmol/L) r = 0.997 |  | 6% bias up to 1000 mg/dL (0.62 mmol/L) of hemoglobin | ≤ 1% bias up to 80 mg/dL of bilirubin | 4% bias up to 1000 mg/dL (11.3 mmol/L) of Intralipid | No Biotin in assay formulation |  |  | n/a | n/a | USP-grade material | 11110130_EN Rev. 02, 2017-05 | IFU Document | 0.636 | 5.625 | 3.5 | 2.75 | 0.289 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097518 |  | dr05 | 00630414595955 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Ecstasy (Xtc) | Xtc300 | 0510 | 400 | 4 x 100 |  |  | 7 | 30 |  | ;#Urine;# | 7 days at 25°C; 30 days at 2–8°C or stored at -20°C. | 23 μL |  |  | EMIT | 300 ng/mL; Semi-quant 75 – 500 ng/mL | n/a | n/a | n/a | qualitative/semi-quantitative assay | n/a |  |  |  | Emit Cal/Control 0 Emit II Plus 6-AM/Ecstasy Cal/Control Levels 1, 2, 3 | 10445406 10470441 10470442 10470443 | expiration date printed on the vial label | Liquid |  | 3 | 4 |  | 60 | 15 |  | No |  | No | [12 empty] | 40 ng/mL | 32 ng/mL | n/a | n/a | n/a | n/a | 231 ng/mL | N/A | 1.5 | 2.2 | 302 ng/mL |  | 1.5 | 7.6 | 391 ng/mL |  | 3.4 | 4.7 | Viva-E Analyzer ± 50% of 300 ng/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | Traceable to Syva Emit Cal/Cont and Emit II Plus 6-AM Ecstasy which are gravimetrically prepared standards- qualified by GC/MS within ± 10% of nominal. | 11110136_EN Rev. 02, 2019-09 | IFU Document | 0.636 | 5.625 | 3.5 | 2.75 | 0.289 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097518 |  | dr05 | 00630414595955 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Ecstasy (Xtc) | Xtc500 | 0511 | 400 | 4 x 100 |  |  | 7 | 30 |  | ;#Urine;# | 7 days at 25°C; 30 days at 2–8°C or stored at -20°C. | 16 μL |  |  | EMIT | 500 ng/mL; Semi-quant 75 – 1000 ng/mL | n/a | n/a | n/a | qualitative/semi-quantitative assay | n/a |  |  |  | Emit Cal/Control 0 Emit II Plus 6-AM/Ecstasy Cal/Control Levels 1, 2, 3, 4 | 10445406 10470441 10470442 10470443 10470444 | expiration date printed on the vial label | Liquid |  | 3 | 4 |  | 60 | 15 |  | No |  | No | [12 empty] | 46 ng/mL | 35 ng/mL | n/a | n/a | n/a | n/a | 391 ng/mL | N/A | 1.9 | 2.3 | 512 ng/mL |  | 4.4 | 4.9 | 681 ng/mL |  | 7.1 | 8 | Viva-E Analyzer ± 50% of 500 ng/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | Traceable to Syva Emit Cal/Cont and Emit II Plus 6-AM Ecstasy which are gravimetrically prepared standards- qualified by GC/MS within ± 10% of nominal. | 11110136_EN Rev. 02, 2019-09 | IFU Document | 0.636 | 5.625 | 3.5 | 2.75 | 0.289 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097590 |  | ge13 | 00630414596501 | Global | General Chemistry | No | Atellica CH Albumin BCG (Alb) | AlbAlb | 1106 | 6800 | 4 x 1700 | N/A | N/A | 10 | 60 | Bromocresol green | ;#Serum;#Lithium Heparin Plasma;# | Up to 3 days at 2–8°C Frozen at -20°C for up to 60 days | 2 μL | No |  | Dye Binding: bromocresol green | 1.0 - 6 g/dL | 10 - 60 g/L | 12.0 g/dL | 120 g/L | 3.2–4.8 g/dL | 32–48 g/L | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 days | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine serum | 97 | 60 |  | No |  | Yes | N/A | 11099300 |  | 1.0 - 12.0 g/dL | 10 - 120 g/L | No | N/A | N/A | 0.0 g/dL | 0 g/L | N/A | N/A | 0.0 g/dL | 0.0 g/dL | N/A | 0 g/L | 0 g/L | N/A | 2.1 g/dL | 21 g/L | 1.7 | 2.4 | 3.4 g/dL | 34 g/L | 1.3 | 2 | 5.1 g/dL | 51 g/L | 1 | 1.6 | ADVIA Chemistry 1800 y = 0.99x + 0.1 g/dL r=0.995 n=121 | ADVIA Chemistry 1800 y = 0.99x + 1 g/L r=0.995 n=121 | ≤ 10% at 400 mg/dL (0.25 mmol/L) of hemoglobin | ≤ 10% at 30 mg/dL (513 µmol/L) of bilirubin, conjugated ≤ 10% at 30 mg/dL (513 µmol/L) of bilirubin, unconjugated | ≤ 10% at 1000 mg/dL (11.3 mmol/L) of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | BCG reference method NIST SRM 927reference materials | 11110153_EN Rev. 03, 2019-07 | IFU Document | 0.535 | 2.875 | 3.5 | 2.75 | 0.243 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097530 |  | ge15 | 00630414595825 | Global | General Chemistry | No | Atellica CH Albumin BCP (AlbP) | AlbPAlbP | 0148 | 6400 | 4 x 1600 | N/A | NA | 3 | 20 | Bromocresol Purple | ;#Serum;#Lithium Heparin Plasma;#Potassium EDTA Plasma;# | Up to 3 days at 2–8°C Frozen at -20°C for up to 30 days Thawed frozen specimens which are turbid must be clarified by centrifugation prior to testing | 4 μL | No |  | Dye Binding: bromocresol purple | 0.5 - 8 g/dL | 5 - 80 g/L | N/A | N/A | 3.4–5.0 g/dL | 34–50 g/L | N/A | N/A | N/A | Atellica CH AlbP CAL | 11099310 | 8 hours | Lyophilized | Preparation required refer to IFU | 3 | 2 | Human serum-based product containing albumin | 30 | 8 |  | No |  | No | N/A | 11099300 |  | N/A | N/A | No | N/A | N/A | 0.2 g/dL | 2 g/L | N/A | N/A | 0.2 g/dL | 0.1 g/dL | 0.4 g/dL | 1 g/L | 2 g/L | 4 g/L | 2.7 g/dL | 27 g/L | 1.17 | 1.17 | 3.6 g/dL | 36 g/L | 1.07 | 1.42 | 7.1 g/dL | 71 g/L | 0.73 | 1.02 | ADVIA Chemistry 1800 ALBP y = 0.99x + 0.01 g/dL r=0.996 n=59 | ADVIA Chemistry 1800 ALBP y = 0.99x + .1 g/L r=0.996 n=59 | ≤ 10% at 600 mg/dL (0.37 mmol/L) of hemoglobin | ≤ 10% at 30 mg/dL (513 µmol/L) of bilirubin, conjugated ≤ 10% at 30 mg/dL (513 µmol/L) of bilirubin, unconjugated | ≤ 10% at 500 mg/dL (5.65 mmol/L) of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | ERM DA470k Reference Material | 11110148_EN Rev. 01, 2017-05 | IFU Document | 0.519 | 2.875 | 3.5 | 2.75 | 0.236 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097605 |  | ge17 | 00630414596365 | Global | General Chemistry | No | Atellica CH Alanine Aminotransferase (ALT) | ALTALT | 1105 | 2550 | 3 x 850 | N/A | N/A | 10 | 90 | L-Alanine; LD (pig heart) | ;#Serum;#Lithium Heparin Plasma;# | Separated specimens may be stored for up to 7 days at 2–8°C or stored frozen for up to 30 days at -20°C or colder. Avoid repeating freezing and thawing | 25 μL | No |  | Modified IFCC | 7 - 1100 U/L | 7 - 1100 U/L | 7 - 3300 U/L | 7 - 3300 U/L | 10–49 U/L | 10–49 U/L | N/A | N/A | N/A | Atellica CH ENZ 2 CAL | 11099318 | 30 days at 2-8 °C | Liquid | Thaw the calibrator at room temperature for 30–45 minutes; Invert 10 times to mix | 3 | 2 | Bovine serum albumin: (alanine aminotransferase (porcine heart) | 131 | 30 |  | No |  | Yes | N/A | 11099300 |  | 7 - 3300 U/L | 7 - 3300 U/L | No | N/A | N/A | 7.0 U/L | 7.0 U/L | N/A | N/A | 2 U/L | 1 U/L | N/A | 1 U/L | 2 U/L | N/A | 27 U/L | 27 U/L | 1.2 | 2.8 | 79 U/L | 79 U/L | 1.7 | 2.2 | 843 U/L | 843 U/L | 0.3 | 1 | ADVIA Chemistry 1800 ALT y = 1.00x -1 U/L r=1.00 n=105 | ADVIA Chemistry 1800 ALT y = 1.00x -1 U/L r=1.00 n=105 | ≤ 10% at 450 mg/dL (0.28 mmol/L) of hemoglobin | ≤ 10% at 30 mg/dL (513 µmol/L) of bilirubin, conjugated ≤ 10% at 30 mg/dL (513 µmol/L) of bilirubin, unconjugated | ≤ 10% at 450 mg/dL (5.09 mmol/L) of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | IFCC reference method, IFCC-454 | 11110168_EN Rev. 04, 2019-06 | IFU Document | 0.719 | 4.187 | 3.5 | 2.75 | 0.327 | 10.636 | 8.89 | 6.985
Atellica® CH | 11097638 |  | ge17 | 00630414009056 | Global | General Chemistry | No | Atellica CH Alanine Aminotransferase P5P (ALTPLc) | ALTPLcALTPLc | 4035 | 2080 | 2 x 1040 | N/A | N/A | 10 | 19 per pack | L-Alanine; LD (pig heart) | ;#Serum;#Lithium Heparin Plasma;# | Separated specimens stored up to 7 days at 2–8°C; Frozen at -20°C for up to 30 days; avoid repeating freezing and thawing | 25 μL | No |  | IFCC Standardization | 9 - 1000 U/L | 9 - 1000 U/L | 6000 U/L | 6000 U/L | 7–40 U/L | 7–40 U/L | N/A | N/A | N/A | Atellica CH ENZ 2 CAL | 11099318 | 30 days | Liquid | Thaw the calibrator at room temperature for 30–45 minutes; Invert 10 times to mix | 3 | 2 | Bovine serum albumin: (alanine aminotransferase (porcine heart) | 130 | 4 |  | No |  | Yes | N/A | 11099300 |  | 9 - 6000 U/L | 9 - 6000 U/L | No | N/A | N/A | 3 U/L | 3 U/L | N/A | N/A | 3 U/L | 2 U/L | 5 U/L | 2 U/L | 3 U/L | 5 U/L | 36 U/L | 36 U/L | 2 | 2.8 | 102 U/L | 102 U/L | 1.8 | 2.7 | 218 U/L | 218 U/L | 0.6 | 1.3 | ADVIA Chemistry 1800 y = 1.02x - 0 U/L r=0.999 n=103 | ADVIA Chemistry 1800 y = 1.02x - 0 U/L r=0.999 n=103 | ≤ 10% at 200 mg/dL (0.12 mmol/L) of hemoglobin | ≤ 10% at 20 mg/dL (342 µmol/L) of bilirubin, conjugated ≤ 10% at 20 mg/dL (342 µmol/L) of bilirubin, unconjugated | ≤ 10% at 400 mg/dL (4.5 mmol/L) of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | IFCC reference method | 11111573_EN Rev. 02, 2019-11 | IFU Document | 0.636 | 5.625 | 3.5 | 2.75 | 0.289 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097603 |  | ge19 | 00630414596389 | Global | General Chemistry | No | Atellica CH Amylase (Amylas) Plasma/ Serum | AmylasAmylas | 0118 | 1050 | 3 x 350 | N/A | N/A | 9 | 30 | α-Glucosidase (microorganism); Ethylidene-4-NP-G7 | ;#Serum;#Lithium Heparin Plasma;# | Separated specimens stored up to 7 days at room temperature; Frozen at -20°C or colder for longer storage | 16 μL | No |  | Ethylidene Blocked-pNPG7 | 5 - 1500 U/L | 5 - 1500 U/L | 4500 U/L | 4500 U/L | 30–118 U/L | 30–118 U/L | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 3 | 1 | N/A | N/A | 1 |  | No |  | Yes | N/A | 11099300 |  | 5 - 4500 U/L | 5 - 4500 U/L | No | N/A | N/A | 4 U/L | 4 U/L | N/A | N/A | 4 U/L | 2 U/L | N/A | 2 U/L | 4 U/L | N/A | 43 U/L | 43 U/L | 1.9 | 2.2 | 134 U/L | 134 U/L | 0.5 | 0.9 | 1334 U/L | 1334 U/L | 0.3 | 0.4 | ADVIA Chemistry 1800 y = 1.05x - 2 U/L r=1.000 n=100 | ADVIA Chemistry 1800 y = 1.05x - 2 U/L r=1.000 n=100 | ≤ 10% at 525 mg/dL (0.336 mmol/L) of hemoglobin | ≤ 10% at 30 mg/dL (513 µmol/L) of bilirubin, conjugated ≤ 10% at 30 mg/dL (513 µmol/L) of bilirubin, unconjugated | ≤ 10% at 650 mg/dL (7.35 mmol/L) of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | IFCC reference method, IFCC-456 | 11110166_EN Rev. 02, 2019-09 | IFU Document | 0.607 | 4.187 | 3.5 | 2.75 | 0.276 | 10.636 | 8.89 | 6.985
Atellica® CH | 11097603 |  | ge19 | 00630414596389 | Global | General Chemistry | No | Atellica CH Amylase (Amylas) (urine) | AmylasAmylas | 0118 | 1050 | 3 x 350 | N/A | N/A | 9 | 30 | α-Glucosidase (microorganism); Ethylidene-4-NP-G7 | ;#Urine;# | Unstable in acid, adjust to pH of 7.0 then store refrigerated | 16 μL | No |  | Ethylidene Blocked-pNPG7 | 5 - 1500 U/L | 5 - 1500 U/L | 4500 U/L | 4500 U/L | ≤ 650 U/L | ≤ 650 U/L | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 3 | 1 | N/A | N/A | 1 |  | No |  | Yes | N/A | 11099300 |  | 5 - 4500 U/L | 5 - 4500 U/L | No | N/A | N/A | 3 U/L | 3 U/L | N/A | N/A | 3 U/L | 2 U/L | N/A | 2 U/L | 3 U/L | N/A | 51 U/L | 51 U/L | 2.2 | 2.9 | 491 U/L | 491 U/L | 0.3 | 1.1 | 1396 U/L | 1396 U/L | 0.5 | 1.5 | ADVIA Chemistry 1800 y = 1.05x + 2 U/L r=1.000 n=100 | ADVIA Chemistry 1800 y = 1.05x + 2 U/L r=1.000 n=100 | N/A | N/A | N/A | No Biotin in assay formulation | N/A | N/A | N/A | N/A | IFCC reference method, IFCC-456 | 11110166_EN Rev. 02, 2019-09 | IFU Document | 0.607 | 4.187 | 3.5 | 2.75 | 0.276 | 10.636 | 8.89 | 6.985
Atellica® CH | 11097607 |  | ge03 | 00630414596341 | Global | General Chemistry | No | Atellica CH Aspartate Aminotransferase (AST) | ASTAST | 1127 | 2550 | 3 x 850 | N/A | N/A | 10 | 90 | L-aspartic acid (580 mmol/L); MDH; LD (pig heart); α-Ketoglutarate; NADH | ;#Serum;#Lithium Heparin Plasma;# | Separated specimens stored up to 3 days at 20–25°C; up to 7 days at 2–8°C; frozen at -20°C or colder for up to 30 days | 25 μL | No |  | Modified IFCC | 8 - 1000 U/L | 8 - 1000 U/L | 6000 U/L | 6000 U/L | <34 U/L | <34 U/L | N/A | N/A | N/A | Atellica CH ENZ 2 CAL | 11099318 | 30 days | Liquid | Thaw the calibrator at room temperature for 30–45 minutes; Invert 10 times to mix | 3 | 2 | Bovine serum albumin: (alanine aminotransferase (porcine heart) | 131 | 30 |  | No |  | Yes | N/A | 11099300 |  | 8 - 6000 U/L | 8 - 6000 U/L | No | N/A | N/A | 2 U/L | 2 U/L | N/A | N/A | 2 U/L | 1 U/L | n/a | 1 U/L | 2 U/L | N/A | 30 U/L | 30 U/L | 2.9 | 3.4 | 116 U/L | 116 U/L | 1.3 | 1.7 | 861 U/L | 861 U/L | 0.4 | 0.9 | ADVIA Chemistry 1800 AST y = 1.04x - 5 U/L r=1.000 n=104 | ADVIA Chemistry 1800 AST y = 1.04x - 5 U/L r=1.000 n=104 | Avoid hemolyzed samples | ≤ 10% at 30 mg/dL (513 µmol/L) of bilirubin, conjugated ≤ 10% at 30 mg/dL (513 µmol/L) of bilirubin, unconjugated | ≤ 10% at 400 mg/dL (4.52 mmol/L) of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | IFCC reference method | 11110170_EN Rev. 03, 2020-08 | IFU Document | 0.724 | 4.187 | 3.5 | 2.75 | 0.329 | 10.636 | 8.89 | 6.985
Atellica® CH | 11097639 |  | ge03 | 00630414009063 | Global | General Chemistry | No | Atellica CH Aspartate Aminotransferase P5P (ASTPLc) | ASTPLcASTPLc | 4036 | 2080 | 2 x 1040 | N/A | N/A | 10 | 30 per pack | L-aspartic acid (580 mmol/L); MDH; LD (pig heart); α-Ketoglutarate; NADH | ;#Serum;#Lithium Heparin Plasma;# | Separated specimens stored up to 3 days at 20–25°C; up to 7 days at 2–8°C; frozen at -20°C or colder for up to 30 days | 25 μL | No |  | IFCC | 8 - 1000 U/L | 8 - 1000 U/L | 6000 U/L | 6000 U/L | 13–40 U/L | 13–40 U/L | N/A | N/A | N/A | Atellica CH ENZ 2 CAL | 11099318 | 30 days | Liquid | Thaw the calibrator at room temperature for 30–45 minutes; Invert 10 times to mix | 3 | 2 | Bovine serum albumin: aspartate aminotransferase (porcine heart)) | 130 | 11 |  | No |  | Yes | N/A | 11099300 |  | 8 - 6000 U/L | 8 - 6000 U/L | No | N/A | N/A | 1 U/L | 1 U/L | N/A | N/A | 1 U/L | 0 U/L | 7 U/L | 0 U/L | 1 U/L | 7 U/L | 34 U/L | 34 U/L | 1.6 | 1.8 | 98 U/L | 98 U/L | 1.4 | 1.5 | 238 U/L | 238 U/L | 0.9 | 1 | ADVIA Chemistry 1800; y = 1.00x - 4 U/L r=0.999 n=110 | ADVIA Chemistry 1800; y = 1.00x - 4 U/L r=0.999 n=110 | Avoid hemolyzed samples | ≤ 10% at 20 mg/dL (342 µmol/L) of bilirubin, conjugated ≤ 10% at 20 mg/dL (342 µmol/L) of bilirubin, unconjugated | ≤ 10% at 200 mg/dL (2.26 mmol/L) of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | IFCC reference method | 11111574_EN Rev. 01, 2017-05 | IFU Document | 0.64 | 5.625 | 3.5 | 2.75 | 0.291 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097595 |  | ge05 | 00630414595672 | Global | General Chemistry | No | Atellica CH Calcium (CPC) | Ca (Serum/plasma)Ca (Serum/plasma) | 0307 | 3560 | 4 x 890 | N/A | N/A | 8 | 30 | Ethanolamine buffer; o‑cresolphthalein complexone; 8‑hydroxyquinoline | ;#Serum;#Lithium Heparin Plasma;# | Up to 8 hours at room temperature; up to 2 days at 2–8°C; Frozen at -20°C or colder for up to 6 months | 16 μL | No |  | CPC | 1 - 15 mg/dL | 0.25 - 3.75 mmol/L | 30.0 mg/dL | 7.50 mmol/L | 8.3–10.6 mg/dL | 2.08–2.65 mmol/L | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 days | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 33 | 3 |  | No |  | Yes | N/A | 11099300 |  | 1 - 30.0 mg/dL | 0.25 - 7.50 mmol/L | No | N/A | N/A | 0.3 mg/dL | 0.08 mmol/L | N/A | N/A | 0.3 mg/dL | 0.2 mg/dL | N/A | 0.05 mmol/L | 0.08 mmol/L | N/A | 6.1 mg/dL | 1.53 mmol/L | 2.4 | 2.9 | 9.1 mg/dL | 2.28 mmol/L | 1.7 | 1.8 | 12.6 mg/dL | 3.15 mmol/L | 1.2 | 1.6 | ADVIA Chemistry 1800 Ca (serum) y = 0.98x + 0.00 mg/dL r=0.999 n=100 | ADVIA Chemistry 1800 Ca (serum) y = 0.98x + 0.00 mmol/L r=0.999 n=100 | ≤ 10% at 1000 mg/dL (0.62 mmol/L) of hemoglobin | ≤ 10% at 30 mg/dL (513 µmol/L) of bilirubin, conjugated ≤ 10% at 30 mg/dL (513 µmol/L) of bilirubin, unconjugated | ≤ 10% at 650 mg/dL (7.35 mmol/L) of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | NIST atomic absorption reference method, NIST SRM 915 and SRM 909b reference materials | 11110158_EN Rev. 02, 2019-10 | IFU Document | 1.133 | 5.625 | 3.5 | 2.75 | 0.515 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097595 |  | ge05 | 00630414595672 | Global | General Chemistry | No | Atellica CH Calcium (CPC) | Ca (Urine)Ca (Urine) | 0307 | 3560 | 4 x 890 | N/A | N/A | 8 | 30 | Ethanolamine buffer; o‑cresolphthalein complexone; 8‑hydroxyquinoline | ;#Urine;# | 2 days at 20–25°C, 4 days at 4–8°C, Frozen for up to 3 weeks at -20°C | 16 μL | No |  | CPC | 1 - 30 mg/dL | 0.25 - 7.50 mmol/L | 60.0 mg/dL | 15.00 mmol/L | 100–300 mg/day | 2.50–7.50 mmol/day | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 days | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 33 | 3 |  | No |  | Yes | N/A | 11099300 |  | 1 - 60.0 mg/dL | 0.25 - 15.00 mmol/L | No | N/A | N/A | 0.5 mg/dL | 0.13 mmol/L | N/A | N/A | 0.5 mg/dL | 0.3 mg/dL | N/A | 0.08 mmol/L | 0.13 mmol/L | N/A | 7.5 mg/dL | 1.88 mmol/L | 3.9 | 4.4 | 11.2 mg/dL | 2.80 mmol/L | 2.5 | 4.5 | 22.1 mg/dL | 5.53 mmol/L | 1.4 | 2.6 | ADVIA Chemistry 1800 Ca; (urine) y = 1.03x - 0.2 mg/dL r=0.999 n= 100 | ADVIA Chemistry 1800 Ca; (urine) y = 1.03x - 0.05 mmol/L r=0.999 n= 100 | N/A | N/A | N/A | No Biotin in assay formulation | N/A | N/A | N/A | N/A | NIST atomic absorption reference method, NIST SRM 915 and SRM 909b reference materials | 11110158_EN Rev. 02, 2019-10 | IFU Document | 1.133 | 5.625 | 3.5 | 2.75 | 0.515 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097609 |  | ge21 | 00630414596327 | Global | General Chemistry | No | Atellica CH Cholesterol_2 (Chol_2) | Chol_2Chol_2 | 0314 | 8400 | 4 x 2100 | N/A | N/A | 10 | 26 | 4-aminoantipyrine ; peroxidase (horseradish); cholesterol oxidase (microbial); phenolcholesterol esterase (Pseudomonas) | ;#Serum;#Lithium Heparin Plasma;# | Up to 8 hours at 25°C; Up to 2 days at 2–8°C; Frozen at -20°C or colder for longer storage | 4 μL | No |  | Colorimetric: cholesterol esterase and cholesterol oxidase conversion followed by a Trinder endpoint | 25 – 618 mg/dL | 0.65 – 16.01 mmol/L | 1236 mg/dL | 32.01 mmol/L | Low-risk levels (desirable); 200 mg/dL Moderate-risk levels (borderline) 200–239 mg/dL High-risk levels ≥ 240 mg/dL | Low-risk levels (desirable); 5.18 mmol/L Moderate-risk levels (borderline) 5.18–6.19 mmol/L High-risk levels ≥ 6.22 mmol/L | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 days | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 50 | 7 |  | No |  | Yes | N/A | 11099300 |  | 25 - 1236 mg/dL | 0.65 - 32.01 mmol/L | No | N/A | N/A | 2 mg/dL | 0.04 mmol/L | N/A | N/A | 2 mg/dL | 1 mg/dL | 25 mg/dL | 0.03 mmol/L | 0.04 mmol/L | 0.65 mmol/L | 115 mg/dL | 2.97 mmol/L | 0.8 | 1.3 | 276 mg/dL | 7.14 mmol/L | 0.5 | 0.8 | 584 mg/dL | 15.13 mmol/L | 0.5 | 0.9 | ADVIA Chemistry 1800 Chol_2 y = 0.97x + 1 mg/dL r=0.997 n=100 | ADVIA Chemistry 1800 Chol_2 y = 0.97x + 0.03 mmol/L r=0.997 n=100 | ≤ 10% at 500 mg/dL (0.311 mmol/L) of hemoglobin | ≤ 10% at 60 mg/dL (1026 µmol/L) of bilirubin, conjugated ≤ 10% at 20 mg/dL (342 µmol/L) of bilirubin, unconjugated | ≤ 10% at 1000 mg/dL (11.3 mmol/L) of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | NCEP/CDC reference method; which uses NIST SRM 909 reference materials | 11110172_EN Rev. 04, 2020-11 | IFU Document | 0.541 | 2.875 | 3.5 | 2.75 | 0.246 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097640 |  | ge22 | 00630414006642 | Global | General Chemistry | No | Atellica CH Creatine Kinase (CK_L) | CK_LCK_L | 0321 | 996 | 3 x 332 | N/A | N/A | 10 | 29 | Imidazole buffer; EDTA; ADP; AMP; Diadenosine Pentaphosphate; NADP;G-6-PDH; HK; N-Acetylcysteine; Mg; Glucose; Creatine Phosphate; | ;#Serum;#Lithium Heparin Plasma;# | Up to 4 hours at 25°C,up to 5 days at 2–8°C Frozen for up to 2 months at -20°C | 4.5 μL | No |  | Enzymatic | 15–1300 U/L | 15–1300 U/L | 7800 U/L | 7800 U/L | Males is 46–171 U/L Females is 34–145 U/L | Males is 46–171 U/L Females is 34–145 U/L | N/A | N/A | N/A | Atellica CH ENZ 3 CAL | 11099319 | 30 days | Liquid | Thaw at room temperature. 30 -45 minutes; Invert 10 times | 3 | 2 | Human serum albumin; creatine kinase MM (human heart) | 202 | 21 |  | No |  | Yes | N/A | 11099300 |  | 15 - 7800 U/L | 15 - 7800 U/L | No | N/A | N/A | 6 U/L | 6 U/L | N/A | N/A | 6 U/L | 1 U/L | 6 U/L | 1 U/L | 6 U/L | 6 U/L | 198 U/L | 198 U/L | 0.8 | 1.1 | 962 U/L | 962 U/L | 0.4 | 0.7 | 1152 U/L | 1152 U/L | 0.3 | 0.7 | ADVIA Chemistry 1800 CK_L y = 0.96x - 3.1 U/L r=1.00 n=177 | ADVIA Chemistry 1800 CK_L y = 0.96x - 3.1 U/L r=1.00 n=177 | ≤ 10% at 125 mg/dL (1.25 g/L) of hemoglobin | ≤ 10% at 60 mg/dL (1026 µmol/L) of bilirubin, conjugated ≤ 10% at 60 mg/dL (1026 µmol/L) of bilirubin, unconjugated | ≤ 10% at 1000 mg/dL (11.3 mmol/L) of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | IFCC reference method | 11110165_EN Rev. 03, 2019-09 | IFU Document | 0.656 | 4.187 | 3.5 | 2.75 | 0.298 | 10.636 | 8.89 | 6.985
Atellica® CH | 11097596 |  | ge25 | 00630414596457 | Global | General Chemistry | No | Atellica CH Creatinine_2 (Crea_2) | Crea_2Crea_2 (serum/plasma) | 0322 | 5888 | 4 x 1472 | N/A | N/A | 7 | 17 | Sodium hydroxide; Picric acid | ;#Serum;#Lithium Heparin Plasma;# | Up to 24 hours at 18–26°C; Up to 7 days at 2–8°C; Frozen at-20°C or colder for up to 3 months | 24 μL | No |  | Jaffe, alkaline picrate, kinetic with blank rate correction | 0.15 - 30.00 mg/dL | 13 - 2652 µmol/L | 60 mg/dL | 5304 μmol/L | Males 0.70–1.30 mg/dL Females 0.55–1.02 mg/dL | Males 62–115 μmol/L Females 49–90 μmol/L | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 days | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 180 | 6 |  | No |  | Yes | N/A | 11099300 |  | 0.15 - 60 mg/dL | 13 - 5304 μmol/L | No | N/A | N/A | 0.08 mg/dL | 7 μmol/L | N/A | N/A | 0.08 mg/dL | 0.03 mg/dL | 0.15 mg/dL | 3 μmol/L | 7 μmol/L | 13 μmol/L | 0.66 mg/dL | 58 µmol/L | 1.2 | 2.8 | 6.35 mg/dL | 561 µmol/L | 0.7 | 1 | 26.00 mg/dL | 2298 µmol/L | 0.2 | 0.6 | ADVIA Chemistry 1800 ECRE_2 y = 0.98x + 0.00 mg/dL r=0.999 n=140 | ADVIA Chemistry 1800 ECRE_2 y = 0.99x + 0.00 µmol/L r=0.999 n=140 | Do not use hemolyzed samples ≤ 10% at 500 mg/dL (0.311 mmol/L) of hemoglobin | ≤ 10% at 20 mg/dL (342 µmol/L) of bilirubin, conjugated ≤ 10% at 10 mg/dL (171 µmol/L) of bilirubin, unconjugated | ≤ 10% at 500 mg/dL (5.65 mmol/L) of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | IDMS reference method via correlation ofpatient samples and NIST SRM967 reference material | 11110159_EN Rev. 04, 2019-07 | IFU Document | 1.096 | 5.625 | 3.5 | 2.75 | 0.498 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097596 |  | ge25 | 00630414596457 | Global | General Chemistry | No | Atellica CH Creatinine_2 (Crea_2) | Crea_2Crea_2 (urine) | 0322 | 5888 | 4 x 1472 | N/A | N/A | 7 | 17 | Sodium hydroxide; Picric acid | ;#Urine;# | Up to 4 days at 2–8°C; Frozen for storage longer than 4 days | 24 μL | No |  | Jaffe, alkaline picrate, kinetic with blank rate correction | 3.00 -245.00 mg/dL | 265 - 21658 µmol/L | 735 mg/dL | 64,975 μmol/L | Males 950–2490 mg/day Female 600–1800 mg/day | Males 8.4–22.0 mmol/dayFemales 5.3–15.9 mmol/day | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 days | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 180 | 6 |  | No |  | Yes | N/A | 11099300 |  | 3.00 - 735 mg/dL | 265 - 64,975 μmol/L | No | N/A | N/A | 0.51 mg/dL | 45 μmol/L | N/A | N/A | 0.51 mg/dL | 0.35 mg/dL | 3.00 mg/dL | 31 μmol/L | 45 μmol/L | 265 μmol/L | 59.62 mg/dL | 5270 µmol/L | 0.3 | 0.6 | 133.31 mg/dL | 11,785 µmol/L | 0.2 | 0.7 | 188.61 mg/dL | 16,673 µmol/L | 0.3 | 0.9 | ADVIA Chemistry 1800 ECRE_2 y = 0.95x + 0.07 mg/dL r=0.999 n=109 | ADVIA Chemistry 1800 ECRE_2 y = 0.95x + 6 µmol/L r=0.999 n=109 | N/A | N/A | N/A | No Biotin in assay formulation | N/A | N/A | N/A | N/A | IDMS reference method via correlation ofpatient samples and NIST SRM967 reference material | 11110159_EN Rev. 04, 2019-07 | IFU Document | 1.096 | 5.625 | 3.5 | 2.75 | 0.498 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097635 |  | pl13 | 00630414595566 | Global | Specific Proteins | No | Atellica CH β2-Microglobulin (B2M) | B2MB2M | 0204 | 400 | 2 x 200 | N/A | N/A | 8 | 30 | Latex particles coated with antihuman β2-microglobulin antibodies (goat); BSA (bovine) | ;#Serum;#Lithium Heparin Plasma;#Potassium EDTA Plasma;# | Up to 7 days at 2–8°C; Frozen at-20°C or colder for up to 2 months | 5 μL | No |  | Latex-enhanced immunoturbidimetric assay | 0.25 – 18.00 mg/L | 0.25 – 18.00 mg/L | 90.00 mg/L | 90.00 mg/L | 1.00–2.40 mg/L for adults | 1.00–2.40 mg/L for adults | N/A | N/A | N/A | Atellica CH B2M CAL | 11099442 | 30 | Lyophilized | 1.0 mL of reagent grade water; stand for 10 minutes at room temperature; Invert to mix | 3 | 2 | Human β2‑microglobulin; bovine serum albumin (50 mg) | 60 | 21 |  | No |  | Yes | NA | 11099300 |  | 0.25 - 90.00 mg/L | 0.25 - 90.00 mg/L | No | NA | NA | 0.20 mg/L | 0.10 mg/L | NA | NA | 0.20 mg/L | 0.10 mg/L | N/A | 0.20 mg/L | 0.10 mg/L | N/A | 2.05 mg/L | 2.05 mg/L | 2.8 | 2.9 | 4.88 mg/L | 4.88 mg/L | 0.6 | 0.9 | 11.15 mg/L | 11.15 mg/L | 0.5 | 0.6 | ADVIA Chemistry 1800 B2M y = 1.02x - 0.18 mg/L r = 0.999 n = 101 | ADVIA Chemistry 1800 B2M y = 1.02x - 0.18 mg/L r = 0.999 n = 101 | ≤ 10% at 1000 mg/dL of hemoglobin | ≤ 10% at 60 mg/dL of bilirubin, conjugated ≤ 10% at 60 mg/dL of bilirubin, unconjugated | ≤ 10% at 1000 mg/dL of Intralipid | No Biotin in assay formulation | 90 mg/dL | 90 mg/dL | 90 mg/dL | 90 mg/dL | World Health Organization (WHO) 1stInternational Standard for β2-microglobulin | 11110198_EN Rev. 03, 2019-11 | IFU Document | 0.323 | 2.875 | 3.5 | 2.75 | 0.147 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097624 |  | pl02 | 00630414595665 | Global | Specific Proteins | No | Atellica CH Complement C3 (C3) | C3C3 | 0301 | 400 | 2 x 200 | N/A | N/A | 10 | 30 | Anti-C3 antibody (goat); Tris/HCI buffer; sodium chloride; sodium azide; Polyethylene glycol | ;#Serum;# | Up to 3 days at 2–8°C; Frozen at-20°C up to 3 weeks | 4.5 μL | No |  | PEG-enhanced immunoturbidimetric | 0.5 - 500.0 mg/dL | 0.01 - 5.00 g/L | 0.5 - 1125.0 mg/dL | 0.01 - 11.25 g/L | Newborns 58 –108 mg/dL 3 months 67 – 124 mg/dL 6 months 74 – 138 mg/dL 9 months 78 – 144 mg/dL 12 months 80 – 150 mg/dL 2–10 years 80 – 150 mg/dL 12–18 years 85 – 160 mg/dL 20 years 82–160 mg/dL 30 years; 84–160 mg/dL 40–70 years 90–170 mg/dL | Newborns 0.60–1.10 g/L 3 months 0.70–1.20 g/L 6 months 0.70–1.40 g/L 9 months; 0.80–1.40 g/L 12 months; 0.80–1.50 g/L 2–10 years 0.80–1.50 g/L 12–18 years 0.90–1.60 g/L 20 years; 0.80–1.60 g/L 30 years; 0.80–1.60 g/L 40–70 years 0.90–1.70 g/L | N/A | N/A | N/A | Atellica CH Liquid Specific Protein Calibrators (LSP CAL) | 11099434 | 28 days | Liquid | No preparation required | 3 | 6 | Proteins derived from human serum; sodium azide | 60 | 14 |  | No |  | Yes | Atellica CH Diluent | 11099300 |  | 0.5 - 1125.0 mg/dL | 0.01 - 11.25 g/L | No | N/A | N/A | 0.2 mg/dL | 0.00 g/L | NA | NA | 0.2 mg/dL | 0.0 mg/dL | N/A | 0.00 g/L | 0.00 g/L | NA | 80.6 mg/dL | 0.81 g/L | 0.6 | 0.2 | 158.1 mg/dL | 1.58 g/L | 0.9 | 1.4 | 346.0 mg/dL | 3.46 g/L | 1.2 | 1.4 | ADVIA Chemistry 1800: y = 0.99x + 0.6 mg/dL r=0.999n=103 | ADVIA Chemistry 1800: y = 0.99x + 0.01 g/L r=0.999 n=103 | ≤ 10% at 1000 mg/dL of hemoglobin | ≤ 10% at 25 mg/dL of bilirubin, conjugated ≤ 10% at 18.8 mg/dL of bilirubin, unconjugated | ≤ 10% at 1000 mg/dL of Lipemia (avian triglycerides) | No Biotin in assay formulation | 2100 mg/dL | 21.00 g/L | 2100 mg/dL | 21.00 g/L | IRMM reference material CRM 470 from IFCC | 11110187_EN Rev. 02, 2019-11 | IFU Document | 0.328 | 2.875 | 3.5 | 2.75 | 0.149 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097623 |  | pl15 | 00630414595658 | Global | Specific Proteins | No | Atellica CH Complement C4 (C4) | C4C4 | 0302 | 400 | 2 x 200 | N/A | N/A | 10 | 30 | Anti-C4 antibody (goat); Tris/HCI buffer; sodium chloride; sodium azide; Polyethylene glycol | ;#Serum;# | Up to 7 days at 2–8°C; Frozen at 20°C for up to up to 3 months | 6 μL | No |  | PEG-enhanced immunoturbidimetric | 0.4 - 140.0 mg/dL | 0.01 - 1.40 g/L | 0.4 - 420.0 mg/dL | 0.01 - 4.20 g/L | 12.0–36.0 mg/dL for adults | 0.12–0.36 g/L for adults | N/A | N/A | N/A | Atellica CH Liquid Specific Protein Calibrators (LSP CAL) | 11099434 | 28 days | Liquid | No preparation required | 3 | 6 | Proteins derived from human serum; sodium azide | 60 | 14 |  | No |  | Yes | Atellica CH Diluent | 11099300 |  | 0.4 - 420.0 mg/dL | 0.01 - 4.20 g/L | No | N/A | N/A | 0.2 mg/dL | 0.00 g/L | NA | NA | 0.2 mg/dL | 0.1 mg/dL | N/A | 0.00 g/L | 0.00 g/L | NA | 15.7 mg/dL | 0.16 g/L | 0.7 | 1.8 | 42.3 mg/dL | 0.42 g/L | 0.9 | 1.3 | 100.8 mg/dL | 1.01 g/L | 1.1 | 1.4 | ADVIA Chemistry 1800: y = 0.96x + 0.4 mg/dL r=0.999 n=103 | ADVIA Chemistry 1800: y = 0.96x + 0.00 g/L r=0.999 n=103 | ≤ 10% at 750 mg/dL of hemoglobin | ≤ 10% at 18.8 mg/dL of bilirubin, conjugated ≤ 10% at 18.8 mg/dL of bilirubin, unconjugated | ≤ 10% at 1000 mg/dL of Lipemia (avian triglycerides) | No Biotin in assay formulation | 300 mg/dL | 3.00 g/L | 300 mg/dL | 3.00 g/L | e IRMM reference material CRM 470 from IFCC | 11110186_EN Rev. 02, 2019-11 | IFU Document | 0.326 | 2.875 | 3.5 | 2.75 | 0.148 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097631 |  | pl17 | 00630414596198 | Global | Specific Proteins | No | Atellica CH C-Reactive Protein_2 (CRP_2) | CRP_2CRP_2 | 0320 | 1000 | 2 x 500 | N/A | N/A | 8 | 30 | CRP antibody synthetic latex (rabbit); sodium azide; sodium chloride; EDTA disodium salt hydrate; sodium azide | ;#Serum;#Lithium Heparin Plasma;# | Up to 3 days at 4–8°C; Frozen at-20°C or colder for up to 6 months | 4 μL | No |  | Latex-enhanced immunoturbidimetric assay | 0.4 - 30.4 mg/dL | 4 - 304 mg/L | 91.2 mg/dL | 912 mg/L | < 1.0 mg/dL for adults | < 10 mg/L for adults | N/A | N/A | N/A | Atellica CH C‑Reactive Protein_2 Calibrators (CRP_2 CAL) | 11099430 | 60 days | Liquid | No preparation required | 3 | 6 | Recombinant human c-reactive protein (CRP) in a stabilized protein matrix; sodium azide | 60 | 30 |  | No |  | Yes | Atellica CH Diluent | 11099300 |  | 0.4 - 91.2 mg/dL | 4 - 912 mg/L | No | N/A | N/A | 0.0 mg/dL | 0.0 mg/L | N/A | N/A | 0.0 mg/dL | 0.0 mg/dL | 0.3 mg/dL | 0.0 mg/L | 0.0 mg/L | 3 mg/L | 0.7 mg/dL | 7 mg/L | 2.3 | 2.3 | 3.3 mg/dL | 33 mg/L | 1.4 | 1.4 | 6.5 mg/dL | 65 mg/L | 0.7 | 0.9 | ADVIA Chemistry 1800 y = 1.02x - 0.1 mg/dL r=0.998 | ADVIA Chemistry 1800 y = 1.02x - 1 mg/L r=0.998 | ≤ 10% at 1000 mg/dL of hemoglobin | ≤ 10% at 60 mg/dL of bilirubin, conjugated ≤ 10% at 60 mg/dL of bilirubin, unconjugated | ≤ 10% at 1000 mg/dL of Intralipid | No Biotin in assay formulation | 100.0 mg/dL | 1000 mg/L | 100.0 mg/dL | 1000 mg/L | IRMM reference material ERMDA-470. | 11110194_EN Rev. 02, 2019-10 | IFU Document | 0.458 | 2.875 | 3.5 | 2.75 | 0.208 | 7.302 | 8.89 | 6.985
Atellica® CH | 11099315 |  | ge23 | 00630414596518 | Global | General Chemistry | No | Atellica CH A‑LYTE Integrated Multisensor (IMT Na K Cl) | Cl (serum/plasma)Cl (serum/plasma) | 0312 | 20000 | 4 x 5000 | A-LYTE IMT Standard A SMN 11099304 A‑LYTE IMT Standard B + Salt Bridge SMN 11099306 A‑LYTE IMT Diluent SMN11099305 A‑LYTE IMT Dilution Check SMN 11099325 | N/A | 18 sec | 14 days or 5000 samples per multisensor | N/A | ;#Serum;#Lithium Heparin Plasma;# | up to 7 days at 2–8°C; Frozen for up to 30 days at -20°C. | 25 μL | No |  | Indirect Integrated Multisensor Technology (IMT) | 50 - 200 mmol/L | 50 - 200 mEq/L | N/A | N/A | Serum-98–107 mmol/L Plasma-98–107 mmol/L | Serum-98–107 mEq/L Plasma-98–107 mEq/L | N/A | N/A | N/A | A‑LYTE IMT Standard A; A‑LYTE IMT Standard B + Salt Bridge | 11099304 11099306 | 30 days onboard | Liquid | No preparation required | 3 | 2 | N/A | automatic calibration every 4 hours | automatic calibration every 4 hours |  | Yes |  | No | N/A | 11099305 |  | N/A | N/A | No | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 38.3 mmol/L | N/A | N/A | 38.3 mEq/L | 78.1 mmol/L | 78.1 mEq/L | 0.9 | 1.5 | 114 mmol/L | 114 mEq/L | 0.7 | 0.9 | 189 mmol/L | 189 mEq/L | 0.5 | 0.7 | ADVIA Chemistry 1800 Cl y = 1.00x + 0.00 mmol/L r=0.998 n=108 | ADVIA Chemistry 1800 Cl y = 1.00&nbsp; + 0.00 mEq/L r=0.998 n=108 | ≤ 10% at 750 mg/dL (0.47 mmol/L)) of hemoglobin | ≤ 10% at 60 mg/dL of bilirubin, conjugated ≤ 10% at 60 mg/dL of bilirubin, unconjugated | ≤ 10% at 3000 mg/dL of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | coulometric reference method, NIST Reference Serum and materials | 11109447_EN Rev. 03, 2019-09 | IFU Document | 0.167 | 6.125 | 2.19 | 3.75 | 0.076 | 15.557 | 5.562 | 9.525
Atellica® CH | 11099315 |  | ge23 | 00630414596518 | Global | General Chemistry | No | Atellica CH A‑LYTE Integrated Multisensor (IMT Na K Cl) | Cl (urine)Cl (urine) | 0312 | 20000 | 4 x 5000 | A-LYTE IMT Standard A SMN 11099304 A‑LYTE IMT Standard B + Salt Bridge SMN 11099306 A‑LYTE IMT Diluent SMN11099305 A‑LYTE IMT Dilution Check SMN 11099325 | N/A | 27 sec | 14 days or 5000 samples per multisensor | N/A | ;#Urine;# | up to 7 days at 2–8°C or stored frozen for up to 30 days at -20°C. 24 hour urine collection for sodium, potassium, and chloride should be made without addition of preservatives. Store refrigerated at 2–8°C or frozen for delayed analysis. | 25 μL | No |  | Indirect Integrated Multisensor Technology (IMT) | 20 - 330 mmol/L | 20 - 330 mEq/L | 2-fold manual urine dilution | 2-fold manual urine dilution | 110–250 mmol/L/24 hr | 110–250 mEq/L/24 hr | N/A | N/A | N/A | A‑LYTE IMT Standard A; A‑LYTE IMT Standard B + Salt Bridge | 11099304 11099306 | 30 days onboard | Liquid | No preparation required | 3 | 2 | N/A | automatic calibration every 4 hours | automatic calibration every 4 hours |  | Yes |  | No | N/A | 11099305 |  | 2-fold manual urine dilution | 2-fold manual urine dilution | No | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 15.8 mmol/L | N/A | N/A | 15.8 mEq/L | 42.7 mmol/L | 42.7 mEq/L | 0.6 | 1.7 | 194 mmol/L | 194 mEq/L | 0.5 | 0.9 | 280 mmol/L | 280 mEq/L | 0.4 | 0.8 | ADVIA Chemistry 1800 Cl y = 0.99x - 0.28 mmol/L r=0.999 n=102 | ADVIA Chemistry 1800 Cl y = 0.99x - 0.28 mEq/L r=0.999 n=102 | ≤ 10% at 500 mg/dL of hemoglobin | ≤ 10% at 60 mg/dL of bilirubin, conjugated ≤ 10% at 60 mg/dL of bilirubin, unconjugated | ≤ 10% at 2000 mg/dL of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | coulometric reference method, NIST Reference Serum and materials | 1109447_EN Rev. 03, 2019-09 | IFU Document | 0.167 | 6.125 | 2.19 | 3.75 | 0.076 | 15.557 | 5.562 | 9.525
Atellica® CH | 11099315 |  | ge34 | 00630414596518 | Global | General Chemistry | No | Atellica CH A‑LYTE Integrated Multisensor (IMT Na K Cl) | K (serum/plasma)K (serum/plasma) | 1607 | 20000 | 4 x 5000 | A-LYTE IMT Standard A SMN 11099304 A‑LYTE IMT Standard B + Salt Bridge SMN 11099306 A‑LYTE IMT Diluent SMN11099305 A‑LYTE IMT Dilution Check SMN 11099325 | N/A | 18 sec | 14 days or 5000 samples per multisensor | N/A | ;#Serum;#Lithium Heparin Plasma;#Urine;# | up to 7 days at 2–8°C or stored frozen for up to 30 days at -20°C. | 25 μL | No |  | Indirect Integrated Multisensor Technology (IMT) | 1 - 10 mmol/L | 1 - 10 mEq/L | N/A | N/A | Serum-3.5–5.1 mmol/L Plasma-3.4–4.5 mmol/L | Serum-3.5–5.1&nbsp; mEq/L ; Plasma-3.4–4.5 mEq/L | N/A | N/A | N/A | A‑LYTE IMT Standard A; A‑LYTE IMT Standard B + Salt Bridge | 11099304 11099306 | 30 days onboard | Liquid | No preparation required | 3 | 2 | N/A | automatic calibration every 4 hours | automatic calibration every 4 hours |  | Yes |  | No | N/A | 11099305 |  | N/A | N/A | No | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 0.429 mmol/L | N/A | N/A | 0.429 mEq/L | 3.95 mmol/L | 3.95 mEq/L | 0.5 | 0.9 | 5.92 mmol/L | 5.92 mEq/L | 0.5 | 0.8 | 7.39 mmol/L | 7.39 mEq/L | 0.5 | 0.8 | ADVIA Chemistry 1800 y = 0.99x - 0.04 mmol/L r=1.000 n=103 | ADVIA Chemistry 1800 y = 0.99x - 0.04 mEq/L r=1.000 n=103 | N/A | <10% up to 60 mg/dL of bilirubin, conjugated and unconjugated | <10% up to 3000 mg/dL of Intralipid® | No Biotin in assay formulation | N/A | N/A | N/A | N/A | flame emission spectrophotometry reference method | 1109447_EN Rev. 03, 2019-09 | IFU Document | 0.167 | 6.125 | 2.19 | 3.75 | 0.076 | 15.557 | 5.562 | 9.525
Atellica® CH | 11099315 |  | ge34 | 00630414596518 | Global | General Chemistry | No | Atellica CH A‑LYTE Integrated Multisensor (IMT Na K Cl) | K (urine)K (urine) | 1607 | 20000 | 4 x 5000 | A-LYTE IMT Standard A SMN 11099304 A‑LYTE IMT Standard B + Salt Bridge SMN 11099306 A‑LYTE IMT Diluent SMN11099305 A‑LYTE IMT Dilution Check SMN 11099325 | N/A | 27 sec | 14 days or 5000 samples per multisensor | N/A | ;#Urine;# | 7 days at 2–8°C or stored frozen for up to 30 days at -20°C.24 hour urine collection for sodium, potassium, and chloride should be made without addition of preservatives. Store refrigerated at 2–8°C or frozen for delayed analysis. | 25 μL | No |  | Indirect Integrated Multisensor Technology (IMT) | 2 - 300 mmol/L | 2 - 300 mEq/L | 2-fold manual urine dilution | 2-fold manual urine dilution | Urine: 25–125 mmol/L/24 hr | Urine: 25–125 mEq/L/24 hr | N/A | N/A | N/A | A‑LYTE IMT Standard A; A‑LYTE IMT Standard B + Salt Bridge | 11099304 11099306 | 30 days onboard | Liquid | No preparation required | 3 | 2 | N/A | automatic calibration every 4 hours | automatic calibration every 4 hours |  | Yes |  | No | N/A | 11099305 |  | N/A | N/A | No | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 1.03 mmol/L | N/A | N/A | 1.03 mEq/L | 31.2 mmol/L | 31.2 mEq/L | 0.5 | 0.9 | 73.7 mmol/L | 73.7 mEq/L | 0.5 | 0.9 | 258 mmol/L | 258 mEq/L | 0.8 | 1.1 | ADVIA Chemistry 1800 y = 1.03x - 0.67 mmol/L r=0.998 n=105 | ADVIA Chemistry 1800 y = 1.03x - 0.67 mEq/L r=0.998 n=105 | <10% up to 500 mg/dL of hemoglobin | ≤ 10% change up to 60 mg/dL of bilirubin, conjugated and unconjugated | <10% change up to 2000 mg/dL of Intralipid® | No Biotin in assay formulation | N/A | N/A | N/A | N/A | flame emission spectrophotometry reference method | 1109447_EN Rev. 03, 2019-09 | IFU Document | 0.167 | 6.125 | 2.19 | 3.75 | 0.076 | 15.557 | 5.562 | 9.525
Atellica® CH | 11099315 |  | ge38 | 00630414596518 | Global | General Chemistry | No | Atellica CH A‑LYTE Integrated Multisensor (IMT Na K Cl) | Na (serum/plasma)Na (serum/plasma) | 1903 | 20000 | 4 x 5000 | A-LYTE IMT Standard A SMN 11099304 A‑LYTE IMT Standard B + Salt Bridge SMN 11099306 A‑LYTE IMT Diluent SMN11099305 A‑LYTE IMT Dilution Check SMN 11099325 | N/A | 18 sec | 14 days or 5000 samples per multisensor | N/A | ;#Serum;#Lithium Heparin Plasma;#Urine;# | up to 7 days at 2–8°C or Frozen for up to 30 days at -20°C. | 25 μL | No |  | Indirect Integrated Multisensor Technology (IMT) | 50 - 200 mmol/L | 50 - 200 mEq/L | N/A | N/A | Serum/ Plasma: 50–200 mmol/L | Serum-136–145 mEq/L Plasma-136–145 mEq/L | N/A | N/A | N/A | A‑LYTE IMT Standard A; A‑LYTE IMT Standard B + Salt Bridge | 11099304 11099306 | 30 days onboard | Liquid | No preparation required | 3 | 2 | N/A | automatic calibration every 4 hours | automatic calibration every 4 hours |  | Yes |  | No | N/A | 11099305 |  | N/A | N/A | No | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 42.4 mmol/L | N/A | N/A | 42.4 mEq/L | 73.6 mmol/L | 73.6 mEq/L | 0.5 | 0.8 | 116 mmol/L | 116 mEq/L | 0.6 | 0.1 | 156 mmol/L | 156 mEq/L | 0.5 | 0.9 | ADVIA Chemistry 1800 y = 1.00x - 2.00 mmol/L r=0.995 n=106 | ADVIA Chemistry 1800 y = 1.00x - 2.00 mEq/L r=0.995 n=106 | <10% up to 750 mg/dL of hemoglobin | <10% up to 60 mg/dL of bilirubin, conjugated and unconjugated | <10% up to 3000 mg/dL of Intralipid® | No Biotin in assay formulation | N/A | N/A | N/A | N/A | flame emission spectrophotometry reference method | 1109447_EN Rev. 03, 2019-09 | IFU Document | 0.167 | 6.125 | 2.19 | 3.75 | 0.076 | 15.557 | 5.562 | 9.525
Atellica® CH | 11099315 |  | ge38 | 00630414596518 | Global | General Chemistry | No | Atellica CH A‑LYTE Integrated Multisensor (IMT Na K Cl) | Na (urine)Na (urine) | 1903 | 20000 | 4 x 5000 | A-LYTE IMT Standard A SMN 11099304 A‑LYTE IMT Standard B + Salt Bridge SMN 11099306 A‑LYTE IMT Diluent SMN11099305 A‑LYTE IMT Dilution Check SMN 11099325 | N/A | 27 sec | 14 days or 5000 samples per multisensor | N/A | ;#Serum;#Lithium Heparin Plasma;#Urine;# | 7 days at 2–8°C or stored frozen for up to 30 days at -20°C.24 hour urine collection for sodium, potassium, and chloride should be made without addition of preservatives. Store refrigerated at 2–8°C or frozen for delayed analysis. | 25 μL | No |  | Indirect Integrated Multisensor Technology (IMT) | 10 - 300 mmol/L | 10 - 300 mEq/L | 2-fold manual urine dilution | 2-fold manual urine dilution | 40–220 mmol/L/24 hr | 40–220 mEq/L/24 hr | N/A | N/A | N/A | A‑LYTE IMT Standard A; A‑LYTE IMT Standard B + Salt Bridge | 11099304 11099306 | 30 days onboard | Liquid | No preparation required | 3 | 2 | N/A | automatic calibration every 4 hours | automatic calibration every 4 hours |  | Yes |  | No | N/A | 11099305 |  | N/A | N/A | No | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 5.35 mmol/L | N/A | N/A | 5.35 mEq/L | 32.2 mmol/L | 32.2 mEq/L | 1 | 1.5 | 177 mmol/L | 177 mEq/L | 0.7 | 1 | 284 mmol/L | 284 mEq/L | 0.6 | 1.1 | ADVIA Chemistry 1800 y = 1.01x + 0.97 mmol/L r=1.000 n=101 | ADVIA Chemistry 1800 y = 1.01x + 0.97 mEq/L r=1.000 n=101 | <10% up to 500 mg/dL of hemoglobin | <10% up to 60 mg/dL of bilirubin, conjugated and unconjugated | <10% up to 2000 mg/dL of Intralipid® | No Biotin in assay formulation | N/A | N/A | N/A | N/A | flame emission spectrophotometry reference method | 1109447_EN Rev. 03, 2019-09 | IFU Document | 0.167 | 6.125 | 2.19 | 3.75 | 0.076 | 15.557 | 5.562 | 9.525
Atellica® IM | 10995441 |  | on02 | 00630414597706 | OUS | Oncology | No | Atellica IM Alpha Fetoprotein (AFP) | Atellica IM AFP | AFP | 100 | 1x100 | Atellica IM CAL D | Atellica IM Multi‑Diluent 2 Atellica IM AFP MCM (SMN 10995443) | 10 | 28 | Rabbit polyclonal anti-AFP antibody | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;#Amniotic fluid;# | RT 8 hours; 2–8°C after 8 hours; Freeze at ≤ -20°C after 48 hours; Freeze only 1 time | 10 μL | No |  | 2‑site sandwich immunoassay | 1.3–1000.0 ng/mL | 1.08–830.00 IU/mL |  |  |  |  | 98.4% of the apparently healthy subjects had AFP values < 8.1 ng/mL (serum) |  |  | Atellica IM CAL D | 10995509 (2-pack) 10995510 (6-pack) | Recon at 2–8°C 28 days Recon at RT 4 hrs | Lyophilized | Add 2.0 mL of special reagent water. Let the vials stand for 15–20 minutes at room temperature | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of alphafetoprotein (human) and carcinoembryonic (human) | 35 | 28 |  | No |  | Yes | Atellica IM Multi‑Diluent 2 | 10995644 |  | 1.3 - 200,000 ng/mL | 1.08–166,000 IU/mL | No | NA | NA | 1.3 ng/mL | 1.08 IU/mL | NP | NP | 1.8 ng/mL | 1.2 ng/mL | 4.5 ng/mL | 1.00 IU/mL | 1.49 IU/mL | 3.74 IU/mL | 8.2 ng/mL (serum) | 6.80 IU/mL (serum) | 6.3 | 9 | 67.5 ng/mL | 56.03 IU/mL | 3.2 | 4.5 | 182.7 ng/mL | 151.66 IU/mL | 3.1 | 3.8 | ADVIA Centaur AFP (Serum) y = 0.958x + 0.121 ng/mL r= 0.997 ADVIA Centaur AFP (Amniotic Fluid) y = 1.05x + 0.027 μg/mL r= 0.99 | ADVIA Centaur AFP (Serum) y = 0.958x + 0.100 IU/mL r= 0.997 | NSI up to 500 mg/dL of hemoglobin | NSI up to 20 mg/dL of bilirubin | NSI up to 1000 mg/dL of triglycerides | No Biotin in assay formulation | 1,000,000 ng/mL | 830,000 IU/mL |  |  | WHO, human AFP (72/225) using highly purified AFP | 10995270_EN Rev. 03, 2020-06 | IFU Document | 0.211 | 6.969 | 2.087 | 1.457 | 0.096 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995442 |  | on02 | 00630414597713 | OUS | Oncology | No | Atellica IM Alpha Fetoprotein (AFP) | Atellica IM AFP | AFP | 500 | 5 x 100 | Atellica IM CAL D | Atellica IM Multi‑Diluent 2, Atellica IM AFP MCM | 10 | 28 | Rabbit polyclonal anti-AFP antibody | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;#Amniotic fluid;# | RT 8 hours; 2–8°C after 8 hours; Freeze at ≤ -20°C after 48 hours; Freeze only 1 time | 10 μL | No |  | 2‑site sandwich immunoassay | 1.3–1000.0 ng/mL | 1.08–830.00 IU/mL |  |  |  |  | 98.4% of the apparently healthy subjects had AFP values < 8.1 ng/mL (serum) |  |  | Atellica IM CAL D | 10995509 (2-pack) 10995510 (6-pack) | Recon at 2–8°C 28 daysRecon at RT 4 hrs | Lyophilized | Add 2.0 mL of special reagent water. Let the vials stand for 15–20 minutes at room temperature | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of alphafetoprotein (human) and carcinoembryonic (human) | 35 | 28 |  | No |  | Yes | Atellica IM Multi‑Diluent 2 | 10995644 |  | 1.3 - 200,000 ng/mL | 1.08–166,000 IU/mL | No | NA | NA | 1.3 ng/mL | 1.08 IU/mL | NP | NP | 1.8 ng/mL | 1.2 ng/mL | 4.5 ng/mL | 1.00 IU/mL | 1.49 IU/mL | 3.74 IU/mL | 8.2 ng/mL (serum) | 6.80 IU/mL (serum) | 6.4 | 9.1 | 67.5 ng/mL | 56.03 IU/mL | 3.2 | 4.5 | 182.7 ng/mL | 151.66 IU/mL | 3.1 | 3.8 | ADVIA Centaur AFP (Serum) y = 0.958x + 0.121 ng/mL r= 0.997 ADVIA Centaur AFP (Amniotic Fluid) y = 1.05x + 0.027 μg/mL r= 0.99 | ADVIA Centaur AFP (Serum) y = 0.958x + 0.100 IU/mL r= 0.997 | NSI up to 500 mg/dL of hemoglobin | NSI up to 20 mg/dL of bilirubin | NSI up to 1000 mg/dL of triglycerides | No Biotin in assay formulation | 1,000,000 ng/mL | 830,000 IU/mL |  |  | WHO, human AFP (72/225) using highly purified AFP | 10995270_EN Rev. 03, 2020-06 | IFU Document | 0.845 | 6.969 | 2.087 | 4.37 | 0.384 | 17.7 | 5.3 | 11.1
Atellica® IM | 10995471 |  | ca01 | 00630414597980 | OUS | Cardiac | No | Atellica IM B‑Type Natriuretic Peptide (BNP) | Atellica IM BNP | BNP | 100 |  | Atellica IM BNP CAL (SMN 10995473) | Atellica IM BNP QC (SMN 10995475) Atellica IM BNP MCM (SMN 10995474) | 10 | 42 |  | ;#EDTA Plasma;# | RT not recommended, at 2–8°C up to 24 hours; at ≤ -20°C after 24 hours up to 9 months, test within 8 hours after thawing | 100 μL | No |  |  | 2.0–5000.0 pg/mL | 0.58–1445.00 pmol/L |  |  |  |  | <100 pg/mL (w/o heart failure) | See IFU for expected results by age | See IFU for expected results by NYHA class | Atellica IM BNP CAL | 10995473 | Recon at 2–8°C 5 days, Recon at ≤ -20°C 60 days; thaw 1 time, Recon at RT 8 hrs | Lyophilized | 2.0 mL/vial | 3 for lot calibration, 2 for pack recalibration | 2 |  | 61 | 42 |  | No |  | Yes | Atellica IM Multi‑Diluent 15 | 11200798 |  | 2.0–50000.0 pg/mL | 0.58–14450.00 pmol/L | Yes | Atellica IM BNP QC | 10995475 | 1.4 pg/mL | 0.40 pmol/L | <1.5 pg/mL | <0.43 pmol/L | 2.4 pg/mL | 1.1 pg/mL | NP | 0.32 pmol/L | 0.69 pmol/L | NP | 12.2 pg/mL | 3.5 pmol/L | 2.6 | 3.5 | 654.6 pg/mL | 189.2 pmol/L | 1.8 | 2.8 | 4325.2 pg/mL | 1250.0pg/mL | 1.5 | 2.5 | ADVIA Centaur BNP y = 0.965x - 0.3 pg/mL r=0.997 | ADVIA Centaur BNP y = 0.965x - 0.091 pmol/L r=0.997 | ≤ 10% change up to 100 mg/dL of hemoglobin | ≤ 5% change up to 25 mg/dL of unconjugated bilirubin≤ 7% change up 25 mg/dL conjugated bilirubin | ≤ 5% change up to 800 mg/dL of triglycerides,1000 mg/dL of cholesterol | ≤ 10% change up to 30 ng/mL of biotin | 100,000 pg/mL | 28,900 pmol/L |  |  | traceable to internal standard, synthetic human BNP (amino acid 77–108) | 10995281_EN Rev. 04, 2019-07 | IFU Document | 0.253 | 6.969 | 2.087 | 1.457 | 0.115 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995472 |  | ca01 | 00630414597997 | OUS | Cardiac | No | Atellica IM B‑Type Natriuretic Peptide (BNP) | Atellica IM BNP | BNP | 500 |  | Atellica IM BNP CAL (SMN 10995473) | Atellica IM BNP QC (SMN 10995475) Atellica IM BNP MCM (SMN 10995474) | 10 | 42 |  | ;#EDTA Plasma;# | RT not recommended, at 2–8°C up to 24 hours; at ≤ -20°C after 24 hours up to 9 months, test within 8 hours after thawing | 100 μL | No |  |  | 2.0–5000.0 pg/mL | 0.58–1445.00 pmol/L |  |  |  |  | <100 pg/mL (w/o heart failure) | See IFU for expected results by age |  | Atellica IM BNP CAL | 10995473 | Recon at 2–8°C 5 days, Recon at ≤ -20°C 60 days; thaw 1 time, Recon at RT 8hrs | Lyophilized | 2.0 mL/vial | 3 for lot calibration, 2 for pack recalibration | 2 |  | 61 | 42 |  | No |  | Yes | Atellica IM Multi‑Diluent 15 | 11200798 |  | 2.0–50000.0 pg/mL | 0.58–14450.00 pmol/L | Yes | Atellica IM BNP QC | 10995475 | 1.4 pg/mL | 0.40 pmol/L | <1.5 pg/mL | <0.43 pmol/L | 2.4 pg/mL | 1.1 pg/mL | NP | 0.32 pmol/L | 0.69 pmol/L | NP | 12.2 pg/mL | 3.5 pmol/L | 2.6 | 3.5 | 654.6 pg/mL | 189.2 pmol/L | 1.8 | 2.8 | 4325.2 pg/mL | 1250.0pg/mL | 1.5 | 2.5 | ADVIA Centaur BNP y = 0.965x - 0.3 pg/mL r=0.997 | ADVIA Centaur BNP y = 0.965x - 0.091 pmol/L r=0.997 | ≤ 10% change up to 100 mg/dL of hemoglobin | ≤ 5% change up to 25 mg/dL of unconjugated bilirubin≤ 7% change up 25 mg/dL conjugated bilirubin | ≤ 5% change up to 800 mg/dL of triglycerides,1000 mg/dL of cholesterol | ≤ 10% change up to 30 ng/mL of biotin | 100,000 pg/mL | 28,900 pmol/L |  |  | traceable to internal standard, synthetic human BNP (amino acid 77–108) | 10995281_EN Rev. 04, 2019-07 | IFU Document | 0.891 | 6.969 | 2.087 | 4.37 | 0.405 | 17.7 | 5.3 | 11.1
Atellica® IM | 10995523 |  | on06 | 00630414598505 | Global | Oncology | No | Atellica IM Carcinoembryonic Antigen (CEA) | Atellica IM CEA | CEA | 100 | 1 x 100 | Atellica IM CAL D | Atellica IM CEA DIL (SMN 10995525) Atellica IM CEA MCM (SMN 10995526) | 14 | 14 | Rabbit polyclonal anti-CEA antibody (~400 ng/mL) | ;#Serum;# | RT 8 hrs 2–8°C after 8 hours; ≤ -20°C after 48 hours; Freeze only 1 time | 50 μL | No |  | 2-site sandwich immunoassay | 0.50–100.00 ng/mL | 0.50–100.00 μg/L |  |  |  |  | See IFU |  |  | Atellica IM CAL D | 10995509 (2-pack) 10995510 (6-pack) | Recon at 2–8°C 28 daysRecon at RT 4 hrs | Lyophilized | 2.0 mL of special reagent water. Let the vials stand for 15–20 minutes at room temperature | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of alphafetoprotein (human) and carcinoembryonic antigen (human) | 29 | 21 |  | No |  | Yes | Atellica IM CEA DIL | 10995525 |  | 0.50-10,000 ng/mL | 0.50-10,000 μg/L | No | NA | NA | 0.26 ng/mL | 0.26 μg/L | NP | NP | 0.88 pg/mL | 0.13 ng/mL | NP | 0.13 μg/L | 0.88 μg/L | NP | 3.12 ng/mL | 3.12 μg/L | 5 | 6 | 33.35 pg/mL | 33.35 μg/L | 1.9 | 2.9 | 84.81 ng/mL | 84.81 μg/L | 1.7 | 2.6 | ADVIA Centaur CEA y = 1.020x - 0.112 ng/mL r=0.997 | ADVIA Centaur CEA y = 1.020x - 0.112 μg/L r=0.997 | ≤ 5% change up to 500 mg/dL of hemoglobin | ≤ 5% change up to 20 mg/dL of unconjugated/conjugated bilirubin | ≤ 5% change up to 1000 mg/dL of triglycerides | No Biotin in assay formulation | 100,000 ng/mL | 100,000 μg/L |  |  | traceable to an internal standard manufactured using highlypurified material | 10995312_EN Rev. 03, 2020-02 | IFU Document | 0.209 | 6.969 | 2.087 | 1.457 | 0.095 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995524 |  | on06 | 00630414598512 | Global | Oncology | No | Atellica IM Carcinoembryonic Antigen (CEA) | Atellica IM CEA | CEA | 500 | 5 x 100 | Atellica IM CAL D | Atellica IM CEA DIL, Atellica IM CEA MCM (SMN 10995526) | 14 | 14 | Rabbit polyclonal anti-CEA antibody (~400 ng/mL) | ;#Serum;# | RT 8 hrs 2–8°C after 8 hours; ≤ -20°C after 48 hours; Freeze only 1 time | 50 μL | No |  | 2-site sandwich immunoassay | 0.50–100.00 ng/mL | 0.50–100.00 μg/L |  |  |  |  | See IFU |  |  | Atellica IM CAL D | 10995509 (2-pack) 10995510 (6-pack) | Recon at 2–8°C 28 days Recon at RT 4 hrs | Lyophilized | 2.0 mL of special reagent water. Let the vials stand for 15–20 minutes at room temperature | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of alphafetoprotein (human) and carcinoembryonic antigen (human) | 29 | 21 |  | No |  | Yes | Atellica IM CEA DIL | 10995525 |  | 0.50-10,000 ng/mL | 0.50-10,000 μg/L | No | NA | NA | 0.26 ng/mL | 0.26 μg/L | NP | NP | 0.88 pg/mL | 0.13 ng/mL | NP | 0.13 μg/L | 0.88 μg/L | NP | 3.12 ng/mL | 3.12 μg/L | 5 | 6 | 33.35 pg/mL | 33.35 μg/L | 1.9 | 2.9 | 84.81 ng/mL | 84.81 μg/L | 1.7 | 2.6 | ADVIA Centaur CEA y = 1.020x - 0.112 ng/mL r=0.998 | ADVIA Centaur CEA y = 1.020x - 0.112 μg/L r=0.998 | ≤ 5% change up to 500 mg/dL of hemoglobin | ≤ 5% change up to 20 mg/dL of unconjugated/conjugated bilirubin | ≤ 5% change up to 1000 mg/dL of triglycerides | No Biotin in assay formulation | 100,000 ng/mL | 100,000 μg/L |  |  | traceable to an internal standard manufactured using highlypurified material | 10995312_EN Rev. 03, 2020-02 | IFU Document | 0.843 | 6.969 | 2.087 | 4.37 | 0.383 | 17.7 | 5.3 | 11.1
Atellica® IM | 10995544 |  | on01 | 00630414598680 | OUS | Oncology | No | Atellica IM Complexed Prostate‑Specific Antigen(cPSA) | Atellica IM cPSA | cPSA | 100 | 1X100 | Atellica IM CAL Y Atellica IM cPSA PRE (SMN 10995546) | Atellica IM cPSA QC, Atellica IM Multi‑Diluent 2, Atellica IM cPSA MCM (SMN 10995545) | 14 | 28 | Goat polyclonal anti‑PSA antibody | ;#Serum;# | RT 8 hrs, 2–8°C after 8hrs, ≤ -20°C after 48 hours. Freeze only 1 time | 35 μL | No |  | 2-site sandwich immunoassay | 0.03–100.00 ng/mL | 0.03–100.00 μg/L |  |  |  |  | See IFU |  |  | Atellica IM CAL Y | 10995520 | Recon at 2–8°C 21 days, Recon at RT 8 hrs | Lyophilized | 2.0 mL of special reagent water, let the vials stand for 15–20 minutes at room temperature. | 3 for lot calibration, 2 for pack recalibration | 2 | PSA (human); goat serum | 63 | 28 |  | No |  | Yes | Atellica IM Multi‑Diluent 2 | 10995644 |  | 0.03- 500 ng/mL | 0.03-500 μg/L | Yes | Atellica IM cPSA QC | 10995547 | 0.03 ng/mL | 0.03 μg/L | NP | NP | 0.10 ng/mL | 0.05 ng/mL | 0.10 ng/mL | 0.05 μg/L | 0.10 μg/L | 0.10 μg/L | 1.07 ng/ml | 1.07 μg/L | 1.4 | 2.5 | 47.28 ng/mL | 47.28 μg/L | 1.2 | 2.5 | 86.8 ng/mL | 86.8 μg/L | 1.6 | 2.7 | ADVIA Centaur cPSA y = 0.97x + 0.182 ng/mL r=0.999 | ADVIA Centaur cPSA y = 0.97x + 0.182 μg/L r=0.999 | ≤ 5% change up to 500 mg/dL of hemoglobin | ≤ 5% change up to 40 mg/dL of bilirubin | ≤ 5% change up to 1000 mg/dL of triglycerides | No Biotin in assay formulation | 9137 ng/mL | 9137 μg/L |  |  | traceable to an internal standard manufactured using highlypurified PSA-ACT | 10995337_EN Rev. 02, 2019-10 | IFU Document | 0.222 | 6.969 | 2.087 | 1.457 | 0.101 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995561 |  | re04 | 00630414598857 | Global | Reproductive Endocrinology | No | Atellica IM Enhanced Estradiol (eE2) | Atellica IM eE2 | eE2 | 100 | 1 x 100 | Atellica IM CAL 30 | Atellica IM eE2 DIL, Atellica IM eE2 MCM (SMN 10995564) | 14 | 42 | Sheep monoclonal anti-estradiol antibody | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | RT 20 hours, 2–8°C after 20 hours, ≤ -20°C after 48 hours up to 6 months, Freeze only 1 time | 80 μL | No |  | Competitive assay | 11.80–3000.00 pg/mL | 43.60–11,010.00 pmol/L |  |  |  |  | See IFU |  |  | Atellica IM CAL 30 | 10995495 (2 pack) 10995497 (6 pack) | Recon at 2–8°C 14 days, Recon at RT 4hrs | Lyophilized | 2.0 mL of special reagent water. Let the vials stand for 15–20 minutes at room temperature | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of USP‑grade estradiol, testosterone, cortisol, and progesterone charcoal-stripped, defibrinated human plasma | 22 | 21 |  | No |  | Yes | Atellica IM eE2 DIL | 10995563 |  | 11.8-15.000 pg/mL | 43.60–11,010.00 pmol/L | No | NA | NA | NP | NP | 18.81 pg/mL | 69.03 pmol/L | 11.54 pg/mL | 7.33 pg/mL | NP | 26.91 pmol/L | 42.36 pmol/L | NP | 19.88 pg/mL | 72.96 pmol/L | 8.2 | 14.7 | 290.49 ng/mL | 1066.10 pmol/L | 1.6 | 3.6 | 2799.58 pg/mL | 10274.46 pmol/L | 1.9 | 2.7 | ADVIA Centaur eE2 (serum) y = 0.96x - 1.23 pg/mL r=0.998 | ADVIA Centaur eE2 (serum) y = 0.96x - 4.52 pmol/L r=0.998 | ≤ 10% change up to 250 mg/dL of hemoglobin | ≤ 10% change upto 20 mg/dL of conjugated/unconjugated bilirubin | ≤ 10% change up to 1000 mg/dL of triglycerides | ≤ 10% change up to 1500 ng/mL of biotin | NA | NA |  |  | standardized using internal standards manufactured analytically which are traceable to ID-GC/MS | (OUS )10995342_EN Rev. 04, 2019-07, (US) 11200383_EN Rev. 04, 2020-02 | IFU Document | 0.273 | 6.969 | 2.48 | 2.913 | 0.124 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995562 |  | re04 | 00630414598864 | Global | Reproductive Endocrinology | No | Atellica IM Enhanced Estradiol (eE2) | Atellica IM eE2 | eE2 | 500 | 5 x 100 | Atellica IM CAL 30 | Atellica IM eE2 DIL, Atellica IM eE2 MCM (SMN 10995564) | 14 | 42 | Sheep monoclonal anti-estradiol antibody | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | RT 20 hours, 2–8°C after 20 hours, ≤ -20°C after 48 hours up to 6 months, Freeze only 1 time | 80 μL | No |  | Competitive assay | 11.80–3000.00 pg/mL | 43.60–11,010.00 pmol/L |  |  |  |  | See IFU |  |  | Atellica IM CAL 30 | 10995495 (2 pack) 10995497 (6 pack) | Recon at 2–8°C 14 days, Recon at RT 4hrs | Lyophilized | 2.0 mL of special reagent water. Let the vials stand for 15–20 minutes at room temperature | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of USP‑grade estradiol, testosterone, cortisol, and progesterone charcoal-stripped, defibrinated human plasma | 22 | 21 |  | No |  | Yes | Atellica IM eE2 DIL | 10995563 |  | 11.8-15.000 pg/mL | 43.60–11,010.00 pmol/L | No | NA | NA | NP | NP | 18.81 pg/mL | 69.03 pmol/L | 11.54 pg/mL | 7.33 pg/mL | NP | 26.91 pmol/L | 42.36 pmol/L | NP | 19.88 pg/mL | 72.96 pmol/L | 8.2 | 14.7 | 290.49 ng/mL | 1066.10 pmol/L | 1.6 | 3.6 | 2799.58 pg/mL | 10274.46 pmol/L | 1.9 | 2.7 | ADVIA Centaur eE2 (serum) y = 0.96x - 1.23 pg/mL r=0.998 | ADVIA Centaur eE2 (serum) y = 0.96x - 4.52 pmol/L r=0.998 | ≤ 10% change up to 250 mg/dL of hemoglobin | ≤ 10% change upto 20 mg/dL of conjugated/unconjugated bilirubin | ≤ 10% change up to 1000 mg/dL of triglycerides | ≤ 10% change up to 1500 ng/mL of biotin | NA | NA |  |  | standardized using internal standards manufactured analytically which are traceable to ID-GC/MS | (OUS )10995342_EN Rev. 04, 2019-07, (US) 11200383_EN Rev. 04, 2019-06 | IFU Document | 1.052 | 10.748 | 2.48 | 6.063 | 0.478 | 27.3 | 6.3 | 15.4
Atellica® IM | 10995577 |  | on07 | 00630414599014 | OUS | Oncology | No | Atellica IM Free Prostate‑Specific Antigen (fPSA) | Atellica IM fPSA | fPSA | 50 | 1 x 50 | Atellica IM fPSA CAL | Atellica IM Multi‑Diluent 2 | 10 | 28 |  | ;#Serum;# | RT 8 hrs2–8°C after 8 hours ≤ -20°C after 48 hours;Freeze only 1 time | 30 μL | No |  |  | 0.01–25.00 ng/mL | 0.01–25.00 μg/L |  |  |  |  | See IFU |  |  | Atellica IM fPSA CAL | 10995578 | Recon at 2–8°C 21 days, Recon at RT 8hrs | Lyophilized | 2.0 mL of special reagent water. Let the vials stand for 15–20 minutes at room temperature | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of fPSA (human), goat serum | 54 | 28 |  | No |  | Yes | Atellica IM Multi‑Diluent 2 | 10995644 |  | 0.01-125 ng/mL | 0.01-125 μg/L | No | NA | NA | NP | NP | NP | NP | 0.01 ng/mL | 0.01 ng/mL | 0.01 ng/mL | 0.01 μg/L | 0.01 μg/L | 0.01 μg/L | 0.05 ng/mL | 0.05 μg/L | 4.6 | 6 | 3.6 ng/mL | 3.6 μg/L | 2.2 | 2.9 | 12.09 ng/mL | 12.09 μg/L | 2.1 | 3.3 | ADVIA Centaur fPSA y = 0.93x - 0.024 ng/mL (μg/L) r=0.999 | ADVIA Centaur fPSA y = 0.93x - 0.024 μg/L r=0.999 | ≤ 5% change up to 1000 mg/dL of hemoglobin | ≤ 5% change up to 25 mg/dL of conjugated/unconjugated bilirubin | ≤ 5% change up to 1000 mg/dL of triglycerides | ≤ 5% change in results up to up to 1500 ng/mL | 15,355 ng/mL | 15,355 μg/L |  |  | traceable to World Health Organization (WHO) InternationalStandard (96/668) | 10995345_EN Rev. 03, 2019-10 | IFU Document | 0.174 | 6.969 | 2.087 | 1.457 | 0.079 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995589 |  | ty02 | 00630414599137 | Global | N/AThyroid | No | Atellica IM Free Thyroxine (FT4) | Atellica IM FT4FT4 | FT4 | 50 | 1 x 50 | Atellica IM CAL A 2- pack SMN 10995500 Atellica IM CAL A 6- pack SMN 10995501 | Atellica IM FT4 MCM SMN 10995590 | 14 | 21 | acridinium-ester-labeled T4 and biotinylated rabbit poloclonal anti-T4 antibody | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | RT 8 hrs, 2–8°C after 8 hours; ≤ -20°C after 48 hours; Freeze only 1 time | 25 μL | No |  | Competitive immunoasay using direct chemiluminescent technology | 0.1 – 12.0 ng/dL | 1.3 – 154.8 pmol/L | NA | NA | 0.89- 1.76 ng/dL | 11.5 - 22.7 pmol/L | NA | NA | NA | Atellica IM CAL A | 10995500 (2-pack) 10995501 (6-pack) | 28 days recon at 2–8°C, 4 hours recon at RT | Lyophilized | Add 5.0 mL of reagent water, let the calibrators stand for 15–20 minutes at RT. Gently swirl and invert the vials until homogeneous | 3 for lot calibration, 2 for pack recalibration | 2 | human plasma with sodium azide, preservatives and protein stabilizers, | 21 | 7 |  | No |  | No | NA | NA |  | NA | NA | No | Use commercially available Controls | NA | 0.1 ng/dL | 1.3 pmol/L | NP | NP | 0.2 ng/dL | 0.1 ng/dL | NP | 1.3 pmol/L | 2.6 pmol/L | NP | 0.4 ng/dL | 5.7 pmol/L | 4.7 | 6.8 | 5.1 ng/dL | 65.8 pmol/L | 2.1 | 4.1 | 10.7 ng/dL | 138.3 pmol/L | 2.5 | 3.5 | ADVIA Centaur FT4 y = 1.018x - 0.02 ng/dL r = 0.981 | ADVIA Centaur FT4 y = 1.018x - 2.6 pmol/L r = 0.981 | NSI up to 300 mg/dL of hemoglobin | NSI up to 20 mg/dL of bilirubin | NSI up to 1000 mg/dL of triglycerides | NSI up to 3500 ng/mL of biotin | NA | NA | NA | NA | Internal standard manufactured using U.S.P.(United States Pharmacopeia) material | 10995348_EN Rev. 06, 2020-11 | IFU Document | 0.187 | 6.969 | 2.087 | 1.457 | 0.085 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995588 |  | ty02 | 00630414599120 | Global | N/AThyroid | No | Atellica IM Free Thyroxine (FT4) | Atellica IM FT4FT4 | FT4 | 250 | 5 x 50 | Atellica IM CAL A 2- pack SMN 10995500 Atellica IM CAL A 6- pack SMN 10995501 | Atellica IM FT4 MCM (SMN 10995590) | 14 | 21 | acridinium-ester-labeled T4 and biotinylated rabbit polyclonal anti- T4 antibody | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | RT 8 hrs, 2–8°C after 8 hours; ≤ -20°C after 48 hours; Freeze only 1 time | 25 μL | No |  | Competitive immunoassay using direct chemiluminescent technology | 0.1 – 12.0 ng/dL | 1.3 – 154.8 pmol/L | NA | NA | 0.89- 1.76 ng/dL | 11.5 - 22.7 pmol/L | NA | NA | NA | Atellica IM CAL A | 10995500 (2-pack) 10995501 (6-pack) | 28 days recon at 2–8°C, 4 hours recon at RT | Lyophilized | Add 5.0 mL of reagent water, let the calibrators stand for 15–20 minutes at RT. Gently swirl and invert the vials until homogeneous | 3 for lot calibration, 2 for pack recalibration | 2 | human plasma with sodium azide, preservatives and protein stabilizers, | 21 | 7 |  | No |  | No | NA | NA |  | NA | NA | No | Use commercially available Controls | NA | 0.1 ng/dL | 1.3 pmol/L | NP | NP | 0.2 ng/dL | 0.1 ng/dL | NP | 1.3 pmol/L | 2.6 pmol/L | NP | 0.4 ng/dL | 5.7 pmol/L | 4.7 | 6.8 | 5.1 ng/dL | 65.8 pmol/L | 2.1 | 4.1 | 10.7 ng/dL | 138.3 pmol/L | 2.5 | 3.5 | ADVIA Centaur FT4 y = 1.018x - 0.02 ng/dL r = 0.981 | ADVIA Centaur FT4 y = 1.018x - 2.6 pmol/L r = 0.981 | NSI up to 300 mg/dL of hemoglobin | NSI up to 20 mg/dL of bilirubin | NSI up to 1000 mg/dL of triglycerides | NSI up to 3500 ng/mL of biotin | NA | NA | NA | NA | Internal standard manufactured using U.S.P.(United States Pharmacopeia) material | 10995348_EN Rev. 06, 2020-11 | IFU Document | 0.737 | 6.969 | 2.087 | 4.37 | 0.335 | 17.7 | 5.3 | 11.1
Atellica® IM | 10995690 |  | re01 | 00630414600031 | OUS | Reproductive Endocrinology | No | Atellica IM Total hCG (ThCG) | Atellica IM ThCG | ThCG | 90 | 1 x 90 | Atellica IM CAL B | Atellica IM ThCG DIL, Atellica IM ThCG MCM (SMN 10995694) | 10 | 21 | Goat polyclonal anti-hCG antibody and Mouse monoclonal anti-hCG antibody | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | RT 8 hours;2–8°C after 8 hours; ≤ -20°C after 48 hours; Freeze only 1 time | 50 μL | No |  | 2-site sandwich immunoassay using direct chemiluminescent | 2.0–1000.0 mIU/mL | 2.0–1000.0 IU/L |  |  |  |  | non-pregnant females 1.5–4.2 mIU/mL (IU/L) | postmenopausal females 1.8–10.1 mIU/mL (IU/L) |  | Atellica IM CAL B | 10995503 (2-pack) 10995504 (6-pack) | Recon at 2–8°C 28 days Recon at RT 4hrs | Lyophilized | Add 5.0 mL of special reagent water into each vial Let the vials stand for 15–20 min at RT | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of digoxin,follicle-stimulating hormone (FSH), luteinizinghormone (LH), prolactin, human chorionic gonadotropin (hCG), and TSH; equine serum | 34 | 28 |  | No |  | Yes | Atellica IM ThCG DIL | 10995691 - 2 pack10995692 - 6 pack10995693 - vial |  | 2.0-200,000.0 mIU/mL | 2.0-200,000.0 IU/L | No | NA | NA | 1.9 mIU/mL | 1.9 IU/L | NA | NA | 2.8 mIU/mL | 1.5 mIU/mL | 2.8 mIU/mL | 1.5 IU/L | 2.8 IU/L | 2.8 IU/L | 2.6 mIU/mL | 2.6 IU/L | 8.5 | 13.4 | 13.0 mIU/mL | 13.0 IU/L | 2.1 | 4.2 | 791.2 mIU/mL | 791.2 IU/L | 1.6 | 2.2 | ADVIA Centaur Total hCG y = 0.994x + 2.973 mIU/mL r=0.999 | ADVIA Centaur Total hCG y = 0.994x + 2.973 IU/L r=0.999 | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up 40 mg/dL conjugated/unconjugated bilirubin | ≤ 10% change up to 3000 mg/dL Intralipid | No Biotin in assay formulation | 400,000 mIU/mL | 400,000 IU/L |  |  | World Health Organization (WHO) 4th IS 75/589 reference material | 10995426_EN Rev. 03, 2020-02 | IFU Document | 0.205 | 6.969 | 2.087 | 1.457 | 0.093 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995689 |  | re01 | 00630414600024 | OUS | Reproductive Endocrinology | No | Atellica IM Total hCG (ThCG) | Atellica IM ThCG | ThCG | 450 | 5 x 90 | Atellica IM CAL B | Atellica IM ThCG DIL Atellica IM ThCG MCM (SMN 10995694) | 10 | 21 | Goat polyclonal anti-hCG antibody and Mouse monoclonal anti-hCG antibody | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | RT 8 hours;2–8°C after 8 hours; ≤ -20°C after 48 hours; Freeze only 1 time | 50 μL | No |  | 2-site sandwich immunoassay using direct chemiluminometric | 2.0–1000.0 mIU/mL | 2.0–1000.0 IU/L |  |  |  |  | non-pregnant females 1.5–4.2 mIU/mL (IU/L) | postmenopausal females 1.8–10.1 mIU/mL (IU/L) |  | Atellica IM CAL B | 10995503 (2-pack) 10995504 (6-pack) | Recon at 2–8°C 28 days Recon at RT 4hrs | Lyophilized | Add 5.0 mL of special reagent water into each vial Let the vials stand for 15–20 min at RT | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of digoxin,follicle-stimulating hormone (FSH), luteinizinghormone (LH), prolactin, human chorionic gonadotropin (hCG), and TSH; equine serum | 34 | 28 |  | No |  | Yes | Atellica IM ThCG DIL | 10995691 - 2 pack10995692 - 6 pack10995693 - vial |  | 2.0-200,000.0 mIU/mL | 2.0-200,000.0 IU/L | No | NA | NA | 1.9 mIU/mL | 1.9 IU/L | NA | NA | 2.8 mIU/mL | 1.5 mIU/mL | 2.8 mIU/mL | 1.5 IU/L | 2.8 IU/L | 2.8 IU/L | 2.6 mIU/mL | 2.6 IU/L | 8.5 | 13.4 | 13.0 mIU/mL | 13.0 IU/L | 2.1 | 4.2 | 791.2 mIU/mL | 791.2 IU/L | 1.6 | 2.2 | ADVIA Centaur Total hCG y = 0.994x + 2.973 mIU/mL (IU/L) r=0.999 | ADVIA Centaur Total hCG y = 0.994x + 2.973 IU/L r=0.999 | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up 40 mg/dL conjugated/unconjugated bilirubin | ≤ 10% change up to 3000 mg/dL Intralipid | No Biotin in assay formulation | 400,000 mIU/mL | 400,000 IU/L |  |  | World Health Organization (WHO) 4th IS 75/589 reference material | 10995426_EN Rev. 03, 2020-02 | IFU Document | 0.821 | 6.969 | 2.087 | 4.37 | 0.373 | 17.7 | 5.3 | 11.1
Atellica® IM | 10995696 |  | ca02 | 00630414600093 | OUS | Cardiac | No | Atellica IM Troponin I Ultra (TnI‑Ultra) | Atellica IM TnI‑Ultra | TnI_UL | 100 |  | Atellica IM APW3 (SMN 10998580) | Atellica IM Multi‑Diluent 11 (SMN 10995642) Atellica IM TnI‑Ultra MCM (SMN 10995698) | 10 | 28 | 1 Goat polyclonal anti-troponin I antibody / 2 Mouse monoclonal | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | 4 hrs at RT, 24 hrs at 2–8°C, 1 month at ≤ -20°C, Freeze only 1 time | 100 μL | No |  | direct chemiluminescent technology | 0.006–30.000 ng/mL | 0.006–30.000 μg/L |  |  |  |  | 99th percentile 0.02 ng/mL (μg/L) (Plasma) | See IFU for expected results by gender |  | Atellica IM TnI-Ultra Cal | Provided in kit | 4 hrs recon at RT, 24 hrs recon at 2–8°C, 30 days recon at ≤ -20°C; thaw 1 time | Lyophilized | Reconstituted 2.0 mL/vial | 3 for lot calibration, 2 for pack recalibration | 2 | bovine cTnI, goat serum | 66 | 28 |  | No |  | Yes | Atellica IM Multi‑Diluent 11 | 10995642 2-pack 1099564 1-vial |  | 0.006–300.000 ng/mL | 0.006–300.000 μg/L | No | NA | NA | 0.006 ng/ml | 0.006 μg/L | 0.015 ng/ml | 0.015 μg/L | 0.015 ng/mL | 0.007 ng/mL | NP | 0.007 μg/L | 0.015 μg/L | NP | 0.074 ng/mL (plasma) | 0.074 μg/L (plasma) | 3 | 4.7 | 0.533 ng/mL (plasma) | 0.533 μg/L (plasma) | 1.6 | 4.3 | 21.270 ng/mL(plasma) | 21.270 μg/L (plasma) | 1.6 | 2.2 | ADVIA Centaur TnI‑Ultra (Lithium Heparin) y = 1.05x + 0.022 ng/mL r=1.00 | ADVIA Centaur TnI‑Ultra (Lithium Heparin) y = 1.05x + 0.022 μg/L r=1.00 | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 40 mg/dL of conjugated bilirubin, upto 60 mg/dL of unconjugated bilirubin | ≤ 10% change up to 2000 mg/dL of Intralipid | ≤ 10% change in results up to 10 ng/mL. | 7606 ng/mL | 7606 μg/L |  |  | NIST SRM 2921 | 10995428_EN Rev. 05, 2019-08 | IFU Document | 0.323 | 6.969 | 2.48 | 2.913 | 0.147 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995697 |  | ca02 | 00630414600109 | OUS | Cardiac | No | Atellica IM Troponin I Ultra (TnI‑Ultra) | Atellica IM TnI‑Ultra | TnI_UL | 500 |  | Atellica IM APW3 (SMN 10998580) | Atellica IM Multi‑Diluent 11 (SMN 10995642) Atellica IM TnI‑Ultra MCM (SMN 10995698) | 10 | 28 | 1 Goat polyclonal anti-troponin I antibody / 2 Mouse monoclonal | ;#Serum;#Sodium Heparin Plasma;#Heparinized Plasma;#EDTA Plasma;# | 4 hrs at RT, 24 hrs at 2–8°C, 1 month at ≤ -20°C, Freeze only 1 time | 100 μL | No |  | direct chemiluminescent technology | 0.006–30.000 ng/mL | 0.006–30.000 μg/L |  |  |  |  | Combined gender 99th percentile using lithium heparin sample type - 0.02 ng/mL(μg/L) | See IFU for expected results by gender |  | Atellica IM TnI-Ultra Cal | Provided in Kit | 4 hrs recon at RT, 24 hrs recon at 2–8°C, 30 days recon at ≤ -20°C; thaw 1 time | Lyophilized | Reconstituted to 2.0 mL/vial | 3 for lot calibration, 2 for pack recalibration | 2 | bovine cTnI, goat serum | 66 | 28 |  | No |  | Yes | Atellica IM Multi‑Diluent 11 | 10995642 - 2 pack 10995641 - vial |  | 0.006–300.000 ng/mL | 0.006–300.000 μg/L | No | NA | NA | 0.006 ng/ml | 0.006 μg/L | 0.015 ng/ml | 0.015 μg/L | 0.015 ng/mL | 0.007 ng/mL | NP | 0.007 μg/L | 0.015 μg/L | NP | 0.074 ng/mL (plasma) | 0.074 μg/L (plasma) | 3 | 4.7 | 0.533 ng/mL (plasma) | 0.533 μg/L (plasma) | 1.6 | 4.3 | 21.270 ng/mL (plasma) | 21.270 μg/L (plasma) | 1.6 | 2.2 | ADVIA Centaur TnI‑Ultra (Lithium Heparin) y = 1.05x + 0.022 ng/mL r=1.00 | ADVIA Centaur TnI‑Ultra (Lithium Heparin) y = 1.05x + 0.022 μg/L r=1.00 | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 40 mg/dL of conjugated bilirubin, up to 60 mg/dL of unconjugated bilirubin | ≤ 10% change up to 2000 mg/dL of Intralipid | ≤ 10% change in results up to 10 ng/mL. | 7606 ng/mL | 7606 μg/L |  |  | NIST SRM 2921 | 10995428_EN Rev. 05, 2019-08 | IFU Document | 1.08 | 6.969 | 2.126 | 6.811 | 0.491 | 17.7 | 5.4 | 17.3
Atellica® IM | 10995714 |  | an01 | 00630414600277 | Global | N/AAnemia | No | Atellica IM Vitamin B12 (VB12) | VB12 | VB12 | 100 | 1 X100 | Atellica IM VB12 DTT/REL (SMN 0995718) Atellica IM T3/T4/VB12 ANC (2pk) (SMN 10995682) Atellica IM T3/T4/VB12 ANC (6pk) (SMN 10995683) | Atellica IM VB12 MCM (SMN 10995717) | 14 | 18 | purified hog intrinsic factor | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | 8 hours at RT 2–8°C after 8 hours; ≤ -20°C after 48 hours | 100 μL | No |  | competitive immunoassay using direct chemiluminescent technology | 45–2000 pg/mL | 33–1476 pmol/L | 45 – 20,000 pg/mL | 33 – 14,760 pmol/L | 211 to 911 pg/mL | 156 to 672 pmol/L | NA | NA | NA | Atellica IM CAL C | 10995506 (2-pack) 10995507 (6-pack) | 28 days recon at 2–8°C 4 hours recon at RT | Lyophilized | Add 5.0 mL of reagent water, let the calibrators stand for 15–20 minutes at RT | 3 for lot calibration, 2 for pack recalibration | 2 | buffered HSA with sodium azide, preservatives and protein stabilizers, | 30 | 18 |  | No |  | Yes | Atellica IM VB12 DIL | 10995716 |  | 45–20000 pg/mL | 33-14760 pmol/L | No | NA | NA | 38 pg/mL | 28 pmol/L | NA | NA | 54 pg/mL | 38 pg/mL | NP | 28 pmol/L | 40 pmol/L | NP | 146 pg/mL (LI heparin) | 108 pmol/L (LI heparin) | 2.5 | 9.2 | 368 pg/mL (serum) | 271 pmol/L (serum) | 2 | 6.4 | 486 pg/mL (serum) | 358 pmol/L (serum) | 2.3 | 5.6 | ADVIA Centaur VB12 y = 1.00x + 7 pg/mL r=0.994 | ADVIA Centaur VB12 y = 1.00x + 5 pmol/L r=0.994 | NSI up to 150 mg/dL of hemoglobin | NSI up to 20 mg/dL of bilirubin | NSI up to 3000 mg/dL of triglycerides | No Biotin in assay formulation | NA | NA | N/A | N/A | Traceable to an internal standard manufactured using U.S.P. (United States Pharmacopeia) material | 10995437_EN Rev. 02, 2019-08 | IFU Document | 0.235 | 6.969 | 2.087 | 1.457 | 0.107 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995715 |  | an01 | 00630414600284 | Global | N/AAnemia | No | Atellica IM Vitamin B12 (VB12) | VB12 | VB12 | 500 | 5 x 100 | Atellica IM VB12 DTT/REL (SMN 0995718) Atellica IM T3/T4/VB12 ANC (2pk) (SMN 10995682) Atellica IM T3/T4/VB12 ANC (6pk) (SMN 10995683) | Atellica IM VB12 MCM (SMN 10995717) | 14 | 18 | purified hog intrinsic factor | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | 8 hours at RT 2–8°C after 8 hours; ≤ -20°C after 48 hours | 100 μL | No |  | competitive immunoassay using direct chemiluminescent technology | 45–2000 pg/mL | 33–1476 pmol/L | 45 – 20,000 pg/mL | 33 – 14,760 pmol/L | 211 to 911 pg/mL | 156 to 672 pmol/L | NA | NA | NA | Atellica IM CAL C | 10995506 (2-pack) 10995507 (6-pack) | 28 days recon at 2–8°C, 4 hours recon at RT | Lyophilized | Add 5.0 mL of reagent water, let the calibrators stand for 15–20 minutes at RT | 3 for lot calibration, 2 for pack recalibration | 2 | buffered HSA with sodium azide, preservatives and protein stabilizers, | 30 | 18 |  | No |  | Yes | Atellica IM VB12 DIL | 10995716 |  | 45–20000 pg/mL | 33-14760 pmol/L | No | NA | NA | 38 pg/mL | 28 pmol/L | NA | NA | 54 pg/mL | 38 pg/mL | NP | 28 pmol/L | 40 pmol/L | NP | 146 pg/mL (LI heparin) | 108 pmol/L (LI heparin) | 2.5 | 9.2 | 368 pg/mL (serum) | 271 pmol/L (serum) | 2 | 6.4 | 486 pg/mL (serum) | 358 pmol/L (serum) | 2.3 | 5.6 | ADVIA Centaur VB12 y = 1.00x + 7 pg/mL r=0.994 | ADVIA Centaur VB12 y = 1.00x + 5 pmol/L r=0.994 | NSI up to 150 mg/dL of hemoglobin | NSI up to 20 mg/dL of bilirubin | NSI up to 3000 mg/dL of triglycerides | No Biotin in assay formulation | NA | NA | N/A | N/A | Traceable to an internal standard manufactured using U.S.P. (United States Pharmacopeia) material | 10995437_EN Rev. 02 2019-08 | IFU Document | 0.957 | 6.969 | 2.087 | 4.37 | 0.435 | 17.7 | 5.3 | 11.1
Atellica® IM | 10995444 |  | id01 | 00630414597737 | Global | Hepatitis | No | Atellica IM Hepatitis A IgM (aHAVM) | Atellica IM aHAVMaHAVM | aHAVM | 100 | 1 x 100 | Atellica IM Multi-Diluent 2 (SMN 10995644) | Atellica IM aHAVM QC | 46 | 41 | Mouse monoclonal anti-HAV antibody F(ab’)2 fragment | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#EDTA Plasma;# | RT not recommended2–8°C for up to 7 days; ≤ -20°C for up to 180 days; | 20 μL | No |  | IgM-capture immunoassay using a 2-pass format | 0.02-7.00 S/CO | 0.02-7.00 S/CO | NA | NA | NA | NA | Nonreactive <0.8 S/CO | Reactive >1.20 S/CO | Equivocal | Atellica IM aHAVM Cal | Provided in kit | at 2–8°C until expiration, on board at RT 8hrs | Liquid | N/A | 3 for lot calibration, 2 for pack recalibration | 2 | Processed human plasma positive for IgM antibodies to HAV; preservatives | 91 | 28 |  | No |  | No | NA | NA |  | NA | NA | Yes | Atellica IM aHAVM QC | 10995445 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 0.49 S/CO (plasma) | 0.49 S/CO (plasma) | 3.9 | 5.5 | 2.98 S/CO (plasma) | 2.98 S/CO (plasma) | 3.3 | 5 | 5.23 S/CO (plasma) | 5.23 S/CO (plasma) | 4.3 | 5.2 | NA | NA | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 60 mg/dL of conjugated bilirubin, and up to 40 mg/dL of unconjugated bilirubin | ≤ 10% change upto 3000 mg/dL of triglycerides | Insignificant effect on the assay up to 500 ng/mL | NA | NA |  |  | Standardization based upon relative clinical agreement with commercially available anti-HAV IgM assays | 10995271_EN Rev. 03, 2019-07 (OUS) 10995272_EN Rev. 03, 2019-06 (US) | IFU Document | 0.359 | 6.969 | 2.48 | 2.913 | 0.163 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995446 |  | id02 | 00630414597751 | OUS | Hepatitis | No | Atellica IM Hepatitis A Total (aHAVT) | Atellica IM aHAVTaHAVT | aHAVT | 100 | 1 x 100 | Atellica IM APW1 (SMN 10995458) | Atellica IM aHAVT QC | 54 | 41 | Mouse monoclonal anti-HAV (~1.0 μg/mL) labeled with acridinium ester and biotinylated mouse monoclonal anti- HAV Fab fragment | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#EDTA Plasma;# | RT not recommended 2–8°C for up to 7 days; ≤ -20°C for up to 30 days | 20 μL | No |  | competitive immunoassay using direct chemiluminescent technology. | 0.00–100.00 mIU/mL | 0.00–100.00 mIU/mL | NA | NA | NA | NA | Nonreactive < 20 mIU/mL | Reactive > 20 mIU/mL |  | Atellica IM aHAVT Cal | Provided in kit | Unopened at 2–8°C until expiration Opened at 2–8°C 60 days on board at RT 8hrs | Liquid | N/A | 3 for lot calibration, 2 for pack recalibration | 2 | Processed human plasma positive for anti-HAV antibodies | 41 | 28 |  | No |  | No | NA | NA |  | NA | NA | Yes | Atellica IM aHAVT QC | 10995448 | NA | NA | NA | NA | NA | N/A | N/A | NA | N/A | NA | 16.17 mIU/mL (plasma) | 16.17 mIU/mL (plasma) | 4.9 | 6.6 | 36.75 mIU/mL (plasma) | 36.75 mIU/mL (plasma) | 2.5 | 3.5 | 80.47 mIU/mL (plasma) | 80.47 mIU/mL (plasma) | 1.6 | 2.4 | NA | NA | NSI up to 500 mg/dL of hemoglobin | NSI - up to 60 mg/dL of conjugated bilirubin, up to 40 mg/dL of unconjugated bilirubin | NSI up to 3000 mg/dL of triglycerides | Insignificant effect on the assay up to 25 ng/mL | NA | NA |  |  | WHO Second International Standard for Anti-Hepatitis A Immunoglobulin (97/646) | 10995273_EN Rev. 03, 2019-07 | IFU Document | 0.387 | 6.969 | 2.48 | 2.913 | 0.176 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995451 |  | id04 | 00630414597812 | OUS | Hepatitis | No | Atellica IM Hepatitis B e antibody (aHBe) | Atellica IM aHBeaHBe | aHBe | 100 | 1 x 100 | NA | Atellica IM aHBe QC | 52 | 26 | Acridinium-ester-labeled mouse monoclonal anti-HBe ; bovine serum albumin; mouse IgG;Biotinylated mouse monoclonal anti-HBe | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#EDTA Plasma;# | RT for up to 24 hrs; 2–8°C for up to 7 days; ≤ -20°C for longer storage | 100 μL | No |  | competitive/neutralizing immunoassay using a 2‑step, singlewash format. | 0.05–4.50 Index | 0.05–4.50 Index | NA | NA | NA | NA | Nonreactive < 0.80 Index | Reactive > 1.20 Index | Equivocal ≥0.80 Index and <1.20 Index | Atellica IM aHBe Cal | Provided in kit | at 2–8°C until expirationon board at RT 8hrs | Liquid | N/A | 3 for lot calibration, 2 for pack recalibration | 2 | Processed human plasma positive for antibodies to HBe antigen | 51 | 21 |  | No |  | No | No | No |  | No | No | Yes | Atellica IM aHBe QC | 10995452 | NA | NA | NA | NA | NA | N/A | N/A | NA | N/A | NA | 0.46 Index (serum) | 0.46 Index (serum) | 4 | 7.7 | 1.41 Index (serum) | 1.41 Index (serum) | 2.3 | 3.6 | 3.61 Index (serum) | 3.61 Index (serum) | 1.2 | 2.1 | NA | NA | ≤ 20% change up to 500 mg/dL of hemoglobin | ≤ 20% change up to 40 mg/dL of conjugated bilirubin, up to 20 mg/dL of unconjugated bilirubin | ≤ 20% change up to 500 mg/dL of triglycerides, and up to 400 mg/dL of cholesterol | Insignificant effect on the assay up to 1500 ng/mL | NA | NA |  |  | Standarization based upon relative clinical agreement withcommercially available anti‑HBe assays. | 10995276_EN Rev. 03, 2019-07 | IFU Document | 0.251 | 6.969 | 2.48 | 2.913 | 0.114 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995481 |  | on03 | 00630414598086 | Global | Oncology | No | Atellica IM CA 125II (CA 125II) | Atellica IM CA 125II | CA125 | 100 | 1x100 | Atellica IM CA 125II CAL | Atellica IM Multi‑Diluent 1 Atellica IM CA 125II MCM (SMN 10995484) | 28 | 28 | Mouse monoclonal anti-M11 antibody (~0.15 μg/mL) | ;#Serum;# | RT: 8 hours; 2–8°C after 8 hours; ≤ -20°C after 24 hours | 50 μL | No |  | 2-site sandwich immunoassay | 2.0–600.0 U/mL | NA |  |  |  |  | <35 U/mL |  |  | Atellica IM CA 125II CAL, | 10995483 | 7 days recon at 2–8°C 8 hours recon at RT | Lyophilized | 2.0 mL of special reagent water. Let the vials stand for 15–20 minutes at room temperature | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of CA 125 | 91 | 28 |  | No |  | Yes | Atellica IM Multi‑Diluent 1 | 10995637 - 2 pack;10995638 - 6 pack;10995639 - vial |  | 2.0–12000.0 U/mL | NA | No | NA | NA | 1.2 U/mL | NA | NP | NA | 3.0 U/mL | 1.1 U/mL | NP | NA | NA | NA | 15.1 U/mL | NA | 1.5 | 2.8 | 172 U/mL | NA | 1.4 | 2.8 | 519 U/mL | NA | 1.4 | 3.2 | ADVIA Centaur CA 125II y = 0.99x + 1.7 U/mL r=1.00 | NA | ≤ 10% change up to 1000 mg/dL of hemoglobin | ≤10% change up 20 mg/dL conjugated/unconjugated bilirubin | ≤10% change upto 900 mg/dL of triglycerides | No Biotin in assay formulation | 70,000 U/mL | NA |  |  | traceable to an internal standard manufactured using highlypurified material | 10995283_EN Rev. 03, 2019-08 | IFU Document | 0.22 | 6.969 | 2.087 | 1.457 | 0.1 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995482 |  | on03 | 00630414598093 | Global | Oncology | No | Atellica IM CA 125II (CA 125II) | Atellica IM CA 125II | CA125 | 500 | 5 x 100 | Atellica IM CA 125II CAL | Atellica IM Multi‑Diluent 1, Atellica IM CA 125II MCM (SMN 10995484) | 28 | 28 | Mouse monoclonal anti-M11 antibody (~0.15 μg/mL) | ;#Serum;# | RT 8 hours; 2–8°C after 8 hours; ≤ -20°C 24hours | 50 μL | No |  | 2-site sandwich immunoassay | 2.0–600.0 U/mL | NA |  |  |  |  | <35 U/mL |  |  | Atellica IM CA 125II CAL, | 10995484 | 7 days recon at 2–8°C 8 hours recon at RT | Lyophilized | 2.0 mL of special reagent water. Let the vials stand for 15–20 minutes at room temperature | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of CA 125 | 91 | 28 |  | No |  | Yes | Atellica IM Multi‑Diluent 1 | 10995637 - 2 pack;10995638 - 6 pack;10995639 - vial |  | 2.0–12000.0 U/mL | NA | No | NA | NA | 1.2 U/mL | NA | NP | NA | 3.0 U/mL | 1.1 U/mL | NP | NA | NA | NA | 15.1 U/mL | NA | 1.5 | 2.8 | 172 U/mL | NA | 1.4 | 2.8 | 519 U/mL | NA | 1.4 | 3.2 | ADVIA Centaur CA 125II y = 0.99x + 1.7 U/mL r=1.01 | NA | ≤ 10% change up to 1000 mg/dL of hemoglobin | ≤10% change up 20 mg/dL conjugated/unconjugated bilirubin | ≤10% change upto 900 mg/dL of triglycerides | No Biotin in assay formulation | 70,000 U/mL | NA |  |  | traceable to an internal standard manufactured using highly purified material | 10995283_EN Rev. 03, 2019-08 | IFU Document | 0.9 | 6.969 | 2.087 | 4.37 | 0.409 | 17.7 | 5.3 | 11.1
Atellica® IM | 10995485 |  | on04 | 00630414598123 | Global | Oncology | No | Atellica IM CA 15‑3 (CA 15‑3) | Atellica IM CA 15‑3 | CA15_3 | 100 | 1 x 100 | Atellica IM CA 15‑3 CAL | Atellica IM Multi‑Diluent 1 Atellica IM CA 15‑3 MCM (SMN 10995488) | 28 | 28 | Mouse monoclonal anti-DF3 antibody (~ 2.0 μg/mL) | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | 2–8°C after 8 hours;≤ -20°C if the sample is not assayed within 24 hours | 20 μL | No |  | 2-step sandwich immunoassay | 0.5–200.0 U/mL | NA |  |  |  |  | <32.4 U/mL |  |  | Atellica IM CA 15‑3 CAL | 10995487 | 28 days recon at 2–8°C 8 hours recon at RT | Lyophilized | 2.0 mL of special reagent water. Let the vials stand for 15–20 minutes at room temperature | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of CA 15‑3 (human) | 91 | 28 |  | No |  | Yes | Atellica IM Multi‑Diluent 1 - | 10995637 - 2 pack;10995638 - 6 pack;10995639 - vial |  | 0.5–4000.0 U/mL | NA | No | NA | NA | 0.0 U/mL | NA | NP | NA | 1.9 U/mL | 0.0 U/mL | NP | NA | NA | NA | 9.6 U/mL | NA | 3.1 | 6.2 | 65.9 U/mL | NA | 1.6 | 3.7 | 165.8 U/mL | NA | 1.4 | 3.4 | ADVIA Centaur CA 15‑3 y = 1.07x + 0.146 U/mL r=0.99 | NA | ≤ 10% change up to 1000 mg/dL of hemoglobin | ≤10% change up 25 mg/dL conjugated bilirubin | ≤10% change upto 900 mg/dL of triglycerides | No Biotin in assay formulation | NA | NA |  |  | traceable to an internal standard manufactured using highly purified material | 10995284_EN Rev. 03, 2019-08 | IFU Document | 0.224 | 6.969 | 2.087 | 1.457 | 0.102 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995486 |  | on04 | 00630414598130 | Global | Oncology | No | Atellica IM CA 15‑3 (CA 15‑3) | Atellica IM CA 15‑3 | CA15_3 | 500 | 5 x 100 | Atellica IM CA 15‑3 CAL | Atellica IM Multi‑Diluent 1, Atellica IM CA 15‑3 MCM | 28 | 28 | Mouse monoclonal anti-DF3 antibody (~ 2.0 μg/mL) | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | 2–8°C after 8 hours;≤ -20°C if the sample is not assayed within 24 hours | 20 μL | No |  | 2-step sandwich immunoassay | 0.5–200.0 U/mL | NA |  |  |  |  | <32.4 U/mL |  |  | Atellica IM CA 15‑3 CAL | 10995487 | 28 days recon at 2–8°C 8 hours recon at RT | Lyophilized | 2.0 mL of special reagent water. Let the vials stand for 15–20 minutes at room temperature | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of CA 15‑3 (human) | 91 | 28 |  | No |  | Yes | Atellica IM Multi‑Diluent 1 - | 10995637 - 2 pack;10995638 - 6 pack;10995639 - vial |  | 0.5–4000.0 U/mL | NA | No | NA | NA | 0.0 U/mL | NA | NP | NA | 1.9 U/mL | 0.0 U/mL | NP | NA | NA | NA | 9.6 U/mL | NA | 3.1 | 6.2 | 65.9 U/mL | NA | 1.6 | 3.7 | 165.8 U/mL | NA | 1.4 | 3 | ADVIA Centaur CA 15‑3 y = 1.07x + 0.146 U/mL r=0.99 | NA | ≤ 10% change up to 1000 mg/dL of hemoglobin | ≤10% change up 25 mg/dL conjugated bilirubin | ≤10% change upto 900 mg/dL of triglycerides | No Biotin in assay formulation | NA | NA |  |  | traceable to an internal standard manufactured using highly purified material | 10995284_EN Rev. 03, 2019-08 | IFU Document | 0.913 | 6.969 | 2.087 | 4.37 | 0.415 | 17.7 | 5.3 | 11.1
Atellica® IM | 10995662 |  | on01 | 00630414599809 | OUS | Oncology | No | Atellica IM Prostate‑Specific Antigen (PSA) | Atellica IM PSA | PSA | 100 | 1x100 | Atellica IM CAL Q | Multi‑Diluent 2, Atellica IM PSA MCM | 10 | 28 | Goat polyclonal anti‑PSA antibody (~77 ng/mL) | ;#Serum;# | 8hrs at RT, 48 hrs at 2-8°C, after 48 hrs at ≤ -20°C | 35 μL | No |  | 2‑site sandwich immunoassay | 0.01–100.00 ng/mL | 0.01–100.00 μg/L |  |  |  |  | See IFU |  |  | Atellica IM CAL Q | 10995517 (2pk) 10995518 (6pk) | 21 days recon at 2–8°C 8 hours recon at RT | Lyophilized | 2.0 mL of special reagent water. Let the vials stand for 15–20 minutes at room temperature | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of PSA | 29 | 28 |  | No |  | Yes | Atellica IM Multi‑Diluent 2 | 10995644 |  | 0.01–10000.00 ng/mL | 0.01–10000.00 μg/L | No | NA | NA | 0.01 ng/mL | 0.01 μg/L | NP | NP | 0.02 ng/mL | 0.01 ng/mL | 0.02 ng/mL | 0.01 μg/L | 0.02 μg/L | 0.02 μg/L | 0.06 ng/mL | 0.06 μg/L | 6.2 | 8.5 | 19.63 ng/mL | 19.63 μg/L | 1.8 | 2.8 | 78.67 ng/mL | 78.67 μg/L | 2.2 | 4 | ADVIA Centaur PSA y = 0.93x + 0.238 ng/mL r=.996 | ADVIA Centaur PSA y = 0.93x + 0.238 μg/L r=.996 | ≤ 5% change up to 500 mg/dL of hemoglobin | ≤ 5% change up to 40 mg/dL of bilirubin | ≤ 5% change up to 1000 mg/dL of triglycerides | No Biotin in assay formulation | 2,500 ng/mL | 2,500 μg/L |  |  | (WHO) International Standard (96/670) | 10995416_EN Rev. 03, 2019-09 | IFU Document | 0.222 | 6.969 | 2.087 | 1.457 | 0.101 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995663 |  | on01 | 00630414599816 | OUS | Oncology | No | Atellica IM Prostate‑Specific Antigen (PSA) | Atellica IM PSA | PSA | 500 | 5 x 100 | Atellica IM CAL Q | Multi‑Diluent 2, Atellica IM PSA MCM | 10 | 28 | Goat polyclonal anti‑PSA antibody (~77 ng/mL) | ;#Serum;# | 8hrs at RT, 48 hrs at 2-8°C, after 48 hrs at ≤ -20°C | 35 μL | No |  | 2‑site sandwich immunoassay | 0.01–100.00 ng/mL | 0.01–100.00 μg/L |  |  |  |  | See IFU |  |  | Atellica IM CAL Q | 10995517 (2pk) 10995518 (6pk) | 21 days recon at 2–8°C 8 hours recon at RT | Lyophilized | 2.0 mL of special reagent water. Let the vials stand for 15–20 minutes at room temperature | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of PSA | 29 | 28 |  | No |  | Yes | Atellica IM Multi‑Diluent 3 | 10995645 |  | 0.01–10000.00 ng/mL | 0.01–10000.00 μg/L | No | NA | NA | 0.01 ng/mL | 0.01 μg/L | NP | NP | 0.02 ng/mL | 0.01 ng/mL | 0.02 ng/mL | 0.01 μg/L | 0.02 μg/L | 0.02 μg/L | 0.06 ng/mL | 0.06 μg/L | 6.2 | 8.5 | 19.63 ng/mL | 19.63 μg/L | 1.8 | 2.8 | 78.67 ng/mL | 78.67 μg/L | 2.2 | 4 | ADVIA Centaur PSA y = 0.93x + 0.238 ng/mL r=.996 | ADVIA Centaur PSA y = 0.93x + 0.238 μg/L r=.996 | ≤ 5% change up to 500 mg/dL of hemoglobin | ≤ 5% change up to 40 mg/dL of bilirubin | ≤ 5% change up to 1000 mg/dL of triglycerides | No Biotin in assay formulation | 2,500 ng/mL | 2,500 μg/L |  |  | (WHO) International Standard (96/670) | 10995416_EN Rev. 03, 2019-09 | IFU Document | 0.9 | 6.969 | 2.087 | 4.37 | 0.409 | 17.7 | 5.3 | 11.1
Atellica® IM | 10995527 |  | id07 | 00630414598543 | OUS | HIV | No | Atellica IM HIV Ag/Ab Combo (CHIV) | Atellica IM CHIVCHIV | CHIV | 100 | 1x100 | Atellica IM PW3 | Atellica IM CHIV QC | 39 | 42 | recombinant HIV antigens/Streptavidin-coated paramagnetic microparticles preformed with biotinylated HIV antigens, recombinant HIV ag labelled with AE | ;#Serum;#EDTA Plasma;# | RT not recommended; at 2–8°C up to 14 days; at ≤ -20°C up to 8 months | 100 μL | No |  | antigen-bridging microparticle chemiluminometric immunoassay | 0.05–12.00 Index | 0.05–12.00 Index | N/A | N/A | see IFU | see IFU | see IFU | see IFU | see IFU | IM CHIV CAL | Provided in Kit | at 2–8°C until expiration 8 hours at RT | Liquid | NA | 3 for lot calibration, 2 for pack recalibration | 2 | Processed human plasma negative for antibodies to HIV spiked with antibodies to HIV‑1 | 24 | 21 |  |  |  | No | NA | N/A |  | N/A | N/A | Yes | Atellica IM CHIV QC | 11200769 | ≤ 2.0 IU/mL | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 1.03 Index | 1.03 Index | 2.2 | 3 | 2.81 Index | 2.81 Index | 1.4 | 2.2 | 5.54 Index | 5.54 Index | 1.6 | 2.8 | N/A | N/A | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 40 mg/dL of conjugated bilirubin, and up to 40 mg/dL of unconjugated bilirubin | ≤10% change upto 1000 mg/dL of triglycerides | Insignificant effect on the assay&nbsp; up to 3500 ng/mL | N/A | N/A | N/A | N/A | Standardization is based upon relative clinical agreement with commercially available anti-HIV combo assays | 10995315_EN Rev. 04, 2019-07 | IFU Document | 0.264 | 6.969 | 2.48 | 2.913 | 0.12 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995597 |  | id08 | 00630414599205 | OUS | Hepatitis | No | Atellica IM Anti‑Hepatitis B core Total (HBcT) | Atellica IM HBcTHBcT | HBcT | 200 | 1 x 200 | Atellica IM APW1; Atellica IM PW3 | Atellica IM HBcT QC | 46 | 28 | Recombinant hepatitis B core antigen | ;#Serum;#EDTA Plasma;# | RT not recommended; 2–8°C up to 7 days; ≤ -20°C for longer storage | 50 μL | No |  | 2‑wash antigen sandwich immunoassay | NA | NA | NA | NA | NA | NA | Nonreactive < 0.50 Index | Reactive ≥ 0.50 Index |  | Atellica IM HBcT CAL | Provided in Kit | at 2–8°C until expiration; RT at 8hrs on board | Liquid | N/A | 3 for lot calibration, 2 for pack recalibration | 2 | Processed human plasma positive for HBc antibodies | 33 | 14 |  |  |  | No | N/A | N/A |  | N/A | N/A | Yes | Atellica IM HBcT QC | 10995598 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 0.38 Index | 0.38 Index | 4.8 | 6.9 | 2.07 Index | 2.07 Index | 4.8 | 7.7 | 5.27 Index | 5.27 Index | 7.2 | 9.5 | NA | NA | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 60 mg/dL of conjugated bilirubin; 40 mg/dL of unconjugated bilirubin | ≤ 10% change up to 1000 mg/dL of intralipids | Insignificant effect on the assay up to 1500 ng/mL | N/A | N/A |  |  | based upon the relative clinical agreement with commercially available anti-HBc Total assays | 10995353_EN Rev. 03, 2019-07 | IFU Document | 0.394 | 6.969 | 2.48 | 2.913 | 0.179 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995449 |  | id03 | 00630414597799 | OUS | Hepatitis | No | Atellica IM Hepatitis B core IgM (aHBcM) | Atellica IM aHBcMaHBcM | aHBcM | 100 | 1 x 100 | N/A | Atellica IM aHBcM QC | 46 | 41 | monoclonal mouse anti-HBc | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#EDTA Plasma;# | RT not recommended 2–8°C up to 7 days 7 days at ≤ -20°C | 15 μL | No |  | IgM capture immunoassay using a 2-step format | 0.05–9.00 Index | 0.05–9.00 Index | NA | NA | NA | NA | Negative <0.80 Index | Equivocal 0.80-0.99 Index | Positive ≥ 1.0 Index | Atellica IM aHBcM CAL | Provided in Kit | at 2–8°C until expiration RT at 8hrs on board | Liquid | N/A | 3 for lot calibration, 2 for pack recalibration | 2 | processed human plasma positive for IgM antibodies to HBc antigen with preservatives | 92 | 27 |  |  |  | No | N/A | N/A |  | N/A | N/A | Yes | Atellica IM aHBcM QC | 10995450 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 0.39 Index | 0.39 Index | 1.01 | 1.55 | 2.11 Index | 2.11 Index | 2.18 | 2.87 | 7,86 | 7,86 | 5.09 | 5.99 | NA | NA | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 60 mg/dL of conjugated bilirubin 40 mg/dL of unconjugated bilirubin | ≤ 10% change up to 1000 mg/dL of triglycerides | Insignificant effect on the assay up to 150 ng/mL | N/A | N/A |  |  | assay standardization is based upon relative clinical agreement with commercially available anti-HBc IgM assays | 10995274_EN Rev. 04, 2019-07 | IFU Document | 0.374 | 6.969 | 2.48 | 2.913 | 0.17 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995604 |  | id05 | 00630414599274 | OUS | Hepatitis | No | Atellica IM Hepatitis B surface Antigen II (HBsII) | Atellica IM HBsIIHBsII | HBsII | 200 | 1 x 200 | Atellica IM PW3 (SMN 10995666) | Atellica IM HBsII QC | 26 | 60 |  | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#EDTA Plasma;#Sodium Citrate Plasma;# | RT not recommended; 2–8°C up to 7 days ; ≤ -20°C for longer storage | 100 μL | No |  |  | 0.1–1000.0 Index | 0.1–1000.0 Index | NA | NA | NA | NA | Non Reactive <1.0 Index | Reactive ≥ 1.00 Index |  | Atellica IM HBsII CAL | Provided in Kit | at 2–8°C until expiration RT at 8hrs on board | Liquid | N/A | 3 for lot calibration, 2 for pack recalibration | 2 |  | 48 | 21 |  |  |  | No | N/A | N/A |  | N/A | N/A | Yes | Atellica IM HBsII QC | 10995605 | 0.018 IU/mL | NA | NA | NA | NA | NA | NA | NA | NA | NA | 0.9 Index | 0.9 Index | 7.2 | 7.5 | 50.2 Index | 50.2 Index | 1.9 | 2.8 | 811.6 Index | 811.6 Index | 1.5 | 2.2 | N/A | NA | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 40 mg/dL of conjugated bilirubin, and up to 40 mg/dL of unconjugated bilirubin | ≤ 10% change up to 1000 mg/dL of triglycerides | ≤ 10% change up to 10 ng/mL | 3.36 mg/mL | N/A |  |  | based upon the relative clinical agreement with commercially available HBsAg assays | 10995358_EN Rev. 03, 2019-07 | IFU Document | 0.394 | 6.969 | 2.48 | 2.913 | 0.179 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995603 |  | id10 | 00630414599267 | OUS | Hepatitis | No | Atellica IM Hepatitis B surface Antigen II Confirmatory (HBsII Conf) | Atellica IM HBsII ConfConf (HBsII) | Conf (HBsII) | 98 | 1 x 98 | NA | Atellica IM HBsII QC | 53 | 41 | human plasma positive for HBs antibodies | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#EDTA Plasma;#Citrate Plasma;# | RT: 24 hrs 2-8C: 7 days, -20C:for longer storage | 200 μL | No |  |  | NA | NA | NA | NA | NA | NA | Not confirmed | Confirmed |  | Atellica IM HBsII CAL | Provided in Kit | at 2–8°C until expiration, RT at 8hrs on board | Liquid | N/A | 3 for lot calibration, 2 for pack recalibration | 2 |  | 43 | 21 |  |  |  | No | N/A | N/A |  | N/A | N/A | Yes | Atellica IM HBsII QC | 10995605 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA |  |  | NA | NA |  |  | NA | NA |  |  | NA | NA | NA | NA | NA | NA | N/A | N/A |  |  | N/A | 10995268_EN Rev. 03, 2019-10 | IFU Document | 0.101 | 4.37 | 2.48 | 2.52 | 0.046 | 11.1 | 6.3 | 6.4
Atellica® IM | 10995569 |  | an03 | 00630414598932 | OUS | N/AAnemia | No | Atellica IM Ferritin (Fer) | Atellica IM FerFer | Fer | 90 | 1 x 90 | N/A | Atellica IM Fer MCM (SMN 10995570) | 14 | 28 | Goat polyclonal and mouse monoclonal anti-ferritin antibody. | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | up to 8 hrs at RT; 48 hrs at 2–8°C. Freeze samples at ≤ -20°C if the sample is not assayed within 48 hours. | 10 μL | No |  | 2-site sandwich immunoassay using direct chemiluminometric | 0.5–1650.0 ng/mL | 1.1–3630.0 pmol/L | 0.5 – 16,500 ng/mL | 1.1 - 36,300 pmol/L | Normal Males 22–322 ng/mL; Normal Females 10–291 ng/mL | Normal Males 48–708 pmol/L; Normal Females 22–640 pmol/L |  |  |  | Atellica IM Calibrator C (CAL C) | 10995506 (2‑pack) 10995507 (6‑pack) | 4 hours at RT; 28 days at 2–8°C | Lyophilized | Reconstitute with 5.0 mL of special reagent water. | 3 for lot calibration, 2 for pack recalibration | 2 | Vitamin B12 and ferrtin; buffered HSA; protein stabilizers; sodium azide | 50 | 28 |  |  |  | Yes | Atellica IM Multi‑Diluent 1 | 10995637 (2‑pack) 10995638 (6‑pack) 10995639 (vial) |  | 0.5 – 16,500 ng/mL | 1.1 - 36,300 pmol/L | No | NA | NA | 0.0 ng/mL | 0.0 pmol/L | NA | NA | 0.7 ng/mL | 0.3 ng/mL | 0.9 ng/mL | 0.7 pmol/L | 1.5 pmol/L | 2.0 pmol/L | 8.2 ng/mL | 18.0 pmol/L | 1.8 | 5.4 | 487.1 ng/mL | 1071.6 pmol/L | 1.6 | 4.2 | 1453.6 ng/mL | 3197.9 pmol/L | 3.4 | 6.3 | ADVIA Centaur FER y = 1.03x - 0.6 ng/mL r= 1.0 n=106 | ADVIA Centaur FER y = 1.03x - 1.3 pmol/L r=1.0 | ≤ 10% change up to 900 mg/dL (0.56 mmol/L) of hemoglobin | ≤ 10% change up to 60 mg/dL conjugated bilirubin (1021 μmol/L) 60 mg/dL unconjugated bilirubin (1021 μmol/L) | ≤ 10% change up to 2000 mg/dL (22.6 mmol/L) | No Biotin in assay formulation | 80,000 ng/mL | 176,000 pmol/L | 80,000 ng/mL | 176,000 pmol/L | Traceable to World Health Organization 2nd International Standard (WHO 80/578) | 10995343_EN Rev. 03, 2019-07 | IFU Document | 0.229 | 6.969 | 2.087 | 1.457 | 0.104 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995568 |  | an03 | 00630414598925 | OUS | Anemia | No | Atellica IM Ferritin (Fer) | Atellica IM FerFer | Fer | 450 | 5 x 90 | N/A | Atellica IM Fer MCM (SMN 10995570) | 14 | 28 | Goat polyclonal and mouse monoclonal anti-ferritin antibody. | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | up to 8 hrs at RT; 48 hrs at 2–8°C. Freeze samples at ≤ -20°C if the sample is not assayed within 48 hours. | 10 μL | No |  | 2-site sandwich immunoassay using direct chemiluminometric | 0.5–1650.0 ng/mL | 1.1–3630.0 pmol/L | 0.5 – 16,500 ng/mL | 1.1 - 36,300 pmol/L | Normal Males 22–322 ng/mL; Normal Females 10–291 ng/mL | Normal Males 48–708 pmol/L; Normal Females 22–640 pmol/L |  |  |  | Atellica IM Calibrator C (CAL C) | 10995506 (2‑pack) 10995507 (6‑pack) | 4 hours at RT; 28 days at 2–8°C | Lyophilized | Reconstitute with 5.0 mL of special reagent water | 3 for lot calibration, 2 for pack recalibration | 2 | Vitamin B12 and ferrtin; buffered HSA; protein stabilizers; sodium azide | 50 | 28 |  |  |  | Yes | Atellica IM Multi‑Diluent 1 | 10995637 (2‑pack) 10995638 (6‑pack) 10995639 (vial) |  | 0.5 – 16,500 ng/mL | 1.1 - 36,300 pmol/L | No | NA | NA | 0.0 ng/mL | 0.0 pmol/L | NA | NA | 0.7 ng/mL | 0.3 ng/mL | 0.9 ng/mL | 0.7 pmol/L | 1.5 pmol/L | 2.0 pmol/L | 8.2 ng/mL | 18.0 pmol/L | 1.8 | 5.4 | 487.1 ng/mL | 1071.6 pmol/L | 1.6 | 4.2 | 1453.6 ng/mL | 3197.9 pmol/L | 3.4 | 6.3 | ADVIA Centaur FER y = 1.03x - 0.6 ng/mL r= 1.0 n=106 | ADVIA Centaur FER y = 1.03x - 1.3 pmol/L r=1.0 | ≤ 10% change up to 900 mg/dL (0.56 mmol/L) of hemoglobin | ≤ 10% change up to 60 mg/dL conjugated bilirubin (1021 μmol/L) 60 mg/dL unconjugated bilirubin (1021 μmol/L) | ≤ 10% change up to 2000 mg/dL (22.6 mmol/L) | No Biotin in assay formulation | 80,000 ng/mL | 176,000 pmol/L | 80,000 ng/mL | 176,000 pmol/L | Traceable to World Health Organization 2nd International Standard (WHO 80/578) | 10995343_EN Rev. 03, 2019-07 | IFU Document | 0.834 | 6.969 | 2.087 | 4.37 | 0.379 | 17.7 | 5.3 | 11.1
Atellica® IM | 10995490 |  | on05 | 00630414598178 | OUS | Oncology | No | Atellica IM CA 19‑9™ (CA 19‑9™) | Atellica IM CA 19‑9 | CA19_9 | 50 | 1 X 50 | N/A | Atellica IM CA 19‑9 DIL (SMN 10995491)Atellica IM CA 19‑9 MCM (SMN 10995492) | 28 | 28 | Mouse monoclonal anti-CA 19‑9 antibody | ;#Serum;# | up to 8 hrs at RT, 2–8°C up to 48 hrs, after 48 hrs at ≤ -20°C. Freeze once | 75 μL | No |  | 2-step sandwich immunoassay using direct chemiluminescent technology | 1.20–700.00 U/mL | NA |  |  |  |  | See IFU |  |  | Atellica IM CA 19‑9 CAL | Provided in Kit | at 2–8°C 14 days RT at 8hrs | Lyophilized | Add 2.0 mL of special reagent water, let the vials stand for 15–20 minutes at room temperature | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of CA 19‑9 (human) | 38 | 27 |  |  |  | Yes | Atellica IM CA 19‑9 DIL | 10995491 |  | 1.20–140,000.00 U/mL | N/A | No | NA | NA | 0.00 U/mL | NA | NA | NA | 3.99 U/mL | 0.00 U/mL | NA | NA | NA | NA | Control 1 25.56 U/mL | NA | 4 | 6 | Control 2 125.24 U/mL | NA | 3.9 | 6.2 | Control 3 301.12 U/mL | NA | 4.4 | 5.7 | ADVIA Centaur CA 19‑9 (x) y = 1.03x - 2.03 U/mL r=0.99 | NA | ≤ 4.0% change up to 1200 mg/dL of hemoglobin | ≤ 4.0% change up to 50 mg/dL of bilirubin | ≤ 4.0% change up to 3000 mg/dL of triglycerides | No Biotin in assay formulation | N/A | N/A |  |  | traceable to internal standard manufactured using highly purified material | 10995285_EN Rev. 03, 2019-07 | IFU Document | 0.317 | 6.969 | 2.48 | 2.913 | 0.144 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995489 |  | on05 | 00630414598161 | OUS | Oncology | No | Atellica IM CA 19‑9™ (CA 19‑9™) | Atellica IM CA 19‑9 | CA19_9 | 250 | 1 X 250 | N/A | Atellica IM CA 19‑9 DIL (SMN 10995491) Atellica IM CA 19‑9 MCM (SMN 10995492) | 28 | 28 | Mouse monoclonal anti-CA 19‑9 antibody | ;#Serum;# | up to 8 hrs at RT, 2–8°C up to 48 hrs, after 48 hrs at ≤ -20°C. Freeze once | 75 μL | No |  | 2-step sandwich immunoassay using direct chemiluminescent technology | 1.20–700.00 U/mL | NA |  |  |  |  | See IFU |  |  | Atellica IM CA 19‑9 CAL | Provided in Kit | at 2–8°C 14 days RT at 8hrs | Lyophilized | Add 2.0 mL of special reagent water, let the vials stand for 15–20 minutes at room temperature | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of CA 19‑9 (human) | 38 | 27 |  |  |  | Yes | Atellica IM CA 19‑9 DIL | 10995491 |  | 1.20–140,000.00 U/mL | N/A | No | NA | NA | 0.00 U/mL | NA | NA | NA | 3.99 U/mL | 0.00 U/mL | NA | NA | NA | NA | Control 1 25.56 U/mL | NA | 4 | 6 | Control 2 125.24 U/mL | NA | 3.9 | 6.2 | Control 3 301.12 U/mL | NA | 4.4 | 5.7 | ADVIA Centaur CA 19‑9 y = 1.03x - 2.03 U/mL r=0.99 | NA | ≤ 4.0% change up to 1200 mg/dL of hemoglobin | ≤ 4.0% change up to 50 mg/dL of bilirubin | ≤ 4.0% change up to 3000 mg/dL of triglycerides | No Biotin in assay formulation | N/A | N/A |  |  | traceable to internal standard manufactured using highly purified material | 10995285_EN Rev. 03, 2019-07 | IFU Document | 0.999 | 6.969 | 2.126 | 6.811 | 0.454 | 17.7 | 5.4 | 17.3
Atellica® IM | 10995585 |  | ty01 | 00630414599090 | Global | N/AThyroid | No | Atellica IM Free Triiodothyronine (FT3) | Atellica IM FT3FT3 | FT3 | 60 | 1 x 60 | Atellica IM CAL A | Atellica IM FT3 MCM (SMN 10995586) | 14 | 28 | acridinium -ester labeled mouse monoclonal anti-T3 antibody | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | up to 8hrs at RT, 8 hrs at 2–8°C, 48hrs at ≤ -20°C | 35 | No |  | Competitive immunoassay using direct chemilumimescent technology | 0.20–20.00 pg/mL | 0.31–30.80 pmol/L | NA | NA | 2.3 - 4.2 pg/mL | 3.5 - 6.5 pmol/L | NA | NA | NA | Atellica IM CAL A | ;2- pack (SMN 10995500) 6- pack (SMN 10995501) | 28 days recon at 2–8°C; 4 hrs recon at RT | Lyophilized | Reconstitute with 5 mL of distilled or deionized water . Let stand 15 - 20 minutes . Gently swirl or mix by inversion | 3 for lot calibration, 2 for pack recalibration | 2 | NA | 82 | 28 |  |  |  | No | NA | N/A |  | N/A | N/A | No | Use commercially available controls | NA | 0.06 pg/mL | 0.09 pmol/L | NA | NA | 0.02 pg/mL | 0.00 pg/mL | NA | 0.00 pmol/L | 0.03 pmol/L | NA | 2.87 pg/mL | 4.42 pmol/L | 0.94 | 1.78 | 4.96 pg/mL | 7.64 pmol/L | 0.67 | 1.07 | 8.71 pg/mL | 13.41 pmol/L | 0.78 | 1.13 | ADVIA Centaur FT3 y = 0.99x + 0.007 pg/mL; r=0.99 | ADVIA Centaur FT3 y = 0.99x + 0.011 pmol/L r=0.99 | ≤ 10% change up to 500 mg/dL (0.3 mmol/L) of hemoglobin | ≤ 10% change up to 20 mg/dL (341 μmol/L) conjugated bilirubin 20 mg/dL (341 μmol/L) unconjugated bilirubin | ≤ 10% change up to 1000 mg/dL (11.3 mmol/L) intralipids | No Biotin in assay formulation | N/A | N/A | N/A | N/A | Internal standard manufactured using U.S.P.(United States Pharmacopeia) material | 10995347_EN Rev. 02, 2020-02 | IFU Document | 0.205 | 6.969 | 2.087 | 1.457 | 0.093 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995584 |  | ty01 | 00630414599083 | Global | N/AThyroid | No | Atellica IM Free Triiodothyronine (FT3) | Atellica IM FT3FT3 | FT3 | 300 | 5 x 60 | Atellica IM CAL A | Atellica IM FT3 MCM (SMN 10995586) | 14 | 28 | acridinium-ester-labeled mouse monoclonal anti-T3 antibody | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | up to 8hrs at RT, 8 hrs at 2–8°C, 48hrs at ≤ -20°C | 35 | No |  | Competitive immunoassay using direct chemiluminescent technology | 0.20–20.00 pg/mL | 0.31–30.80 pmol/L | NA | NA | 2.3 - 4.2 pg/mL | 3.5 - 6.5 pmol/L | NA | NA | NA | Atellica IM CAL A | ;2- pack (SMN 10995500) 6- pack (SMN 10995501) | 28 days recon at 2-8°C; 4 hrs recon at RT | Lyophilized | Reconstitute with 5 mL of distilled or deionized water . Let stand 15 - 20 minutes . Gently swirl or mix by inversion | 3 for lot calibration, 2 for pack recalibration | 2 | NA | 82 | 28 |  |  |  | No | NA | N/A |  | N/A | N/A | No | Use commercially available controls | NA | 0.06 pg/mL | 0.09 pmol/L | NA | NA | 0.02 pg/mL | 0.00 pg/mL | NA | 0.00 pmol/L | 0.03 pmol/L | NA | 2.87 pg/mL | 4.42 pmol/L | 0.94 | 1.78 | 4.96 pg/mL | 7.64 pmol/L | 0.67 | 1.07 | 8.71 pg/mL | 13.41 pmol/L | 0.78 | 1.13 | ADVIA Centaur FT3 y = 0.99x + 0.007 pg/mL; r=0.99 | ADVIA Centaur FT3 y = 0.99x + 0.011 pmol/L r=0.99 | ≤ 10% change up to 500 mg/dL (0.3 mmol/L) of hemoglobin | ≤ 10% change up to 20 mg/dL (341 μmol/L) conjugated bilirubin 20 mg/dL (341 μmol/L) unconjugated bilirubin | ≤ 10% change up to 1000 mg/dL (11.3 mmol/L) intralipids | No Biotin in assay formulation | N/A | N/A | N/A | N/A | Internal standard manufactured using U.S.P.(United States Pharmacopeia) material | 10995347_EN Rev. 02, 2020-02 | IFU Document | 0.823 | 6.969 | 2.087 | 4.37 | 0.374 | 17.7 | 5.3 | 11.1
Atellica® IM | 10995679 |  | ty08 | 00630414599922 | Global | N/AThyroid | No | Atellica IM Total Triiodothyronine (T3) | Atellica IM T3T3 | T3 | 120 | 1 x 120 | Atellica IM CAL A | Atellica IM T3 DIL (SMN 10995680) Atellica IM T3 MCM (SMN 10995681) | 14 | 28 | acridinium-ester-labeled mouse monoclonal anti-T3 antibody | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | up to 8hrs at RT, 8 hrs at 2–8°C, 48hrs at ≤ -20°C | 30 | No |  | Competitive immunoassay using direct chemiluminescent technology | 0.10–8.00 ng/mL | 0.15–12.30 nmol/L | NA | NA | 0.6 - 1.81 ng/mL | 0.92 - 2.79 nmol/L | NA | NA | NA | Atellica IM CAL A | 2-pack (SMN 10995500) 6-pack (SMN 10995501) | 28 days recon at 2–8°C; 4 hrs recon at RT | Lyophilized | Reconstitute with 5 mL of distilled or deionized water . Let stand 15 - 20 minutes . Gently swirl or mix by inversion | 3 for lot calibration, 2 for pack recalibration | 2 | human plasma | 56 | 21 |  |  |  | Manual | Atellica IM T3 DIL | 10995680 |  | N/A | N/A | No | Use commercially available controls | NA | 0.10 ng/mL | 0.15 nmol/L | NA | NA | 0.38 ng/mL | 0.00 ng/mL | NA | 0.00 nmol/L | 0.59 nmol/L | NA | 0.74 ng/mL | 1.14 nmol/L | 2.8 | 5.2 | 1.23 ng/mL | 1.89 nmol/L | 2.5 | 6.7 | 7.93 ng/mL | 12.21 nmol/L | 1.8 | 9.3 | ADVIA Centaur T3; y = 1.08x - 0.02 ng/mL r=0.99 | ADVIA Centaur T3; y = 1.08x - 0.02 ng/mL r=0.99 | ≤ 10% change up to 250 mg/dL (0.15 mmol/L) hemoglobin | ≤ 10% change up to 20 mg/dL (341 μmol/L) conjugated bilirubin 20 mg/dL (341 μmol/L) unconjugated bilirubin | ≤ 10% change up to1000 mg/dL (11.3 mmol/L) intralipids | No Biotin in assay formulation | N/A | N/A | N/A | N/A | Internal standard manufactured using U.S.P.(United States Pharmacopeia) material | 10995424_EN Rev. 03, 2020-06 | IFU Document | 0.216 | 6.969 | 2.087 | 1.457 | 0.098 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995678 |  | ty08 | 00630414599915 | Global | N/AThyroid | No | Atellica IM Total Triiodothyronine (T3) | Atellica IM T3T3 | T3 | 600 | 5 x 120 | Atellica IM CAL A | Atellica IM T3 DIL (SMN 10995680) Atellica IM T3 MCM (SMN 10995681) | 14 | 28 | acridinium-ester-labeled mouse monoclonal anti-T3 antibody | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | up tp 8hrs at RT 8 hrs at 2–8°C 48hrs at ≤ -20°C | 30 | No |  | Competitive immunoassay using direct chemiluminescent technology | 0.10–8.00 ng/mL | 0.15–12.30 nmol/L | NA | NA | 0.6 - 1.81 ng/mL | 0.92 - 2.79 nmol/L | NA | NA | NA | Atellica IM CAL A | 2-pack (SMN 10995500) 6-pack (SMN 10995501) | 28 days recon at 2–8°C; 4 hrs recon at RT | Lyophilized | Reconstitute with 5 mL of distilled or deionized water . Let stand 15 - 20 minutes . Gently swirl or mix by inversion | 3 for lot calibration, 2 for pack recalibration | 2 | human plasma | 56 | 21 |  |  |  | Manual | Atellica IM T3 DIL | 10995680 |  | N/A | N/A | No | Use commercially available controls | NA | 0.10 ng/mL | 0.15 nmol/L | NA | NA | 0.38 ng/mL | 0.00 ng/mL | NA | 0.00 nmol/L | 0.59 nmol/L | NA | 0.74 ng/mL | 1.14 nmol/L | 2.8 | 5.2 | 1.23 ng/mL | 1.89 nmol/L | 2.5 | 6.7 | 7.93 ng/mL | 12.21 nmol/L | 1.8 | 9.3 | ADVIA Centaur T3; y = 1.08x - 0.02 ng/mL r=0.99 | ADVIA Centaur T3; y = 1.08x - 0.02 ng/mL r=0.99 | ≤ 10% change up to 250 mg/dL (0.15 mmol/L) hemoglobin | ≤ 10% change up to 20 mg/dL (341 μmol/L) conjugated bilirubin 20 mg/dL (341 μmol/L) unconjugated bilirubin | ≤ 10% change up to1000 mg/dL (11.3 mmol/L) intralipids | No Biotin in assay formulation | N/A | N/A | N/A | N/A | Internal standard manufactured using U.S.P.(United States Pharmacopeia) material | 10995424_EN Rev. 03, 2020-06 | IFU Document | 0.865 | 6.969 | 2.087 | 4.37 | 0.393 | 17.7 | 5.3 | 11.1
Atellica® IM | 10995600 |  | id09 | 00630414599236 | OUS | Hepatitis | No | Atellica IM Hepatitis B e Antigen (HBeAg) | Atellica IM HBeAgHBeAg | HBeAg | 50 | 1 x 50 | Atellica IM PW3 (SMN 10995666) | Atellica IM HBeAg QC Atellica IM Multi‑Diluent 10 (SMN10995640) | 39 | 60 |  | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#EDTA Plasma;# | RT not recommended; at 2–8°C up to 3 days (primary tubes), up to 7 days (separated samples); ≤ -20°C for longer storage | 100 μL | No |  |  | 0.05–1000.00 Index | 0.05–1000.00 Index | NA | NA | NA | NA | See IFU |  |  | Atellica IM HBeAg CAL | Provided in Kit | at 2–8°C until expiration 8 hours at RT | Liquid | N/A | 3 for lot calibration, 2 for pack recalibration | 2 |  | 74 | 30 |  |  |  | Yes | Atellica IM Multi‑Diluent 10 | 10995640 |  | na | N/A | Yes | Atellica IM HBeAg QC | 10995601 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 0.91 Index | 0.91 Index | 6.9 | 8.3 | 219.56 Index | 219.56 Index | 1.3 | 1.7 | 770.11 Index | 770.11 Index | 1.1 | 1.8 | NA | NA | NA | NA | NA | Insignificant effect on the assay up to 1500 ng/mL | N/A | N/A |  |  | standardization is based upon relative clinical agreement with commercially available HBeAg assays | 10995355_EN Rev. 03, 2019-07 | IFU Document | 0.264 | 6.969 | 2.48 | 2.913 | 0.12 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995610 |  | id11 | 00630414599335 | OUS | HIV | No | Atellica IM HIV 1/O/2 Enhanced (EHIV) | Atellica IM EHIVEHIV | EHIV | 200 | 1x200 | Atellica IM PW3 | Atellica IM EHIV QC | 39 | 28 | recombinant HIV antigens/Streptavidin-coated paramagnetic microparticles preformed with biotinylated HIV antigens, recombinant HIV ag labelled with AE | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#EDTA Plasma;# | at 2–8°C up to 7 days; ≤ -20°C for longer storage | 50 μL | No |  | antigen-bridging microparticle chemiluminometric immunoassay | 0.05–50.00 Index | 0.05–50.00 Index | N/A | N/A | see IFU | see IFU | see IFU | see IFU | see IFU | Atellica IM EHIV CAL | Provided in Kit | at 2–8°C until expiration RT at 8hrs on board | Liquid | NA | 3 for lot calibration, 2 for pack recalibration | 2 | Processed human plasma negative for antibodies to HIV spiked with antibodies to HIV‑1 | 29 | 14 |  |  |  | No | NA | N/A |  | N/A | N/A | Yes | Atellica IM EHIV QC | 10995612 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 0.55 Index | 0.55 Index | 5.1 | 10.5 | 6.37 | 6.37 | 2.6 | 2.8 | 25.66 | 25.66 | 4.3 | 5.4 | NA | NA | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 30 mg/dL of conjugated bilirubin, 20 mg/dL of unconjugated bilirubin | ≤ 10% change up to 3000 mg/dL of triglycerides | Insignificant effect on the assay up to 3500 ng/mL | N/A | N/A | N/A | N/A | standardization is based upon relative clinical agreement with commercially available anti-HIV assays | 110995364_EN Rev. 04, 2020-11 | IFU Document | 0.323 | 6.969 | 2.48 | 2.913 | 0.147 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995656 |  | re06 | 00630414599748 | Global | Reproductive Endocrinology | No | Atellica IM Prolactin (PRL) | Atellica IM PRL | PRL | 50 | 1 x 50 | Atellica IM CAL B | Atellica IM Multi‑Diluent 1 Atellica IM PRL MCM (SMN 10995657) | 14 | 28 | goat polyclonal anti-prolactin antibody | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | up to 8hrs at RT, 8 hrs at 2–8°C, 48 hrs at ≤ -20°C | 25 μL | No |  | 2-site sandwich immunoassay | 0.30–200.00 ng/mL | 6.36–4240.00 μIU/mL |  |  |  |  | See IFU |  |  | Atellica IM CAL B | 10995503 (2‑pack) 10995504 (6‑pack) | at 2–8°C 28 days | Lyophilized | Add 5.0 mL of special reagent water into each vial Let the vials stand for 15–20 min at RT | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of digoxin,follicle-stimulating hormone (FSH), luteinizinghormone (LH), prolactin, human chorionic gonadotropin (hCG), and TSH; equine serum | 34 | 28 |  |  |  | Yes | Atellica IM Multi‑Diluent 1 | 10995637 (2‑pack) 10995638 (6‑pack) 10995639 (vial) |  | 0.30–1000.00 ng/mL | 6.36–21200 μIU/mL | No | NA | NA | 0.20 ng/mL | 4.24 μIU/mL | NA | NA | 0.35 ng/mL | 0.06 ng/mL | NA | 1.27 μIU/mL | 7.42 μIU/mL | NA | Control 1 16.69 ng/mL | Control 1 353.83 μIU/mL | 1.4 | 3.3 | Control 2 37.23 ng/mL | Control 2 789.28 μIU/mL | 1.5 | 2.3 | Control 3 182.25 ng/mL | Control 3 3863.70 μIU/mL | 3.1 | 4.9 | ADVIA Centaur PRL y = 0.94x + 0.26 ng/mL r=1.0 | ADVIA Centaur PRL y = 0.94x + 5.49 μIU/mL r=1.0 | insignificant effect up to 500 mg/dL of hemoglobin | insignificant effect up to 20 mg/dL of bilirubin | insignificant effect up to 1000 mg/dL of triglycerides | No Biotin in assay formulation | 30,000 ng/mL | 636,000 μIU/mL |  |  | traceable to the World Health Organization (WHO) 3rd IRP for human Prolactin (84/500) | 10995414_EN Rev. 03, 2020-02 (OUS), 11200387_EN Rev. 03, 2020-02 (US) | IFU Document | 0.205 | 6.969 | 2.087 | 1.457 | 0.093 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995655 |  | re06 | 00630414599731 | Global | Reproductive Endocrinology | No | Atellica IM Prolactin (PRL) | Atellica IM PRL | PRL | 250 | 5 x 50 | Atellica IM CAL B | Atellica IM Multi‑Diluent 1 Atellica IM PRL MCM (SMN 10995657) | 14 | 28 | goat polyclonal anti-prolactin antibody | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | up to 8hrs at RT, 8 hrs at 2–8°C, 48 hrs at ≤ -20°C | 25 μL | No |  | 2-site sandwich immunoassay | 0.30–200.00 ng/mL | 6.36–4240.00 μIU/mL |  |  |  |  | See IFU |  |  | Atellica IM CAL B | 10995503 (2‑pack) 10995504 (6‑pack) | at 2–8°C 28 days | Lyophilized | Add 5.0 mL of special reagent water into each vial Let the vials stand for 15–20 min at RT | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of digoxin,follicle-stimulating hormone (FSH), luteinizinghormone (LH), prolactin, human chorionic gonadotropin (hCG), and TSH; equine serum | 34 | 28 |  |  |  | Yes | Atellica IM Multi‑Diluent 1 | 10995637 (2‑pack) 10995638 (6‑pack) 10995639 (vial) |  | 0.30–1000.00 ng/mL | 6.36–21200 μIU/mL | No | NA | NA | 0.20 ng/mL | 4.24 μIU/mL | NA | NA | 0.35 ng/mL | 0.06 ng/mL | NA | 1.27 μIU/mL | 7.42 μIU/mL | NA | Control 1 16.69 ng/mL | Control 1 353.83 μIU/mL | 1.4 | 3.3 | Control 2 37.23 ng/mL | Control 2 789.28 μIU/mL | 1.5 | 2.3 | Control 3 182.25 ng/mL | Control 3 3863.70 μIU/mL | 3.1 | 4.9 | ADVIA Centaur PRL y = 0.94x + 0.26 ng/mL r=1.0 | ADVIA Centaur PRL y = 0.94x + 5.49 μIU/mL r=1.0 | insignificant effect up to 500 mg/dL of hemoglobin | insignificant effect up to 20 mg/dL of bilirubin | insignificant effect up to 1000 mg/dL of triglycerides | No Biotin in assay formulation | 30,000 ng/mL | 636,000 μIU/mL |  |  | traceable to the World Health Organization (WHO) 3rd IRP for human Prolactin (84/500) | 10995414_EN Rev. 03, 2020-02 (OUS), 11200387_EN Rev. 03, 2020-02 (US) | IFU Document | 0.818 | 6.969 | 2.087 | 4.37 | 0.372 | 17.7 | 5.3 | 11.1
Atellica® IM | 10995660 |  | re07 | 00630414599786 | Global | Reproductive Endocrinology | No | Atellica IM Progesterone (PRGE) | Atellica IM PRGE | PRGE | 90 | 1 x 90 | Atellica IM CAL E | Atellica IM Multi‑Diluent 3 Atellica IM PRGE MCM (SMN 10995661) | 14 | 28 | mouse monoclonal anti-progesterone antibody | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | up to 8hrs at RT, 8 hrs at 2–8°C, 48 hrs at ≤ -20°C, Freeze samples only 1 time and mix thoroughly after thawing. | 10 μL | No |  | competitive immunoassay | 0.21–60.00 ng/mL | 0.67–190.80 nmol/L |  |  |  |  | See IFU for Normals, Pregnant and Postmenopausal Females | Males 0.28–1.22 ng/mL 0.89–3.88 nmol/L | See IFU for Pediatric Reference Intervals by Age | Atellica IM CAL E | 10995512 (2‑pack) 10995513 (6‑pack) | at 2–8°C 28 days | Lyophilized | Add 2.0 mL of special reagent water into each vial. Let vials stand for 15-20 minutes at room temperature. | 3 for lot calibration, 2 for pack recalibration | 2 | After reconsitution, low and high levels of cortisol, progesterone, and testosterone; human serum | 90 | 28 |  |  |  | Yes | Atellica IM Multi‑Diluent 3 | 10995645 |  | 0.21–600.00 ng/mL | 0.67–1908.0 nmol/L | No | NA | NA | 0.10 ng/mL | 0.32 nmol/L | NA | NA | 0.31 ng/mL | 0.18 ng/mL | NA | 0.57 nmol/L | 0.98 nmol/L | NA | Control 1 0.64 ng/mL | Control 1 2.04 nmol/L | 6.1 | 12.9 | Control 3 18.63 ng/mL | Control 3 59.24 nmol/L | 1.9 | 2.9 | Serum 53.84 ng/mL | Serum 171.21 nmol/L | 1.6 | 3.8 | ADVIA Centaur Progesterone y = 0.980x + 0.014 ng/mL r=0.99 | ADVIA Centaur Progesterone y = 0.980x + 0.044 nmol/L r=0.99 | ≤ 10% change up to 250 mg/dL (0.15 mmol/L) hemoglobin | ≤ 10% change up to 20 mg/dL (341 μmol/L) conjugated bilirubin 20 mg/dL (341 μmol/L) unconjugated bilirubin | ≤ 10% change up to 500 mg/dL (5.65 mmol/L) intralipids | No Biotin in assay formulation | N/A | N/A |  |  | standardized using internal standards manufactured analytically which are traceable to gas chromatography–mass spectroscopy (GCMS) | 10995415_EN Rev. 03, 2020-02 (OUS), 11200386_EN Rev.03, 2020-02 (US) | IFU Document | 0.246 | 6.969 | 2.48 | 2.913 | 0.112 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995659 |  | re07 | 00630414599779 | Global | Reproductive Endocrinology | No | Atellica IM Progesterone (PRGE) | Atellica IM PRGE | PRGE | 450 | 5 x 90 | Atellica IM CAL E | Atellica IM Multi‑Diluent 3 Atellica IM PRGE MCM (SMN 10995661) | 14 | 28 | mouse monoclonal anti-progesterone antibody | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | up to 8hrs at RT, 8 hrs at 2–8°C, 48 hrs at ≤ -20°C | 10 μL | No |  | competitive immunoassay | 0.21–60.00 ng/mL | 0.67–190.80 nmol/L |  |  |  |  | See IFU for Normals, Pregnant and Postmenopausal Females | Males 0.28–1.22 ng/mL 0.89–3.88 nmol/L | See IFU for Pediatric Reference Intervals by Age | Atellica IM CAL E | 10995512 (2‑pack) 10995513 (6‑pack) | at 2–8°C 28 days | Lyophilized | Add 2.0 mL of special reagent water into each vial. Let vials stand for 15-20 minutes at room temperature. | 3 for lot calibration, 2 for pack recalibration | 2 | After reconsitution, low and high levels of cortisol, progesterone, and testosterone; human serum | 90 | 28 |  |  |  | Yes | Atellica IM Multi‑Diluent 3 | 10995645 |  | 0.21–600.00 ng/mL | 0.67–1908.0 nmol/L | No | NA | NA | 0.10 ng/mL | 0.32 nmol/L | NA | NA | 0.31 ng/mL | 0.18 ng/mL | NA | 0.57 nmol/L | 0.98 nmol/L | NA | Control 1 0.64 ng/mL | Control 1 2.04 nmol/L | 6.1 | 12.9 | Control 3 18.63 ng/mL | Control 3 59.24 nmol/L | 1.9 | 2.9 | Serum 53.84 ng/mL | Serum 171.21 nmol/L | 1.6 | 3.8 | ADVIA Centaur Progesterone y = 0.980x + 0.014 ng/mL r=0.99 | ADVIA Centaur Progesterone y = 0.980x + 0.044 nmol/L r=0.99 | ≤ 10% change up to 250 mg/dL (0.15 mmol/L) hemoglobin | ≤ 10% change up to 20 mg/dL (341 μmol/L) conjugated bilirubin 20 mg/dL (341 μmol/L) unconjugated bilirubin | ≤ 10% change up to 500 mg/dL (5.65 mmol/L) intralipids | No Biotin in assay formulation | N/A | N/A |  |  | standardized using internal standards manufactured analytically which are traceable to gas chromatography–mass spectroscopy (GCMS) | 10995415_EN Rev. 03, 2020-02 (OUS), 11200386_EN Rev.03, 2020-02 (US) | IFU Document | 0.957 | 10.748 | 2.48 | 6.063 | 0.435 | 27.3 | 6.3 | 15.4
Atellica® IM | 10995708 |  | re10 | 00630414600222 | Global | Reproductive Endocrinology | No | Atellica IM Testosterone II (TSTII) | Atellica IM TSTII | TSTII | 100 | 1 x 100 | Atellica IM APW1 (SMN 10995458) | Atellica IM Multi‑Diluent 3 Atellica IM TSTII MCM (SMN 10995709) | 14 | 18 | sheep monoclonal antibody | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#EDTA Plasma;# | up to 48 hours at RT, 7 days at 2–8ºC, 7 days at ≤ -20°C | 20 μL | No |  | competitive immunoassay using direct chemiluminescent technology | 7.00–1500.00 ng/dL | 0.24–52.05 nmol/L |  |  |  |  | See IFUs |  |  | Atellica IM TSTII CAL | Provided in Kit | 14 days recon at 2–8°C 6 hrs recon at RT | Lyophilized | 2.0 mL of special reagent water. Let the vials stand for 30 minutes at room temperature | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of USP-grade estradiol, testosterone, cortisol, and progesterone spiked in charcoal-stripped, defibrinated human plasma;ked in | 91 | 28 |  |  |  | Yes | Atellica IM Multi‑Diluent 3 | 10995645 |  | 7.00–3000 ng/dL | 0.24–104.1 nmol/L | No | NA | NA | NA | NA | NA | NA | 4.43 ng/dL | 1.81 ng/dL | 4.43 ng/dL | 0.06 nmol/L | 0.15 nmol/L | 0.15 nmol/L | 17.81 ng/dL | 0.62 nmol/L | 2.2 | 5.9 | 290.14 ng/dL | 10.07 nmol/L | 1.8 | 2.9 | 1014.48 ng/dL | 35.2 nmol/L | 2.7 | 4.7 | ADVIA Centaur TSTII y = 0.97x + 2.04 ng/dL r=1.0 | ADVIA Centaur TSTII y = 0.97x + 0.07 nmol/L r=1.0 | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 15 mg/dL of conjugated bilirubin 20 mg/dL of unconjugated bilirubin | ≤ 10% change up to 1000 mg/dL of triglycerides | ≤ 10% change in results up to 30 ng/mL of biotin | N/A | N/A |  |  | traceable to the primary testosterone standard National Measurement Institute (NMI) M914 | 10995435_EN Rev. 04, 2019-08 (OUS), 11200389_EN Rev. 04, 2019-06 (US) | IFU Document | 0.381 | 6.969 | 2.48 | 2.913 | 0.173 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995707 |  | re10 | 00630414600215 | Global | Reproductive Endocrinology | No | Atellica IM Testosterone II (TSTII) | Atellica IM TSTII | TSTII | 500 | 1 x 500 | Atellica IM APW1 (SMN 10995458) | Atellica IM Multi‑Diluent 3 Atellica IM TSTII MCM (SMN 10995709) | 14 | 18 | sheep monoclonal antibody | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#EDTA Plasma;# | up to 48 hours at RT, 7 days at 2–8ºC, 7 days at ≤ -20°C | 20 μL | No |  | competitive immunoassay using direct chemiluminescent technology | 7.00–1500.00 ng/dL | 0.24–52.05 nmol/L |  |  |  |  | See IFUs |  |  | Atellica IM TSTII CAL | Provided in Kit | 14 days recon at 2–8°C 6 hrs recon at RT | Lyophilized |  | 3 for lot calibration, 2 for pack recalibration | 2 |  | 91 | 28 |  |  |  | Yes | Atellica IM Multi‑Diluent 3 | 10995645 |  | 7.00–3000 ng/dL | 0.24–104.1 nmol/L | No | NA | NA | NA | NA | NA | NA | 4.43 ng/dL | 1.81 ng/dL | 4.43 ng/dL | 0.06 nmol/L | 0.15 nmol/L | 0.15 nmol/L | 17.81 ng/dL | 0.62 nmol/L | 2.2 | 5.9 | 290.14 ng/dL | 10.07 nmol/L | 1.8 | 2.9 | 1014.48 ng/dL | 35.2 nmol/L | 2.7 | 4.7 | ADVIA Centaur TSTII y = 0.97x + 2.04 ng/dL r=1.0 | ADVIA Centaur TSTII y = 0.97x + 0.07 nmol/L r=1.0 | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 15 mg/dL of conjugated bilirubin 20 mg/dL of unconjugated bilirubin | ≤ 10% change up to 1000 mg/dL of triglycerides | ≤ 10% change in results up to 30 ng/mL of biotin | N/A | N/A |  |  | traceable to the primary testosterone standard National Measurement Institute (NMI) M914 | 10995435_EN Rev. 04, 2019-08(OUS), 11200389_EN Rev. 04, 2019-06 (US) | IFU Document | 1.415 | 6.969 | 2.126 | 6.811 | 0.643 | 17.7 | 5.4 | 17.3
Atellica® CH | 11097637 |  | di03 | 00630414595580 | Global | Diabetes | No | Atellica CH Fructosamine (Fruc) | FrucFruc | 0628 | 400 | 2 x 200 | N/A | N/A | 10 | 30 | Proteinase-K; 4-aminoantipyrine: R2: fructosaminase; peroxidase (horseradish) N-Ethyl-N-Sulphohydroxypropyl-m-toluidine (TOOS) | ;#Serum;#Lithium Heparin Plasma;#Potassium EDTA Plasma;# | Up to 3 days at 20-25C; Up to 14 days at 4-8C; Up to 60 days frozen at -20C | 16 μL | No |  | Colorimetric end-point reaction | 30–1000 μmol/L | 30–1000 μmol/L | N/A | N/A | 122–236 μmol/L | 122–236 μmol/L | N/A | N/A | N/A | Atellica CH Fruc CAL | 11099432 | 28 days | Lyophilized | Let the vials stand for 10–15 minutes at room temperature; Add 1.0 mL of room temperature deionized water: Let the vials stand for 5 minutes at room temperature; , swirl and gently invert the vials at least 5 times | 3 | 2 | human serum albumin | 180 | 30 |  | No |  | No | N/A | 11099300 |  | N/A | N/A | No | N/A | N/A | N/A | N/A | N/A | N/A | 10 μmol/L | 6 μmol/L | 28 μmol/L | N/A | N/A | N/A | 35 μmol/L | 35 μmol/L | 4 | 5.5 | 324 μmol/L | 324 μmol/L | 0.7 | 0.9 | 636 μmol/L | 636 μmol/L | 0.4 | 1.4 | ADVIA Chemistry 1800 y = 1.02x + 3 μmol/L r=0.994 n= 112 | ADVIA Chemistry 1800 y = 1.02x + 3 μmol/L r=0.994 n= 112 | ≤ 3% change up to 750 mg/dL of hemoglobin. | ≤ 8% change up to 5 mg/dL of conjugated bilirubin; ≤ 8% change up to 7.5 mg/dL of unconjugated bilirubin | ≤ 6% change up to 1000 mg/dL of Trig. | No Biotin in assay formulation | N/A | N/A | N/A | N/A | Internal standard | 11110200_EN Rev. 02, 2019-07 | IFU Document | 0.341 | 2.875 | 3.5 | 2.75 | 0.155 | 7.302 | 8.89 | 6.985
Atellica® CH | 11318376 |  | pl33 | 05055273215564 | OUS | Specific Proteins | Yes | Soluble Transferrin Receptors (Randox Alliance Application) | sTfRsTfR | 8092 | 85 | 1 x 85 | Atellica CH Empty Reagent Packs (11097534) | N/A | 9 | 14 | Antibody-latex Reagent; Mouse monoclonal anti-STFR antibodies | ;#Serum;#Lithium Heparin Plasma;# | If not analysed immediately, serum should be stored at -20°C. | 8.9 | No |  | Turbidimetry | 0.50 – 11.77 mg/L | 0.50 – 11.77 mg/L | N/A | N/A | 0.65 - 1.88 mg/L | 0.65 - 1.88 mg/L | N/A | N/A | N/A | Randox STFR Calibrator | 11306493 | 28 days | Lyophilized | 1. Open vial very carefully, avoiding any loss of material. 2. Reconstitute each vial with 1 ml of double deionised water. 3. Replace stopper, close vial and mix carefully for 30 minutes. 4. Ensure contents are completely dissolved before use. | 3 | 6 | Human Source | 6 | 6 |  | No |  | No | N/A | 11099300 |  | N/A | N/A | No | Randox STFR CONTROL | 11309043 | N/A | N/A | N/A | N/A | 0.13 mg/L | 0.7 mg/L | 0.50 mg/L | 0.7 mg/L | 0.13 mg/L | 0.50 mg/L | 0.61 mg/L | 0.61 mg/L | 8.8 | 9.9 | 1.95 mg/L | 1.95 mg/L | 2.5 | 5 | 4.75 mg/L | 4.75 mg/L | 1.5 | 2.7 | Roche cobas 501 y=0.97x -0.023 mg/L r=0.997 | Roche cobas 501 y=0.97x -0.023 mg/L r=0.997 | ≤ 10% change up to 175 mg/dL of hemoglobin. | ≤ 10% change up to 60 mg/dL of free and conjugate bilirubin. | ≤ 10% change up to 400 mg/dL of Intralipd | No Biotin in assay formulation | N/A | N/A | N/A | N/A | WHO Reference Reagent Recombinant Soluble Transferrin Receptor (NIBSC code:07/202) | IFU: Revised 27 Feb 18 ml and App. Sheet: (s): Revised 28 Mar 18 ml | IFU Document | 0.1201 | 4.8 | 1.96 | 1.88 | 0.0545 | 12.2 | 5 | 4.8
Atellica® IM | 10997840 |  | ca02 | 00630414006703 | OUS | Cardiac | No | Atellica IM High Sensitivity Troponin I (TnIH) | Atellica IM TnIH | TnIH | 100 | 1X100 | Atellica IM APW3 ( SMN 10998580) | Atellica IM Multi‑Diluent 11 (10995642)Atellica IM TnIH MCM (SMN 10997842) | 10 | 31 | Bovine serum albumin (BSA) conjugated to a recombinant monoclonal (sheep) Fab anti-human cTnI | ;#Serum;#Lithium Heparin Plasma;# | RT for up to 8 hours 2–8°C for 24 hours ≤ -20°C for up to 40 days ≤ -70°C for up to 1 year | 100 μL | No |  | direct chemiluminescent technology | 2.51–25,000.00 pg/mL | 2.51–25,000.00 ng/L |  |  |  |  | Combined gender and lithium heparin sample type were used to determine the overall 99th percentile of 45.20 pg/mL (ng/L) |  |  | Atellica IM TnIH CAL | Provided in Kit | 4 hrs recon at 2–8°C 30 days recon at ≤ -20°C 4 hrs onboard at RT | Lyophilized | Low Calibrator is liquid ready to use. High Calibrator - Reconstituted to 1.00 mL/vial | 3 for lot calibration, 2 for pack recalibration | 2 | HEPES buffer; bovine serum albumin (BSA) | 47 | 31 |  |  |  | Yes | Atellica IM Multi‑Diluent 11 | 10995642 (2‑pack) |  | 2.50 - 125,000 pg/mL | 2.50 - 125,000 ng/L | No | NA | NA | NA | NA | NA | NA | 1.27 pg/mL | 0.58 pg/mL | 2.51 pg/mL | 0.58 ng/L | 1.27 ng/L | 2.51 ng/L | 9.54 | N/A | 5 | 7.5 | 147.37 | N/A | 1.7 | 3 | 21 392.14 | N/A | 1.9 | 2.5 | See IFU | See IFU | <10% change up to 500 mg/dL | <10% change up to 40 mg/dL of conjugated bilirubin . 60 mg/dL of unconjugated bilirubin | <10% change up to 2000 mg/dL of triglycerides | ≤ 10% change in results up to 1500 ng/mL of biotin. | 618,000 pg/mL | 618,000 ng/L |  |  | Traceable to an internal standard manufactured using human heart homogenate | 11200497_EN Rev. 06, 2019-08 | IFU Document | 0.286 | 6.875 | 2.125 | 2.438 | 0.13 | 17.463 | 5.398 | 6.193
Atellica® IM | 10997841 |  | ca02 | 00630414006758 | OUS | Cardiac | No | Atellica IM High Sensitivity Troponin I (TnIH) | Atellica IM TnIH | TnIH | 500 | 5 X 100 | Atellica IM APW3 (SMN 10998580) | Atellica IM Multi‑Diluent 11 (SMN 10995642) Atellica IM TnIH MCM (SMN 10997842) | 10 | 28 | Bovine serum albumin (BSA) conjugated to a recombinant monoclonal (sheep) Fab anti-human cTnI | ;#Serum;#Lithium Heparin Plasma;# | RT for up to 8 hours 2–8°C for 24 hours ≤ -20°C for up to 40 days ≤ -70°C for up to 1 year | 100 μL | No |  | direct chemiluminescent technology | 2.51–25,000.00 pg/mL | 2.51–25,000.00 ng/L |  |  |  |  | Combined gender and lithium heparin sample type were used to determine the overall 99th percentile of 45.20 pg/mL (ng/L) |  |  | Atellica IM TnIH CAL | Provided in Kit | 4 hrs recon at 2–8°C 30 days recon at ≤ -20°C 4 hrs onboard at RT | Lyophilized | Low calibrator is liquid ready to use. High calibrator - Reconstituted to 1.00 mL/vial | 3 for lot calibration, 2 for pack recalibration | 2 |  | 47 | 31 |  |  |  | Yes | Atellica IM Multi‑Diluent 11 | 10995642 (2‑pack) |  | 2.50 - 125,000 pg/mL | 2.50 - 125,000 ng/L | No | NA | NA | NA | NA | NA | NA | 1.27 pg/mL | 0.58 pg/mL | 2.51 pg/mL | 0.58 ng/L | 1.27 ng/L | 2.51 ng/L | 9.54 | N/A | 5 | 7.5 | 147.37 | N/A | 1.7 | 3 | 21 392.14 | N/A | 1.9 | 2.5 | See IFU | See IFU | <10% change up to 500 mg/dL | <10% change up to 40 mg/dL of conjugated bilirubin . 60 mg/dL of unconjugated bilirubin | <10% change up to 2000 mg/dL of triglycerides | ≤ 10% change in results up to 1500 ng/mL of biotin. | 618,000 pg/mL | 618,000 ng/L |  |  | Traceable to an internal standard manufactured using human heart homogenate | 11200497_EN Rev. 06, 2019-08 | IFU Document | 0.953 | 6.875 | 2.125 | 6.75 | 0.433 | 17.463 | 5.398 | 17.145
Atellica® IM | 10995456 |  | id06 | 00630414597867 | OUS | Hepatitis | No | Atellica IM Hepatitis C Antibody (aHCV) | Atellica IM aHCVaHCV | aHCV | 200 | 1 x 200 | Atellica IM APW1 (SMN 10995458) | Atellica IM aHCV QC | 46 | 41 | anti-human IgG monoclonal antibody | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#EDTA Plasma;# | 2–8°C for up to 7 days, ≤ -20°C for longer storage | 10 μL | No |  | indirect two wash sandwich immunoassay | 0.02–11.00 Index | 0.02–11.00 Index | NA | NA | NA | NA | nonreactive <0.80 Index | reactive ≥ 1.00 Index | equivocal ≥ 0.80 Index and <1.00 Index | Atellica IM aHCV CAL | Provided in Kit | at 2–8°C until expiration 8 hours at RT | Liquid | N/A | 3 for lot calibration, 2 for pack recalibration | 2 | processed human plasma negative and positive for anti-HCV with preservatives | 60 | 28 |  |  |  | No | NA | NA |  | NA | NA | Yes | Atellica IM aHCV QC | 10995457 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 0.80 Index | 0.80 Index | 3.4 | 11.5 | 3.97 Index | 3.97 Index | 4.7 | 7.2 | 8.61 Index | 8.61 Index | 6.3 | 7.2 | NA | NA | ≤ 10% change up to up to 500 mg/dL of hemoglobin | ≤ 10% change up to 60 mg/dL of conjugated bilirubin 40 mg/dL of unconjugated bilirubin | ≤ 10% change up to 1000 mg/dL of triglycerides | Insignificant effect on the assay up to 1500 ng/mL | NA | NA |  |  | Based upon relative clinical agreement with commercially available anti-HCV assays | 10995360_EN Rev. 03, 2019-07 | IFU Document | 0.389 | 6.969 | 2.48 | 2.913 | 0.177 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995580 |  | re05 | 00630414599045 | Global | Reproductive Endocrinology | No | Atellica IM Follicle Stimulating Hormone (FSH) | Atellica IM FSH | FSH | 190 | 1 x 190 | Atellica IM CAL B | Atellica IM Multi‑Diluent 1 Atellica IM FSH MCM(SMN 10995582) | 14 | 28 | mouse monoclonal anti-FSH antibody | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | RT up to 8 hours > 8 hours at 2–8°C > 48 hr at ≤ -20°C Freeze 1 time | 50 μL | No |  | 2-site samdwich immonoassay | 0.30–200.00 mIU/mL | 0.30–200.00 IU/L |  |  |  |  | See IFU |  |  | Atellica IM CAL B | 10995503 (2‑pack) 10995504 (6‑pack) | 28 days recon at 2–8°C 4 hours at RT | Lyophilized | Add 5.0 mL of special reagent water into each vial Let the vials stand for 15–20 min at RT | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of digoxin,follicle-stimulating hormone (FSH), luteinizinghormone (LH), prolactin, human chorionic gonadotropin (hCG), and TSH; equine serum | 36 | 14 |  |  |  | Yes | Atellica IM Multi‑Diluent 1 | 10995637 (2‑pack) 10995638 (6‑pack) 10995639 (vial) |  | 0.30–400.00 mIU/mL | 0.30–400.00 IU/L | No | NA | NA | 0.02 mIU/mL | 0.02 IU/L | NA | NA | 0.55 mIU/mL | 0.20 mIU/mL | Not Performed | 0.20 IU/L | 0.55 IU/L | Not Performed | 1.75 mIU/mL | 1.75 IU/L | 6.3 | 8.2 | 50.78 mIU/mL | 50.78 IU/L | 1.5 | 2.9 | 169.96 mIU/mL | 169.96 IU/L | 2 | 2.9 | ADVIA Centaur FSH y = 0.98x + 1.92 mIU/mL r=1.00 | ADVIA Centaur FSH y = 0.98x + 1.92 IU/L r=1.00 | ≤ 10% change up to up to 150 mg/dL of hemoglobin | ≤ 10% change up to 20 mg/dL of conjugated bilirubin 20 mg/dL of unconjugated bilirubin | ≤ 10% change up to 500 mg/dL of lipemia | No Biotin in assay formulation | 1000 mIU/mL | 1000 IU/L |  |  | Traceable to World Health Organization (WHO) 2nd International Standard for human FSH (IS 94/632) | 10995346_EN Rev. 03, 2020-02 (OUS) 11200384_EN Rev. 03, 2020-02 (US) | IFU Document | 0.205 | 6.969 | 2.087 | 1.457 | 0.093 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995581 |  | re05 | 00630414599052 | Global | Reproductive Endocrinology | No | Atellica IM Follicle Stimulating Hormone (FSH) | Atellica IM FSH | FSH | 950 | 5 x 190 | Atellica IM CAL B | Atellica IM Multi‑Diluent 1 Atellica IM FSH MCM (SMN 10995582) | 14 | 28 | mouse monoclonal anti-FSH antibody | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | RT up to 8 hours > 8 hours at 2–8°C >48 hr at ≤ -20°C Freeze 1 time | 50 μL | No |  | 2-site samdwich immonoassay | 0.30–200.00 mIU/mL | 0.30–200.00 IU/L |  |  |  |  | See IFU |  |  | Atellica IM CAL B | 10995503 (2‑pack) 10995504 (6‑pack) | 28 days recon at 2–8°C 4 hours at RT | Lyophilized | Add 5.0 mL of special reagent water into each vial Let the vials stand for 15–20 min at RT | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of digoxin,follicle-stimulating hormone (FSH), luteinizinghormone (LH), prolactin, human chorionic gonadotropin (hCG), and TSH; equine serum | 36 | 14 |  |  |  | Yes | Atellica IM Multi‑Diluent 1 | 10995637 (2‑pack) 10995638 (6‑pack) 10995639 (vial) |  | 0.30–400.00 mIU/mL | 0.30–400.00 IU/L | No | NA | NA | 0.02 mIU/mL | 0.02 IU/L | NA | NA | 0.20 mIU/mL | 0.20 mIU/mL | Not Performed | 0.20 IU/L | 0.20 IU/L | Not Performed | 1.75 mIU/mL | 1.75 IU/L | 6.3 | 8.2 | 50.78 mIU/mL | 50.78 IU/L | 1.5 | 2.9 | 169.96 mIU/mL | 169.96 IU/L | 2 | 2.9 | ADVIA Centaur FSH y = 0.98x + 1.92 mIU/mL r=1.00 | ADVIA Centaur FSH y = 0.98x + 1.92 IU/L r=1.00 | ≤ 10% change up to up to 150 mg/dL of hemoglobin | ≤ 10% change up to 20 mg/dL of conjugated bilirubin 20 mg/dL of unconjugated bilirubin | ≤ 10% change up to 500 mg/dL of lipemia | No Biotin in assay formulation | 1000 mIU/mL | 1000 IU/L |  |  | Traceable to World Health Organization (WHO) 2nd International Standard for human FSH (IS 94/632) | 10995346_EN Rev. 03, 20120-02 (OUS), 11200384_EN Rev.03, 2020-02 (US) | IFU Document | 0.812 | 6.969 | 2.087 | 4.37 | 0.369 | 17.7 | 5.3 | 11.1
Atellica® IM | 10995635 |  | re02 | 00630414599533 | Global | Reproductive Endocrinology | No | Atellica IM Luteinizing Hormone (LH) | Atellica IM LH | LH | 110 | 1 x 110 | Atellica IM CAL B | Atellica IM Multi‑Diluent 1 Atellica IM LH MCM (SMN 10995636) | 14 | 28 | mouse monoclonal anti-LH antibody | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | RT up to 8 hours > 8 hours at 2–8°C >48 hr at ≤ -20°C Freeze 1 time | 25 μL | No |  | 2-site sandwich immunoassay | 0.07–200.00 mIU/mL | 0.07–200.00 IU/L |  |  |  |  | See IFU |  |  | Atellica IM CAL B | 10995503 (2‑pack) 10995504 (6‑pack) | 28 days recon at 2–8°C 4 hours at RT | Lyophilized | Add 5.0 mL of special reagent water into each vial Let the vials stand for 15–20 min at RT | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of digoxin,follicle-stimulating hormone (FSH), luteinizinghormone (LH), prolactin, human chorionicgonadotropin (hCG), and TSH; equine serum | 65 | 28 |  |  |  | Yes | Atellica IM Multi‑Diluent 1 | 10995637 (2‑pack) 10995638 (6‑pack) 10995639 (vial) |  | 0.07–400.00 mIU/mL | 0.07–400.00 IU/L | No | NA | NA | 0.06 mIU/mL | 0.06 IU/L | NA | NA | 1.44 mIU/mL | 0.07 mIU/mL | Not Performed | 0.07 IU/L | 1.44 IU/L | Not Performed | 18.74 mIU/mL | 18.74 IU/L | 1.8 | 3.4 | 63.64 mIU/mL | 63.64 IU/L | 1.4 | 2.2 | 177.25 mIU/mL | 177.25 IU/L | 1.1 | 2.1 | ADVIA Centaur LH y = 1.06x + 2.06 mIU/mL r=1.00 | ADVIA Centaur LH y = 1.06x + 2.06 IU/L r=1.00 | ≤ 10% change up to up to 500 mg/dL of hemoglobin | ≤ 10% change up to 20 mg/dL of conjugated bilirubin 20 mg/dL of unconjugated bilirubin | ≤ 10% change up to 3000mg/dL of lipemia | No Biotin in assay formulation | 18,000 mIU/mL | 18,000 IU/L |  |  | Traceable to World Health Organization (WHO) 2nd International Standard for human LH (IS 80/552). | 10995369_EN Rev. 03, 2020-02 (OUS) 11200385_EN Rev. 03, 2020-02 (US) | IFU Document | 0.213 | 6.969 | 2.087 | 1.457 | 0.097 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995634 |  | re02 | 00630414599526 | Global | Reproductive Endocrinology | No | Atellica IM Luteinizing Hormone (LH) | Atellica IM LH | LH | 550 | 5 x 110 | Atellica IM CAL B | Atellica IM Multi‑Diluent 1 Atellica IM LH MCM (SMN 10995636) | 14 | 28 | mouse monoclonal anti-LH antibody | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | RT up to 8 hours > 8 hours at 2–8°C >48 hr at ≤ -20°C Freeze 1 time | 25 μL | No |  | 2-site sandwich immunoassay | 0.07–200.00 mIU/mL | 0.07–200.00 IU/L |  |  |  |  | See IFU |  |  | Atellica IM CAL B | 10995503 (2‑pack) 10995504 (6‑pack) | 28 days recon at 2–8°C 4 hours at RT | Lyophilized | Add 5.0 mL of special reagent water into each vial Let the vials stand for 15–20 min at RT | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of digoxin,follicle-stimulating hormone (FSH), luteinizinghormone (LH), prolactin, human chorionic gonadotropin (hCG), and TSH; equine serum | 65 | 28 |  |  |  | Yes | Atellica IM Multi‑Diluent 1 | 10995637 (2‑pack) 10995638 (6‑pack) 10995639 (vial) |  | 0.07–400.00 mIU/mL | 0.07–400.00 IU/L | No | NA | NA | 0.06 mIU/mL | 0.06 IU/L | NA | NA | 1.44 mIU/mL | 0.07 mIU/mL | Not Performed | 0.07 IU/L | 1.44 IU/L | Not Performed | 18.74 mIU/mL | 18.74 IU/L | 1.8 | 3.4 | 63.64 mIU/mL | 63.64 IU/L | 1.4 | 2.2 | 177.25 | 177.25 IU/L | 1.1 | 2.1 | ADVIA Centaur LH y = 1.06x + 2.06 mIU/mL r=1.00 | ADVIA Centaur LH y = 1.06x + 2.06 IU/L r=1.00 | ≤ 10% change up to up to 500 mg/dL of hemoglobin | ≤ 10% change up to 20 mg/dL of conjugated bilirubin 20 mg/dL of unconjugated bilirubin | ≤ 10% change up to 3000mg/dL of lipemia | No Biotin in assay formulation | 18,000 mIU/mL | 18,000 IU/L |  |  | Traceable to World Health Organization (WHO) 2nd International Standard for human LH (IS 80/552). | 10995369_EN Rev. 03, 2020-02 (OUS) 11200385_EN Rev. 03, 2020-02 (US) | IFU Document | 0.851 | 6.969 | 2.087 | 4.37 | 0.387 | 17.7 | 5.3 | 11.1
Atellica® IM | 10995572 |  | an02 | 00630414598963 | OUS | Anemia | No | Atellica IM Folate (Fol) | Atellica IM FolFol | Fol | 140 | 1 x 140 | Atellica IM Fol DTT/REL (SMN 10995576) Atellica IM APW1 (SMN 10995458) | Atellica IM RBC Fol (SMN 10995440) Atellica IM Fol MCM (SMN 10995575) | 14 | 14 | Bovine and human serum albumi | ;#Serum;#Whole blood;# | Serum < 8 hrs at RT > 8 hrs at 2–8°C >48 hrs at ≤ -20°C | 100 μL | No |  | Competitive immunoassay using direct chemiluminescent technology | 0.35–24.00 ng/mL | 0.79–54.36 nmol/L | 0.35 - 48.00 ng/mL | 0.79 - 108.72 nmol/L | Serum Folate; Deficient 0.35–3.37 ng/mL Indeterminate 3.38–5.38 ng/mL Normal > 5.38 ng/mL | Serum Folate; Deficient 0.79–7.63 nmol/L Indeterminate 7.64–12.19 nmol/L Normal > 12.19 nmol/L |  |  |  | Atellica IM Fol CAL | Provided in Kit | 7 days at 2–8°C<br>8 hours at RT<br>28 days ≤ -20°C | Lyophilized | Add 3.0 mL of reagent water into each calibrator vial. Let the calibrators stand for 15–20 minutes at room temperature. | 3 for lot calibration, 2 for pack recalibration | 2 | Human serum albumin | 14 | 7 |  |  |  | Yes | Atellica IM Fol DIL | 10995574 |  | 0.35 - 48.00 ng/mL | 0.79 - 108.72 nmol/L | No | NA | NA | < 0.01 ng/mL | < 0.02 nmol/L | NA | NA | 0.38 ng/mL | 0.19 ng/mL | NA | 0.43 nmol/L | 0.86 nmol/L | NA | 4.13 ng/mL | 9.35 nmol/L | 2.4 | 5.9 | 9.23 ng/mL | 20.91 nmol/L | 2.6 | 6.5 | 21.91 ng/mL | 49.63 nmol/L | 3.3 | 7.1 | ADVIA Centaur Folate y = 0.94x - 0.01 ng/mL r=0.99 | ADVIA Centaur Folate y = 0.94x - 0.02 nmol/L r=0.99 | NA | ≤ 10% change up to 20 mg/dL conjugated bilirubin, 20 mg/dL unconjugated bilirubin | ≤ 10% change up to 2000 mg/dL Lipemia | ≤ 10% change up to 38 ng/mL of biotin | NA | NA | NA | NA | Traceable to an internal standard manufactured using highly purified material (N‑5‑methyl tetrahydrofolate) | 10995344_EN Rev. 03, 2019-07 | IFU Document | 0.411 | 6.969 | 2.48 | 2.913 | 0.187 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995573 |  | an02 | 00630414598970 | OUS | Anemia | No | Atellica IM Folate (Fol) | Atellica IM FolFol | Fol | 700 | 5 x 140 | Atellica IM Fol DTT/REL (SMN 10995576) Atellica IM APW1 (SMN 10995458) | Atellica IM RBC Fol (SMN 10995440) Atellica IM Fol MCM (SMN 10995575) | 14 | 14 | Bovine and human serum albumi | ;#Serum;#Whole blood;# | Serum < 8 hrs at RT >8 hrs at 2–8°C >48 hrs at ≤ -20°C | 100 μL | No |  | Competitive immunoassay using direct chemiluminescent technology | 0.35–24.00 ng/mL | 0.79–54.36 nmol/L | 0.35 - 48.00 ng/mL | 0.79 - 108.72 nmol/L | Serum Folate; Deficient 0.35–3.37 ng/mL Indeterminate 3.38–5.38 ng/mL Normal > 5.38 ng/mL | Serum Folate; Deficient 0.79–7.63 nmol/L Indeterminate 7.64–12.19 nmol/L Normal > 12.19 nmol/L |  |  |  | Atellica IM Fol CAL | Provided in Kit | 7 days at 2–8°C<br>8 hours at RT<br>28 days ≤ -20°C | Lyophilized | Add 3.0 mL of reagent water into each calibrator vial. Let the calibrators stand for 15–20 minutes at room temperature. | 3 for lot calibration, 2 for pack recalibration | 2 | Human serum albumin | 14 | 7 |  |  |  | Yes | Atellica IM Fol DIL | 10995574 |  | 0.35-48.00 ng/mL | 0.79-108.72 nmol/L | No | NA | NA | < 0.01 ng/mL | < 0.02 nmol/L | NA | NA | 0.38 ng/mL | 0.19 ng/mL | NA | 0.43 nmol/L | 0.86 nmol/L | NA | 4.13 ng/mL | 9.35 nmol/L | 2.4 | 5.9 | 9.23 ng/mL | 20.91 nmol/L | 2.6 | 6.5 | 21.91 (ng/mL) | 49.63 (nmol/L) | 3.3 | 7.1 | ADVIA Centaur Folate y = 0.94x - 0.01 ng/mL r=0.99 | ADVIA Centaur Folate y = 0.94x - 0.02 nmol/L r=0.99 | NA | ≤ 10% change up to 20 mg/dL conjugated bilirubin, 20 mg/dL unconjugated bilirubin | ≤ 10% change up to 2000 mg/dL Lipemia | ≤ 10% change up to 38 ng/mL of biotin | NA | NA | NA | NA | Traceable to an internal standard manufactured using highly purified material (N‑5‑methyl tetrahydrofolate) | 10995344_EN Rev. 03, 2019-07 | IFU Document | 1.208 | 6.969 | 2.126 | 6.811 | 0.549 | 17.7 | 5.4 | 17.3
Atellica® IM | 10995621 |  | bo02 | 00630414599441 | OUS | Bone Metabolism | No | Atellica IM Intact Parathyroid Hormone (PTH) | Atellica IM PTH | PTH | 190 |  | NA | Atellica IM PTH QC (10995626) Atellica IM Multi‑Diluent 13 (10995643) Atellica IM PTH MCM (SMN 10995625) | 10 | 28 | Two mouse monoclonal anti-human PTH antibodies | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#EDTA Plasma;# | Serum: 25°C 8 hours 2–8°C 8 hours -20°C 1 month Plasma EDTA: 25°C 25 hours 2–8°C 14 days -20°C Not recommended Plasma Lithium Heparin & Sodium Heparin: 25°C 9 hours 2–8°C 72 hours -20°C Not recommended | 25 μL | No |  | 2-site sandwich immunoassay using direct chemiluminescent technology | 6.3–2000.0 pg/mL (US); 4.6-2000 pg/mL (OUS) | 0.7–212.0 pmol/L (US); 0.5-212.0 pmol/L (OUS) |  |  |  |  | Plasma: 18.4–80.1 pg/mL (1.95–8.49 pmol/L) | Serum: 18.5–88.0 pg/mL (1.96–9.33 pmol/L) |  | IM PTH CAL low | Provided in Kit | 60 days recon at ≤ -20°C 8 hours recon at RT | Lyophilized | Reconstitute to 1mL/vial | 3 for lot calibration, 2 for pack recalibration | 2 | Human plasma-based product | 91 | 14 |  |  |  | Yes | Atellica IM Muti-Diluent 13 | 10995643 |  | 6.3–10,000pg/mL (US); 4.6-10,000pg/mL (OUS) | 0.7 - 8480 pmol/L; 0.5-8480 pmol/L (OUS) | Yes | Atellica IM PTH QC | 10995626 (2x1.0mL) | NA | NA | NA | NA | 1.5 pg/mL | 1.4 pg/mL | 2.3 pg/mL | 0.15 pmol/L | 0.16 pmol/L | 0.24 pmol/L | 42.0 pg/mL (Control 1 | 4.4.pmol/L | 4.6 | 4.7 | 240.7 pg/mL (Control 2) | 25.5 pmol/L (Control 2) | 1.7 | 2.9 | 868.7 pmol/L (Control 3) | 92.1 pg/mL (Control 3) | 1.4 | 2.3 | ADVIA Centaur PTH y = 1.03x - 1.08 pg/mL r=1.00 | ADVIA Centaur PTH y = 1.03x - 0.11 pmol/L r=1.00 | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 60 mg/dL of conjugated bilirubin 60 mg/dL of unconjugated bilirubin | ≤ 10% change up to 1000 mg/dL of Lipemia | ≤ 10% change up to 1500 ng/mL | 100,000 pg/mL | 10,600 pmol/L |  |  | The Atellica IM PTH assay standardization is maintained with internal standards using purified human PTH (1-84) | 10995367_EN Rev. 04, 2021-01-27 | IFU Document | 0.308 | 6.969 | 2.087 | 1.457 | 0.14 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995622 |  | bo02 | 00630414599458 | OUS | Bone Metabolism | No | Atellica IM Intact Parathyroid Hormone (PTH) | Atellica IM PTH | PTH | 950 |  | NA | Atellica IM PTH QC (10995626) Atellica IM Multi‑Diluent 13 (10995643) Atellica IM PTH MCM (SMN 10995625) | 10 | 28 | Two mouse monoclonal anti-human PTH antibodies | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#EDTA Plasma;# | Serum: 25°C 8 hours 2–8°C 8 hours -20°C 1 month Plasma EDTA: 25°C 25 hours 2–8°C 14 days -20°C not recommended Plasma Lithium Heparin & Sodium Heparin: 25°C 9 hours 2–8°C 72 hours -20°C not recommended | 25 μL | No |  | 2-site sandwich immunoassay using direct chemiluminescent technology | 6.3–2000.0 pg/mL (US); 4.6-2000.0 pg/mL (OUS) | 0.7–212.0 pmol/L (US); 0.5-212 pmol/L (OUS) |  |  |  |  | Plasma: 18.4–80.1 pg/mL (1.95–8.49 pmol/L) | Serum: 18.5–88.0 pg/mL (1.96–9.33 pmol/L) |  | IM PTH CAL low | Provided in Kit | 60 days recon at ≤ -20°C 8 hours recon at RT | Lyophilized | Reconstitute to 1.0 mL/vial | 3 for lot calibration, 2 for pack recalibration | 2 | Human plasma-based product | 91 | 14 |  |  |  | Yes | Atellica IM Multi-Diluent 13 | 10995643 |  | 6.3–10,000pg/mL (US); 4.6-10,000 pg/mL (OUS) | 0.7 - 1060 pmol/L (US); 0.5-1060 pmol/L (OUS) | Yes | Atellica IM PTH QC | 10995626 (2x1.0mL) | NA | NA | NA | NA | 1.5 pg/mL | 1.4 pg/mL | 2.3 pg/mL | 0.15 pmol/L | 0.16 pmol/L | 0.24 pmol/L | 42.0 pg/mL (Control 1) | 4.4 pmol/L | 4.6 | 4.7 | 240.7 pg/mL (Control 2) | 25.5 pmol/L (Control) | 1.7 | 2.9 | 868.7 pmol/L (Control 3) | 92.1 pg/mL (Control) | 1.4 | 2.3 | ADVIA Centaur PTH y = 1.03x - 1.08 pg/mL r=1.00 | ADVIA Centaur PTH y = 1.03x - 0.11 pmol/L r=1.00 | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 60 mg/dL of conjugated bilirubin 60 mg/dL of unconjugated bilirubin | ≤ 10% change up to 1000 mg/dL of Lipemia | ≤ 10% change up to 1500 ng/mL | 100,000 pg/mL | 10,600 pmol/L |  |  | The Atellica IM PTH assay standardization is maintained with internal standards using purified human PTH (1-84) | 10995367_EN Rev. 04, 2021-01-27 | IFU Document | 0.999 | 6.969 | 2.087 | 4.37 | 0.454 | 17.7 | 5.3 | 11.1
Atellica® IM | 11202699 |  | se01 | 00630414246550 | OUS | Sepsis | No | Atellica IM BRAHMS Procalcitonin (PCT) | Atellica IM BRAHMS PCTPCT | PCT | 100 | 1 X 100 | NA | Atellica IM BRAHMS PCT MCM (SMN 10998265) | 18 | 60 | Mouse monoclonal anti-PCT antibody | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#EDTA Plasma;# | < 8 hrs at RT, > 8 hrs at 2–8°C, >48 hrs at ≤ -20°C, 5 freeze thaw cycles | 100 μL | No |  | 2-site sandwich immunoassay using direct chemiluminescent technology | 0.03–50.00 ng/mL | 0.03–50.00 ng/mL | 0.03–1000.00 ng/mL | 0.03–1000.00 ng/mL | < 0.05 ng/mL | < 0.05 ng/mL |  |  |  | Atellica IM BRAHMS PCT CAL | Included in the kit | 8 hours at RT; 24 hrs at 2–8°C, 60 days at ≤ -20°C; thaw up to 2x | Lyophilized | Recon with 2.0 mL of special reagent water, let stand for 15–20 minutes at RT | 3 for lot calibration, 2 for pack recalibration | 2 | Recombinant PCT; equine serum | 82 | 35 |  |  |  | Yes | Atellica IM Multi‑Diluent 1 | 10995637 (2‑pack) 10995638 (6‑pack) 10995639 (vial) |  | 0.03–1000.00 ng/mL | 0.03–1000.00 ng/mL | Yes | Atellica IM BRAHMS PCT QC | 11202700 | NA | NA | NA | NA | 0.03 ng/mL | 0.00 ng/mL | 0.04 ng/mL | 0.00 ng/mL | 0.03 ng/mL | 0.04 ng/mL | 0.05 ng/mL | 0.05 ng/mL | 9.7 | 12.1 | 1.52 ng/mL | 1.52 ng/mL | 1.4 | 2.6 | 19.14 ng/mL | 19.14 ng/mL | 1.5 | 2.2 | B·R·A·H·M·S PCT sensitive Kryptor y= 1.02 x - 0.02 ng/mL r= 0.98; N= 522 | B·R·A·H·M·S PCT sensitive Kryptor y= 1.02 x - 0.02 ng/mL r= 0.98; N= 522 | 500 mg/dL at conc 0.36 ng/mL -3.6% bias, at conc 1.73 ng/mL -3.4% bias | Conjugted Bili 40 mg/dL at conc 0.34 ng/mL -3.1% bias at conc 1.61 ng/mL -1.6% bias Unconjugated 40 mg/dL at conc 0.34 ng/mL 6.6% bias, at conc 1.67 ng/mL 3.8% bias | 1000 mg/dL at conc 0.35 ng/mL -2.9% bias, at conc 1.62 ng/mL 1.4% bias | No Biotin in assay formulation | 2000 ng/mL | 2000 ng/mL | 2000 ng/mL | 2000 ng/mL | Specified by BRAHMS with N-terminal amino acid sequencing (Edmann’s method) and Mass analysis | 11202687_EN Rev. 03, 2019-07 | IFU Document | 0.343 | 6.969 | 2.48 | 2.913 | 0.156 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995478 |  | on08 | 00630414598055 | Global | Oncology | No | Atellica IM BR 27.29 (BR) | Atellica IM BR | BR | 50 | 1x50 | Atellica IM BR CAL (SMN 10995515) Atellica IM BR PRE (SMN 10995480) | Atellica IM Multi‑Diluent 1 (SMN 10995637, 2-pack; SMN 10995638, 6-pack; SMN 10995639, vial) Atellica IM BR MCM (SMN 10995479) | 14 | 28 | Mouse monoclonal anti-CA 27.29 antibody | ;#Serum;#EDTA Plasma;# | < 8 hrs at RT > 8 hrs at 2–8°C >48 hrs at ≤ -20°C | 25 μL | Yes |  | Competitive immunoassay using direct chemiluminescent technology | 3.5–450.0 U/mL | NA |  |  |  |  | < 38.6 U/mL | See IFU for other results for other conditions |  | Atellica IM BR CAL | 10995515 (2‑pack) | 28 days recon at 2–8°C 32 hours recon at RT | Lyophilized | 2.0 mL of special reagent water | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of CA 27.29 (human) | 32 | 10 |  |  |  | Yes | Atellica IM Multi‑Diluent 1 | 10995637 (2‑pack) 10995638 (6‑pack) 10995639 (vial) |  | 3.5–9000 U/mL | NA | No | NA | NA | < 0.1 U/mL | NA | NA | NA | 4.1 U/mL | 1.0 U/mL | Not Performed | NA | NA | NA | 31.2 U/mL (Control) | NA | 4.5 | 7.3 | 154.6 U/mL (Control) | NA | 1.7 | 3.3 | 317.1 U/mL | NA | 2.1 | 4.1 | ADVIA Centaur BR y = 1.02x + 0.1 U/mL r=1.00 | NA | ≤ 5% change up to 500 mg/dL hemoglobin | ≤ 5% change up to 20 mg/dL conjugated bilirubin, 20 mg/dL unconjugated bilirubin | ≤ 5% change up to 3000 mg/dL Triglycerides | No Biotin in assay formulation | NA | NA |  |  | Traceable to an internal standard manufactured using highly purified CA 27.29 | 10995282_EN Rev. 03, 2019-08 | IFU Document | 0.218 | 6.969 | 2.087 | 1.457 | 0.099 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995477 |  | on08 | 00630414598048 | Global | Oncology | No | Atellica IM BR 27.29 (BR) | Atellica IM BR | BR | 250 | 5x50 | Atellica IM BR CAL (SMN 10995515) Atellica IM BR PRE (SMN 10995480) | Atellica IM Multi‑Diluent 1 (SMN 10995637, 2-pack; SMN 10995638, 6-pack; SMN 10995639, vial) Atellica IM BR MCM (SMN 10995479) | 14 | 28 | Mouse monoclonal anti-CA 27.29 antibody | ;#Serum;#EDTA Plasma;# | < 8 hrs at RT > 8 hrs at 2–8°C >48 hrs at ≤ -20°C | 25 μL | Yes |  | Competitive immunoassay using direct chemiluminescent technology | 3.5–450.0 U/mL | NA |  |  |  |  | < 38.6 U/mL | See IFU for other results for other conditions |  | Atellica IM BR CAL | 10995515 (2‑pack) | 28 days recon at 2–8°C 32 hours recon at RT | Lyophilized | 2.0 mL of special reagent water | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of CA 27.29 (human) | 32 | 10 |  |  |  | Yes | Atellica IM Multi‑Diluent 1 | 10995637 (2‑pack) 10995638 (6‑pack) 10995639 (vial) |  | 3.5–9000 U/mL | NA | No | NA | NA | < 0.1 U/mL | NA | NA | NA | 4.1 U/mL | 1.0 U/mL | Not Performed | NA | NA | NA | 31.2 U/mL (Control) | NA | 4.5 | 7.3 | 154.6 U/mL (Control) | NA | 1.7 | 3.3 | 317.1 U/mL | NA | 2.1 | 4.1 | ADVIA Centaur BR y = 1.02x + 0.1 U/mL r=1.00 | NA | ≤ 5% change up to 500 mg/dL hemoglobin | ≤ 5% change up to 20 mg/dL conjugated bilirubin, 20 mg/dL unconjugated bilirubin | ≤ 5% change up to 3000 mg/dL Triglycerides | No Biotin in assay formulation | NA | NA |  |  | Traceable to an internal standard manufactured using highly purified CA 27.29 | 10995282_EN Rev. 03, 2019-08 | IFU Document | 0.713 | 6.969 | 2.087 | 4.37 | 0.324 | 17.7 | 5.3 | 11.1
Atellica® IM | 10995649 |  | ca07 | 00630414599670 | Global | Cardiac | No | Atellica IM Myoglobin (MYO) | Atellica IM MYO | MYO | 50 | 1 X 100 | Atellica IM CAL U (SMN 10995519) | Atellica IM Multi‑Diluent 10 (SMN 10995640) Atellica IM MYO MCM (SMN 10995650) | 10 | 28 | Goat polyclonal anti-myoglobin antibody / mouse monoclonal anti-myoglobin antibody | ;#Serum;#Heparinized Plasma;# | Up to 8 hrs at RT, 48hr 2-8C, freeze 1 time | 10 μL | No |  | direct chemiluminescent technology | 3.00–1000.00 ng/mL | 3.00–1000.00 μg/L |  |  |  |  | A myoglobin value of >110 ng/mL (ug/L) is highly suggestive of myocardial infarct |  |  | Atellica IM CAL U | 10995519 (2‑pack) | 4 hrs at RT, 24 hrs recon at 2–8°C, 60 days recon at -80 to -60°C | Lyophilized | reconstitute 2.0mL / Vial | 3 for lot calibration, 2 for pack recalibration | 2 |  | 33 | 14 |  |  |  | Yes | Atellica IM Multi‑Diluent 10 | 10995640 |  | 3.00–20,000 ng/mL | 3.00-20,000.00 μg/L | No | NA | NA | 2.53 ng/mL | 2.53 μg/L | NA | NA | 3.57 ng/mL | 2.56 ng/mL | NA | 2.56 μg/L | 3.57 μg/L | NA | 155.95 mg/dL (Control) | 155.95 μg/L (Control) | 2 | 2.9 | 361.51 mg/dL (Control) | 361.51 μg/L (Control) | 2.1 | 3 | 832.09 mg/dL | 832.09 μg/L | 2.2 | 2.9 | ADVIA Centaur Myoglobin y = 0.96x + 2.80 ng/mL r=1.00 | ADVIA Centaur Myoglobin y = 0.96x + 2.80 μg/L r=1.00 | ≤ 5% change up to 1000 mg/dL of hemoglobin | ≤ 5% change up to 40 mg/dL of conjugated bilirubin 40 mg/dL of unconjugated bilirubin | ≤ 5% change up to 1000 mg/dL of triglycerides | No Biotin in assay formulation | 200,000 ng/mL | 200,000 μg/L |  |  | Traceable to an internal standard manufactured using highly purified material | 10995411_EN Rev. 02, 2019-07 | IFU Document | 0.176 | 6.969 | 2.087 | 1.457 | 0.08 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995555 |  | re09 | 00630414598796 | Global | Reproductive Endocrinology | No | Atellica IM DHEA‑SO4 (DHEAS) | Atellica IM DHEAS | DHEAS | 50 | 1 x 50 | NA | Atellica IM Multi‑Diluent 1 Atellica IM DHEAS MCM (SMN 10995556) | 14 | 16 | DHEAS (~280 pg/mL) labeled with acridinium ester, Biotinylated mouse monoclonal anti-DHEAS antibody | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | RT up to 4 hours > 4 hours, at 2–8°C > 6 days, at ≤ -20°C for 1 month, Freeze 1 time | 25 μL | No |  | Quantitative competitive immunoassay | 3.00–1500.00 μg/dL | 0.08–40.71 μmol/L |  |  |  |  | See IFUs |  |  | Atellica IM DHEAS CAL | Provided in Kit | at 2–8°C until expiration 4 hours at RT | Liquid |  | 3 for lot calibration, 2 for pack recalibration | 2 | DHEAS; human serum | 70 | 14 |  |  |  | Yes | Atellica IM Multi‑Diluent 1 | 10995637 (2‑pack) 10995638 (6‑pack) 10995639 (vial) |  | 3.00–3000.00 μg/dL | 0.08–81.42 μmol/L | No | NA | NA | NA | NA | NA | NA | 4.27 μg/dL | 2.35 μg/dL | NA | 0.06 μmol/L | 0.12 μmol/L | NA | 74.56 μg/dL | 2.02 μmol/L | 2.3 | 5.6 | 636.71 μg/dL | 17.28 μmol/L | 2.2 | 4.4 | 1236.97 μg/dL | 33.57 μmol/L | 2.5 | 4.5 | ADVIA Centaur DHEAS (x) y = 1.07x - 1.48 μg/dL r=0.99 | ADVIA Centaur DHEAS y = 1.07x - 0.04 μmol/L r=0.99 | ≤ 10% change up to 500 mg/dL hemoglobin | ≤ 10% change up to 5 mg/dL conjugated bilirubin 5 mg/dL unconjugated bilirubin | ≤ 10% change up to 1000 mg/dL triglycerides | ≤ 10% change in results up to 12.5 ng/mL of biotin | NA | NA |  |  | Traceable to an internal standard manufactured using purified DHEAS | 10995340_EN Rev. 04, 2019-07 (OUS); 11200382_EN Rev. 04, 2019-06 (US) | IFU Document | 0.317 | 6.969 | 2.48 | 2.913 | 0.144 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995591 |  | on09 | 00630414599151 | Global | Oncology | No | Atellica IM HER‑2/neu (H2n) | Atellica IM H2n | H2n | 50 | 1 x 50 | Atellica IM H2n CAL (SMN 10995592, 2-pack) | Atellica IM H2n QC (SMN 10995494) Atellica IM Multi‑Diluent 1 (SMN 10995637, 2-pack; SMN 10995638, 6-pack; SMN10995639, vial) Atellica IM H2n MCM (SMN 10995593) | 14 | 41 | Mouse monoclonal anti-human HER‑2/neu antibody | ;#Serum;# | Up to 8 hrs at RT > 8 hrs at 2–8°C >24 hrs at ≤ -20°C | 20 μL | No |  | 2-site sandwich immunoassay using direct chemiluminescent technology | 0.5–350.0 ng/mL | NA |  |  |  |  | ULN defined as the 95th percentile of the observed results, was 15.2 ng/mL. | See IFU for information on monitoring metastatic breast cancer | HER‑2/neu Changes and Clinical Status | Atellica IM H2n CAL | 10995592 (2‑pack) | 60 days open at 2–8°C 8 hours at RT | Liquid | N/A | 3 for lot calibration, 2 for pack recalibration | 2 | Low or high levels of HER‑2/neu | 58 | 14 |  |  |  | Yes | Atellica IM Multi‑Diluent 1 | 10995637 (2‑pack) 10995638 (6‑pack) 10995639 (vial) |  | 0.5–7000.0 ng/mL | NA | Yes | Atellica IM H2n QC | 10995594 | < 0.1 ng/mL | NA | NA | NA | 0.0 ng/mL | < 0.1 ng/mL | NA | NA | NA | NA | 15.2 ng/mL (Control) | NA | 2.1 | 2.9 | 97.8 ng/mL | NA | 1.7 | 2.9 | 333.9 ng/mL | NA | 1.3 | 3 | ADVIA Centaur H2n y = 0.98x - 0.11ng/mL r=1.00 | NA | Insignificant effect up to 1000 mg/dL hemoglobin | Insignificant effect up to 25 mg/dL conjugated bilirubin, 25 mg/dL unconjugated bilirubin | Insignificant effect up to 900 mg/dL Triglycerides | No Biotin in assay formulation | 20,000 ng/mL | NA |  |  | Traceable to an internal standard manufactured using highly purified material | 10995349_EN Rev. 02, 2019-07 | IFU Document | 0.183 | 6.969 | 2.087 | 1.457 | 0.083 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995548 |  | is01 | 00630414598727 | Global | Immunosupressant Drugs | No | Atellica IM Cyclosporine (CsA) | Atellica IM CsA | CsA | 50 |  | Atellica IM CsA CAL Atellica IM CsA PRE (10995552) | Atellica IM Multi‑Diluent 12 Atellica IM CsA MCM (10995551) | 14 | 42 |  | ;#Whole blood;# | 6hr at RT; 7 days at 2-8C; frozen at -20C up to 1 month | 100 μL | Yes |  |  | 30.00–1500.00 ng/mL | 24.95–1247.25 nmol/L |  |  |  |  | See IFU |  |  | Atellica IM CsA CAL | 10995549 | 21 weeks at 2–8°C 4 hours at RT | Liquid |  | 3 for lot calibration, 2 for pack recalibration | 2 |  | 32 | 28 |  |  |  | Yes | Atellica IM Multi‑Diluent 12 | 10995550 |  | 30.00–7500.00 ng/mL | 24.95–6236.25 nmol/L | No | NA | NA | NA | NA | 29.40 ng/mL | 24.45 nmol/L | 13.50 ng/mL | 8.60 ng/mL | NA | 7.15 nmol/L | 11.23 nmol/L | NA | 127.14 ng/mL | 105.72 nmol/L | 2.4 | 7.6 | 284.55 ng/mL | 236.60 nmol/L | 3.3 | 5.6 | 640.89 ng/mL | 532.90 nmol/L | 2.7 | 6.5 | ADVIA Centaur CsA y = 0.97x - 4.23 ng/mL r=0.99 | ADVIA Centaur CsA y = 0.97x - 3.52 nmol/L r=0.99 | NA | ≤ 10% change up to 60 mg/dL conjugated bilirubin 40 mg/dL unconjugated bilirubin | ≤ 10% change up to 900 mg/dL triglycerides 300 mg/dL cholesterol | ≤ 10% change up to 50 ng/mL | NA | NA |  |  | Traceable to an internal standard manufactured using highly purified cyclosporine (USP grade) | 10995338_EN Rev. 04, 2019-07 | IFU Document | 0.196 | 6.969 | 2.087 | 1.457 | 0.089 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995695 |  | li02 | 00630414600086 | OUS | Liver Fibrosis | No | Atellica IM Tissue Inhibitor of Matrix Metalloproteinase 1 (TIMP‑1) | Atellica IM TIMP-1TIMP-1 | TIMP-1 | 50 | 1 x 50 |  | N/A | 14 | 60 | Mouse monoclonal antibodies (MAb) | ;#Serum;# | 48 hr at RT, 7 days at 2-8C. Freeze samples at ≤ -20°C if the sample is not assayed within 7 days | 25 μL | No |  | 2-site sandwich immunoassay using direct chemiluminescent technology | 3.5–1300.0 ng/mL | 3.5–1300.0 ng/mL | 3.5–6500.0 ng/mL | 3.5–6500.0 ng/mL | Refer to the Atellica IM ELF Test instructions for use for ELF score expected values | Refer to the Atellica IM ELF Test instructions for use for ELF score expected values |  |  |  | Atellica IM ELF CAL | 10995565 | 8 hours at RT; 5 days at 2–8°C; 90 days at -20°C | Lyophilized | Add 2.0 mL of reagent water into each calibrator vial, Let the calibrators stand for 15–20 minutes at RT, Gently swirl | 3 for lot calibration, 2 for pack recalibration | 2 | low or high levels of HA, PIIINP (bovine), TIMP-1 with bovine serum albumin, buffer, sodium azide and preservatives | 64 | 60 |  |  |  | Yes | Atellica IM Multi‑Diluent 10 | 10995640 |  | 3.5–6500.0 ng/mL | 3.5–6500.0 ng/mL | Yes | Atellica IM ELF QC | 10995566 | NA | NA | NA | NA | 1.1 ng/mL | 0.4 ng/mL | NA | 0.4 ng/mL | 1.1 ng/mL | NA | 81.3 ng/mL | 81.3 ng/mL |  |  | 254.5 ng/mL | 254.5 ng/mL | 2.7 | 3.3 | 533.5 ng/mL | 533.5 ng/mL | 2.8 | 3.9 | ADVIA Centaur TIMP-1 y=0.94x + 4.6 ng/mL r=1.00; N=112 | ADVIA Centaur TIMP-1 y=0.94x + 4.6 ng/mL r=1.00; N=112 | < 10% change up to 500 mg/dL hemoglobin | < 10% change up to 40 mg/dL conjugated bilirubin 40 mg/dL unconjugated bilirubin | < 10% change up to 1000 mg/dL triglycerides | No Biotin in assay formulation | 30,000 ng/mL | 30,000 ng/mL | 30,000 ng/mL | 30,000 ng/mL | Traceable to an internal standard manufactured using highly purified material | 10995427_EN Rev. 02, 2019-08 | IFU Document | 0.202 | 6.969 | 2.087 | 1.457 | 0.092 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995653 |  | li03 | 00630414599717 | OUS | Liver Fibrosis | No | Atellica IM N‑terminal Propeptide of Type III Procollagen (PIIINP) | Atellica IM PIIINPPIIINP | PIIINP | 50 | 1 x 50 | Atellica IM PW3 (SMN 10995666) | N/A | 14 | 30 | Mouse monoclonal antibodies (MAb) | ;#Serum;# | 48 hr at RT, 7 days at 2-8C. Freeze samples at ≤ -20°C if the sample is not assayed within 7 days | 20 | No |  | 2-site sandwich immunoassay using direct chemiluminescent technology | 0.50–150.00 ng/mL | 0.50–150.00 ng/mL | 0.50–750.00 ng/mL | 0.50–750.00 ng/mL | Refer to the Atellica IM ELF Test instructions for use for ELF score expected values | Refer to the Atellica IM ELF Test instructions for use for ELF score expected values |  |  |  | Atellica IM ELF CAL | 10995565 | 8 hours at RT; 5 days at 2–8°C; 90 days at -20°C | Lyophilized | Add 2.0 mL of reagent water into each calibrator vial, Let the calibrators stand for 15–20 minutes at RT, Gently swirl | 3 for lot calibration, 2 for pack recalibration | 2 | low or high levels of HA, PIIINP (bovine), TIMP-1 with bovine serum albumin, buffer, sodium azide and preservatives | 63 | 30 |  |  |  | Yes | Atellica IM Multi‑Diluent 13 | 10995643 |  | 0.50–750.00 ng/mL | 0.50–750.00 ng/mL | Yes | Atellica IM ELF QC | 10995566 | NA | NA | NA | NA | 0.35 ng/mL | 0.08 ng/mL | NA | 0.08 ng/mL | 0.35 ng/mL | NA | 3.31 ng/mL | 3.31 ng/mL | 4.9 | 5.7 | 6.63 ng/mL | 6.63 ng/mL | 2.7 | 3.5 | 12.1 ng/mL | 12.1 ng/mL | 2.4 | 3.3 | ADVIA Centaur PIIINP y = 0.97x + 0.09 ng/mL ; r=1.00, N= 138 | ADVIA Centaur PIIINP y = 0.97x + 0.09 ng/mL ; r=1.00, N= 138 | < 10% change up to 500 mg/dL hemoglobin | < 10% change up to 40 mg/dL conjugated bilirubin 40 mg/dL unconjugated bilirubin | < 10% change up to 1000 mg/dL triglycerides | <10% change up to 50 ng/mL of biotin | 2500 ng/mL | 2500 ng/mL | 2500 ng/mL | 2500 ng/mL | Traceable to an internal standard manufactured using bovine PIIINP | 10995413_EN Rev. 03, 2019-07 | IFU Document | 0.2 | 6.969 | 2.087 | 1.457 | 0.091 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995595 |  | li01 | 00630414599199 | OUS | Liver Fibrosis | No | Atellica IM Hyaluronic Acid (HA) | Atellica IM HAHA | HA | 50 | 1 x 50 |  | N/A | 39 | 28 | Mouse monoclonal antibodies (MAb) | ;#Serum;# | 48 hr at RT, 7 days at 2-8C. Freeze samples at ≤ -20°C if the sample is not assayed within 7 days | 20 μL | No |  | 2-site sandwich assay using direct chemiluminescent technology | 1.60–1000.00 ng/mL | 1.60–1000.00 ng/mL | 1.60 – 5000.00 ng/mL | 1.60 – 5000.00 ng/mL | Refer to the Atellica IM ELF Test instructions for use for ELF score expected values | Refer to the Atellica IM ELF Test instructions for use for ELF score expected values |  |  |  | Atellica IM ELF CAL | 10995565 | 8 hours at RT; 5 days at 2–8°C; 90 days at -20°C | Lyophilized | Add 2.0 mL of reagent water into each calibrator vial, Let the calibrators stand for 15–20 minutes at RT, Gently swirl | 3 for lot calibration, 2 for pack recalibration | 2 | low or high levels of HA, PIIINP (bovine), TIMP-1 with bovine serum albumin, buffer, sodium azide and preservatives | 24 | 14 |  |  |  | Yes | Atellica IM Multi‑Diluent 13 | 10995643 |  | 1.60 – 5000.00 ng/mL | 1.60 – 5000.00 ng/mL | Yes | Atellica IM ELF QC | 10995566 | NA | NA | NA | NA | 1.10 ng/mL | 0.56 ng/mL | NA | 0.56 ng/mL | 1.10 ng/mL | NA | 21.66 ng/mL | 21.66 ng/mL | 3.1 | 4.3 | 54.92 ng/mL | 54.92 ng/mL | 3 | 3.8 | 194.17 ng/mL | 194.17 ng/mL | 3 | 4.6 | ADVIA Centaur HA y = 0.98x + 1.89 ng/mL r=1.00, N= 128 | ADVIA Centaur HA y = 0.98x + 1.89 ng/mL r=1.00, N= 128 | < 10% change up to 500 mg/dL hemoglobin | < 10% change up to 40 mg/dL conjugated bilirubin 40 mg/dL unconjugated bilirubin | < 10% change up to 1000 mg/dL triglycerides | No Biotin in assay formulation | 25,000 ng/mL | 25,000 ng/mL | 25,000 ng/mL | 25,000 ng/mL | Traceable to an internal standard manufactured using highly purified material | 10995350_EN Rev. 02, 2019-07 | IFU Document | 0.198 | 6.969 | 2.087 | 1.457 | 0.09 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995672 |  | re08 | 00630414599854 | Global | Reproductive Endocrinology | No | Atellica IM Sex Hormone Binding Globulin (SHBG) | Atellica IM SHBG | SHBG | 50 | 1 x 50 | Atellica IM SHBG CAL Atellica IM Multi‑Diluent 1 | Atellica IM Multi‑Diluent 1 (SMN 10995639) Atellica IM SHBG MCM (SMN 10995674) | 14 | 60 | Mouse monoclonal anti-SHBG antibody, | ;#Serum;#Lithium Heparin Plasma;# | RT up to 4 hours > 4 hours at 2–8°C > 6 days at ≤ -20°C for 1 month Freeze 1 time | 10 μL | No |  | Sandwich immunoassay | 1.60–180.00 nmol/L | 0.15–17.10 μg/mL |  |  |  |  | See IFU |  |  | Atellica IM SHBG CAL | 10995673(2‑pack) | 60 days recon at ≤ -20°C thaw 1 time | Lyophilized | Add 2.0 mL of special reagent water into each vial. Let the vials stand for 15–20 minutes at room temperature | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of SHBG (human); equine serum | 58 | 35 |  |  |  | Yes | Atellica IM Multi‑Diluent 1 | 10995637(2‑pack) 10995638(6‑pack) 10995639(vial) |  | 1.60–360.00 nmol/L | 0.15–34.20 μg/mL | No | NA | NA | NA | NA | NA | NA | 1.11 nmol/L | 0.59 nmol/L | 1.11 nmol/L | 0.06 μg/mL | 0.11 μg/mL | 0.11 μg/mL | 18.95 nmol/L | 1.80 μg/mL | 3.1 | 3.6 | 67.85 nmol/L | 6.45 μg/mL | 3.3 | 7 | 138.93 nmol/L | 13.20 μg/mL | 3.4 | 6.6 | ADVIA Centaur SHBG y = 0.99x - 1.93 nmol/L r=0.99 | ADVIA Centaur SHBG y = 0.99x - 0.18 μg/mL r=0.99 | < 10% change up to 500 mg/dL hemoglobin | < 10% change up to 20 mg/dL conjugated bilirubin 20 mg/dL unconjugated bilirubin | < 10% change up to 1000 mg/dL triglycerides | ≤ 10% change in results up to 250 ng/mL biotin | 1000 nmol/L | 95 μg/mL |  |  | Standardized to World Health Organization (WHO) 2nd International Standard for SHBG, NIBSC code 08/266 | 10995421_EN Rev. 04, 2019-08 (OUS) 11200388_EN Rev. 04, 2019-06 (US) | IFU Document | 0.174 | 6.969 | 2.087 | 1.457 | 0.079 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995719 |  | bo01 | 00630414600321 | OUS | Bone Metabolism | No | Atellica IM Vitamin D Total (VitD) | Atellica IM VitD | VitD | 100 |  | NA | Atellica IM VitD QC (10995724) Atellica IM VitD DIL (10995721) Atellica IM VitD MCM (10995723) | 14 | 28 | Mouse monoclonal anti-25(OH)vitamin D antibody | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#EDTA Plasma;# | < 24 hours at RT >24 hours at 2–8°C for up to 7 days &gt;7 days at ≤ -20°C | 20 μL | No |  | Competitive immunoassay using direct chemiluminescent technology | 4.20–150.00 ng/mL | 10.50–375.00 nmol/L |  |  |  |  | See IFU |  |  | Atellica IM VitD CAL | Provided in Kit | 28 days recon at 2–8°C 120 days recon at ≤ -20°C 10 hours recon at RT | Lyophilized | Reconstituted to 2.0 mL/vial | 3 for lot calibration, 2 for pack recalibration | 2 | Defibinated human plasma containing albumin | 28 | 28 |  |  |  | Yes | Atellica IM VitD DIL | 10995721 |  | 4.20–300.00 ng/mL | 10.50–750.00 nmol/L | Yes | Atellica IM VitD QC | 10995724 (3x2.0mL) | NA | NA | NA | NA | 2.73 ng/mL | 1.26 ng/mL | 2.73 ng/mL | 3.15 nmol/L | 6.83 nmol/L | 6.83 nmol/L | 24.76 ng/mL | 61.90 nmol/L | 3.4 | 6.9 | 64.33 ng/mL | 160.83 nmol/L | 3.7 | 6.2 | 122.18 ng/mL | 305.45 nmol/L | 2.2 | 2.8 | ADVIA Centaur VitD y = 1.05x + 2.16 ng/mL r=0.99 | ADVIA Centaur VitD y = 1.05x + 5.40 nmol/L r=0.99 | ≤ 10% change up to 155 mg/dL of hemoglobin | ≤ 10% change up to 40 mg/dL of conjugated bilirubin 40 mg/dL of unconjugated bilirubin | ≤ 10% change up to 540 mg/dL of triglycerides | No Biotin in assay formulation | NA | NA |  |  | Traceable to the ID LC/MS/MS 25(OH)vitamin D RMP | 10995438_EN Rev. 02, 2019-08 | IFU Document | 0.378 | 6.969 | 2.48 | 2.913 | 0.172 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995720 |  | bo01 | 00630414600338 | OUS | Bone Metabolism | No | Atellica IM Vitamin D Total (VitD) | Atellica IM VitD | VitD | 500 |  | NA | Atellica IM VitD QC (10995724) Atellica IM VitD DIL (10995721) Atellica IM VitD MCM ( SMN 10995723) | 14 | 28 | Mouse monoclonal anti-25(OH)vitamin D antibody | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#EDTA Plasma;# | < 24 hours at RT, >24 hours at 2–8°C for up to 7 days, >7 days at ≤ -20°C | 20 μL | No |  | Competitive immunoassay using direct chemiluminescent technology | 4.20–150.00 ng/mL | 10.50–375.00 nmol/L |  |  |  |  | See IFU |  |  | Atellica IM VitD CAL | Provided in Kit | 28 days recon at 2–8°C 120 days recon at ≤ -20°C 10 hours recon at RT | Lyophilized | Reconstitute to 2.0 mL/vial | 3 for lot calibration, 2 for pack recalibration | 2 | Defribrinated human plasma containing albumin | 28 | 28 |  |  |  | Yes | Atellica IM VitD DIL | 10995721 |  | 4.20–300.00 ng/mL | 10.50–750.00 nmol/L | Yes | Atellica IM VitD QC | 10995724 (3x2.0mL) | NA | NA | NA | NA | 2.73 ng/mL | 1.26 ng/mL | 2.73 ng/mL | 3.15 nmol/L | 6.83 nmol/L | 6.83 nmol/L | 24.76 ng/mL | 61.90 nmol/L | 3.4 | 6.9 | 64.33 ng/mL | 160.83 nmol/L | 3.7 | 6.2 | 122.18 ng/mL | 305.45 nmol/L | 2.2 | 2.8 | ADVIA Centaur VitD y = 1.05x + 2.16 ng/mL r=0.99 | ADVIA Centaur VitD y = 1.05x + 5.40 nmol/L r=0.99 | ≤ 10% change up to 155 mg/dL of hemoglobin | ≤ 10% change up to 40 mg/dL of conjugated bilirubin 40 mg/dL of unconjugated bilirubin | ≤ 10% change up to 540 mg/dL of triglycerides | No Biotin in assay formulation | NA | NA |  |  | Traceable to the ID LC/MS/MS 25(OH)vitamin D RMP | 10995438_EN Rev. 02, 2019-08 | IFU Document | 1.54 | 10.748 | 2.48 | 6.063 | 0.7 | 27.3 | 6.3 | 15.4
Atellica® IM | 10995712 |  | ty05 | 00630414600253 | Global | N/AThyroid | No | Atellica IM Thyroid Uptake (TUp) | Atellica IM TUpTUp | TUp | 50 | 1 x 50 | Atellica IM CAL A | NA | 14 | 4 | Acridinium ester labeled T3 -BBG (bovine gamma globulin) and unlabeled T3 | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | Up to 8 hrs at RT, Up to 48 hrs at 2-8°C; Frozen at <-20 °C | 20 | No |  | Competitive immunoasay using direct chemiluminescent technology | NA | NA | NA | NA | 0.75 - 1.23 TU Ratio | 0.75 - 1.23 TU Ratio | NA | NA | NA | Atellica IM CAL A | 10995500 10995501 | 28 days | Lyophilized | Add 5 mL of reagent grade water, let stand 15-20 minutes at RT , gently invert and swirl until homogeneous | 3 for lot calibration, 2 for pack recalibration | 2 | Human plasma | 33 | 4 |  |  |  | No | NA | NA |  | NA | NA | No | Use commercially available control material | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 1.14 T Uptake Ratio | 1.14 T Uptake Ratio | 2.1 | 3.7 | 1.44 T Uptake Ratio | 1.44 T Uptake Ratio | 1.6 | 3.4 | 1.79 T Uptake Ratio | 1.79 T Uptake Ratio | 1.4 | 3.2 | ADVIA Centaur TUp FTI y = 1.05x - 0.07 Index Units r = 0.98 | ADVIA Centaur TUp FTI y = 1.05x - 0.07 Index Units r = 0.98 | insignificant effect up to 1200 mg/dL of hemoglobin | insignificant effect up to 30 mg/dL of bilirubin | insignificant effect up to 2000 mg/dL of triglycerides | No Biotin in assay formulation | NA | NA | NA | NA | Internal Standard | 10995436_EN Rev. 04, 2020-06 | IFU Document | 0.202 | 6.969 | 2.087 | 1.457 | 0.092 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995711 |  | ty05 | 00630414600246 | Global | N/AThyroid | No | Atellica IM Thyroid Uptake (TUp) | Atellica IM TUpTUp | TUp | 250 | 5 x 50 | Atellica IM CAL A | NA | 14 | 4 | Acridinium ester labeled T3 -BBG (bovine gamma globulin) and unlabeled T3 | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | Up to 8 hrs at RT, Up to 48 hrs at 2-8°C; Frozen at <-20 °C | 20 | No |  | Competitive immunoassay using direct chemiluminescent technology | NA | NA | NA | NA | 0.75 - 1.23 TU Ratio | 0.75 - 1.23 TU Ratio | NA | NA | NA | Atellica IM CAL A | 10995500 10995501 | 28 days | Lyophilized | Add 5 mL of reagent grade water, let stand 15-20 minutes at RT , gently invert and swirl until homogeneous | 3 for lot calibration, 2 for pack recalibration | 2 | Human plasma | 33 | 4 |  |  |  | No | NA | NA |  | NA | NA | No | Use commercially available control material | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 1.14 T Uptake Ratio | 1.14 T Uptake Ratio | 2.1 | 3.7 | 1.44 T Uptake Ratio | 1.44 T Uptake Ratio | 1.6 | 3.4 | 1.79 T Uptake Ratio | 1.79 T Uptake Ratio | 1.4 | 3.2 | ADVIA Centaur TUp FTI y = 1.02x + 0.00 (ACS:180 Index Units) r = 0.99 | ADVIA Centaur TUp FTI y = 1.02x + 0.00 (ACS:180 Index Units) r = 0.99 | insignificant effect up to 1200 mg/dL of hemoglobin | insignificant effect up to 30 mg/dL of bilirubin | insignificant effect up to 2000 mg/dL of triglycerides | No Biotin in assay formulation | NA | NA | NA | NA | Internal Standard | 10995436_EN Rev. 04, 2020-06 | IFU Document | 0.812 | 6.969 | 2.087 | 4.37 | 0.369 | 17.7 | 5.3 | 11.1
Atellica® IM | 10995615 |  | in01 | 00630414599380 | Global | Inflamation | No | Atellica IM Total IgE (tIgE) | Atellica IM tlgE | tlgE | 50 |  | Atellica IM CAL 80 | Atellica IM tIgE DIL Atellica IM tIgE MCM (SMN 10995617) | 14 | 28 | Lite Reagent: Goat anti-human IgE antibody, Solid Phase: Mouse anti-human IgE antibody | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | Up to 8 hrs at RT, Up to 48 hrs at 2-8°C, Frozen at <-20 °C | 30 μL | No |  | 2-site sandwich immunoassay | 1.5–3000.0 IU/mL | NA |  |  |  |  | Observed range for normal samples was 0-378 IU/mL | See IFU for age related ranges |  | Atellica IM CAL 80 | 10995497 (2‑pack) | 60 days recon at 2–8°C, 8 hours recon at RT | Lyophilized | Reconstitue with 2.0mLof special reagent water. Let stand for 15-20 minutes at RT | 3 for lot calibration, 2 for pack recalibration | 2 | Human IgE: equine serum, preservatives | 50 | 28 |  |  |  | Yes | Atellica IM tIgE DIL | 10995616 |  | 1.5–15000.0 IU/mL | NA | No | NA | NA | NA | NA | NA | NA | 1.4 IU/mL | 1.0 IU/mL | NA | NA | NA | NA | 9.4 U/mL | NA | 2.6 | 3.5 | 329.2 IU/mL (Control) | NA | 1.4 | 2.1 | 2408.8 IU/mL | NA | 3.5 | 12.6 | ADVIA Centaur tIgE y = 1.04x + 0.48 IU/mL r=0.99 | NA | insignificant effect up to 500 mg/dL of hemoglobin | insignificant effect up to 20 mg/dL of bilirubin | insignificant effect up to 3000 mg/dL of triglycerides | No Biotin in assay formulation | 10,000 IU/mL | NA |  |  | traceable to the World Health Organization (WHO) International Reference Preparation (75/502) | 10995366_EN Rev. 02, 2020-02 | IFU Document | 0.207 | 6.969 | 2.087 | 1.457 | 0.094 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995614 |  | in01 | 00630414599373 | Global | Inflamation | No | Atellica IM Total IgE (tIgE) | Atellica IM tlgE | tlgE | 250 |  | Atellica IM CAL 80 | Atellica IM tIgE DIL Atellica IM tIgE MCM (SMN 10995617) | 14 | 28 | Lite Rgt: Goat anti-human IgE antibody, Solid Phase: Mouse anti-human IgE antibody | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | Up to 8 hrs at RT, Up to 48 hrs at 2-8°C, Frozen at <-20 °C | 30 μL | No |  | 2-site sandwich immunoassay | 1.5–3000.0 IU/mL | NA |  |  |  |  | Observed range for normal samples was 0–378 IU/mL | See IFU for age related ranges |  | Atellica IM CAL 80 | 10995497 (2‑pack) | 60 days recon at 2–8°C, 8 hours recon at RT | Lyophilized | Reconstitue with 2.0 mL of special reagent water. Let stand for 15–20 minutes at RT | 3 for lot calibration, 2 for pack recalibration | 2 | Human IgE;equine serum; preservatives | 50 | 28 |  |  |  | Yes | Atellica IM tIgE DIL | 10995616 |  | 1.5–15000.0 IU/mL | NA | No | NA | NA | NA | NA | NA | NA | 1.4 IU/mL | 1.0 IU/mL | NA | NA | NA | NA | 9.4 U/mL | NA | 2.6 | 3.5 | 329.2 IU/mL (Control) | NA | 1.4 | 2.1 | 2408.8 IU/mL | NA | 3.5 | 12.6 | ADVIA Centaur tIgE y = 1.04x + 0.48 IU/mL r=0.99 | NA | insignificant effect up to 500 mg/dL of hemoglobin | insignificant effect up to 20 mg/dL of bilirubin | insignificant effect up to 3000 mg/dL of triglycerides | No Biotin in assay formulation | 10,000 IU/mL | NA |  |  | traceable to the World Health Organization (WHO) International Reference Preparation (75/502) | 10995366_EN Rev. 02, 2020-02 | IFU Document | 0.827 | 6.969 | 2.087 | 4.37 | 0.376 | 17.7 | 5.3 | 11.1
Atellica® IM | 10995606 |  | me02 | 00630414599298 | Global | N/AMetabolic | No | Atellica IM Homocysteine (HCY) | Atellica IM HCYHCY | HCY | 100 | 1 x 100 | Atellica IM HCY CAL SMN 10995498 | Atellica IM HCY DIL SMN 10995498 Atellica IM HCY MCM SMN 10995609 | 14 | 28 | Mouse monoclonal anti-SAH labeled with acridinium ester | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | 48 hrs at 2-8°C, frozen at <-20 °C up to 13 weeks | 20 μL | No |  | Competitive immunoassay using direct chemiluminescent technology | 0.50–65.00 μmol/L | 0.50–65.00 μmol/L | 0.50– 260.00 μmol/L | 0.50– 260.00 μmol/L | 3.7 – 13.9 μmol/L | 3.7 – 13.9 μmol/L | NA | NA | NA | Atellica IM HCY CAL | 10995498 | 4 hours | Liquid | Ready to use Warm to RT and gently mix | 3 for lot calibration, 2 for pack recalibration | 2 | S-adenosylhomocysteine | 42 | 14 |  |  |  | Yes | Atellica IM HCY DIL | 10995608 |  | 0.50–130.00 μmol/L | 0.50–130.00 μmol/L | No | NA | NA | NA | NA | NA | NA | 0.67 μmol/L | < 0.01 μmol/L | NA | < 0.01 μmol/L | 0.67 μmol/L | NA | 5.53 μmol/L | 5.53 μmol/L | 2.9 | 3.3 | 11.55 μmol/L | 11.55 μmol/L | 1.6 | 2.7 | 26.38 μmol/L | 26.38 μmol/L | 1.4 | 2.5 | ADVIA Centaur HCY y = 1.06 + 0.22 μmol/L r=0.98 | ADVIA Centaur HCY y = 1.06 + 0.22 μmol/L r=0.98 | ≤ 10% change up to 1000 mg/dL of hemoglobin | ≤ 10% change up to 25.0 mg/dL of bilirubin | ≤ 10% change up to 1300 mg/dL of lipid | ≤ 10% change up to 3500 ng/mL of biotin | NA | NA | NA | NA | Internal Standard | 10995362_EN Rev. 03, 2020-11 | IFU Document | 0.341 | 6.969 | 2.48 | 2.913 | 0.155 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995607 |  | me02 | 00630414599304 | Global | N/AMetabolic | No | Atellica IM Homocysteine (HCY) | Atellica IM HCYHCY | HCY | 500 | 5 x 100 | Atellica IM HCY CAL SMN 10995498 | Atellica IM HCY DIL SMN 10995498 Atellica IM HCY MCM SMN 10995609 | 14 | 28 | Mouse monoclonal anti-SAH labeled with acridinium ester | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | 48 hrs at 2-8°C,; frozen at <-20 °C up to 13 weeks | 20 μL | No |  | Competitive immunoassay using direct chemiluminescent technology | 0.50–65.00 μmol/L | 0.50–65.00 μmol/L | 0.50– 260.00 μmol/L | 0.50– 260.00 μmol/L | 3.7 – 13.9 μmol/L | 3.7 – 13.9 μmol/L | NA | NA | NA | Atellica IM HCY CAL | 10995498 | 4 hours | Liquid | Ready to use Warm to RT and gently mix | 3 for lot calibration, 2 for pack recalibration | 2 | S-adenosylhomocysteine | 42 | 14 |  |  |  | Yes | Atellica IM HCY DIL | 10995608 |  | 0.50–130.00 μmol/L | 0.50–130.00 μmol/L | No | NA | NA | NA | NA | NA | NA | 0.67 μmol/L | < 0.01 μmol/L | NA | < 0.01 μmol/L | 0.67 μmol/L | NA | 5.53 μmol/L | 5.53 μmol/L | 2.9 | 3.3 | 11.55 μmol/L | 11.55 μmol/L | 1.6 | 2.7 | 26.38 μmol/L | 26.38 μmol/L | 1.4 | 2.5 | ADVIA Centaur HCY y = 1.06 + 0.22 μmol/L r=0.98 | ADVIA Centaur HCY y = 1.06 + 0.22 μmol/L r=0.98 | ≤ 10% change up to 1000 mg/dL of hemoglobin | ≤ 10% change up to 25.0 mg/dL of bilirubin | ≤ 10% change up to 1300 mg/dL of lipid | ≤ 10% change up to 3500 ng/mL of biotin | NA | NA | NA | NA | Internal Standard | 10995362_EN Rev. 03, 2020-11 | IFU Document | 1.346 | 10.748 | 2.48 | 6.063 | 0.612 | 27.3 | 6.3 | 15.4
Atellica® IM | 10995558 |  | th03 | 00630414598826 | Global | Therapeutic Drug Monitoring | No | Atellica IM Digoxin (Dig) | Atellica IM Dig | Dig | 50 |  | Atellica IM CAL B (10995503 or 10995504) | Atellica IM Dig MCM (10995559) | 10 | 28 | Mouse monoclonal anti-digoxin antibody | ;#Serum;# | Up to 8 hrs at RT | 50 μL | No |  | Competitive immunoassay using direct chemiluminescent technology. | 0.10–5.00 ng/mL | 0.13–6.40 nmol/L |  |  |  |  | 0.80–2.00 ng/mL (1.02–2.56 nmol/L) |  |  | Atellica IM CAL B | 10995503 (2‑pack) 10995504 (6‑pack) | 28 days recon at 2–8°C 4 hours at RT | Lyophilized |  | 3 for lot calibration, 2 for pack recalibration | 2 |  | 51 | 28 |  |  |  | No | NA | NA |  | NA | NA | No | NA | NA | < 0.01 ng/mL | < 0.01 nmol/L | NA | NA | 0.27 ng/mL | 0.10 ng/mL | NA | 0.13 nmol/L | 0.35 nmol/L | NA | 0.94 ng/mL | 1.20 nmol/L | 2.5 | 3.7 | 2.10 ng/mL | 2.69 nmol/L | 1.4 | 2.4 | 3.83 ng/mL | 4.9 nmol/L | 1.9 | 2.6 | ADVIA Centaur DIG y = 1.02x - 0.08 ng/mL r=1.00 | ADVIA Centaur DIG y = 1.02x - 0.10 nmol/L r=1.00 | insignificant effect up to 500 mg/dL of hemoglobin | insignificant effect up to 20 mg/dL of bilirubin | insignificant effect up to 1000 mg/dL of triglycerides | No Biotin in assay formulation | NA | NA |  |  | traceable to an internal standard manufactured using U.S.P. (United States Pharmacopeia) material | 10995341_EN Rev. 02, 2020-02 | IFU Document | 0.178 | 6.969 | 2.087 | 1.457 | 0.081 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995557 |  | th03 | 00630414598819 | Global | Therapeutic Drug Monitoring | No | Atellica IM Digoxin (Dig) | Atellica IM Dig | Dig | 250 |  | Atellica IM CAL B | Atellica IM Dig MCM (10995559) | 10 | 28 | Mouse monoclonal anti-digoxin antibody | ;#Serum;# | Up to 8 hrs at RT | 50 μL | No |  | Competitive immunoassay using direct chemiluminescent technology | 0.10–5.00 ng/mL | 0.13–6.40 nmol/L |  |  |  |  | 0.80–2.00 ng/mL (1.02–2.56 nmol/L) |  |  | Atellica IM CAL B | 10995503 (2‑pack) 10995504 (6‑pack) | 28 days recon at 2–8°C 4 hours at RT | Lyophilized |  | 3 for lot calibration, 2 for pack recalibration | 2 |  | 51 | 28 |  |  |  | No | NA | NA |  | NA | NA | No | NA | NA | < 0.01 ng/mL | < 0.01 nmol/L | NA | NA | 0.27 ng/mL | 0.10 ng/mL | NA | 0.13 nmol/L | 0.35 nmol/L | NA | 0.94 ng/mL | 1.20 nmol/L | 2.5 | 3.7 | 2.10 ng/mL | 2.69 nmol/L | 1.4 | 2.4 | 3.83 ng/mL | 4.9 nmol/L | 1.9 | 2.6 | ADVIA Centaur DIG y = 1.02x - 0.08 ng/mL r=1.00 | ADVIA Centaur DIG y = 1.02x - 0.10 nmol/L r=1.00 | insignificant effect up to 500 mg/dL of hemoglobin | insignificant effect up to 20 mg/dL of bilirubin | insignificant effect up to 1000 mg/dL of triglycerides | No Biotin in assay formulation | NA | NA |  |  | traceable to an internal standard manufactured using U.S.P. (United States Pharmacopeia) material | 10995341_EN Rev. 02, 2020-02 | IFU Document | 0.678 | 6.969 | 2.087 | 4.37 | 0.308 | 17.7 | 5.3 | 11.1
Atellica® IM | 10995554 |  | th04 | 00630414598789 | Global | Therapeutic Drug Monitoring | No | Atellica IM Digitoxin (Dgtn) | Atellica IM Dgtn | Dgtn | 50 |  | Atellica IM CAL Z | Atellica IM Multi‑Diluent 5 | 10 | 28 | Mouse monoclonal anti-digitoxin antibody | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | Up to 8 hrs at RT | 10 μL | No |  | competitive immunoassay using direct chemiluminescent technology | 1.5–80.0 ng/mL | 2.0–104.6 nmol/L |  |  |  |  | 10–30 ng/mL (13.1–39.2 nmol/L) |  |  | Atellica IM CAL Z | 10995521 (2‑pack) | 28 days recon at 2–8°C 8 hours at RT | Lyophilized |  | 3 for lot calibration, 2 for pack recalibration | 2 |  | 49 | 7 |  |  |  | Yes | Atellica IM Multi‑Diluent 5 | 10995647 |  | 1.5–160.0 ng/mL | 2.0–209.2 nmol/L | No | NA | NA | < 0.01 ng/mL | < 0.1 nmol/L | NA | NA | 2.3 ng/mL | 0.5 ng/mL | NA | 0.7 nmol/L | 3.0 nmol/L | NA | 11.6 ng/mL | 15.1 nmol/L | 2.4 | 4.3 | 18.9 ng/mL | 24.7 nmol/L | 1.7 | 3.7 | 44.4 ng/mL | 58.1 nmol/L | 1.9 | 3.5 | ADVIA Centaur DGTN y = 0.99x + 0.26 ng/mL r=1.00 | ADVIA Centaur DGTN y = 0.99x + 0.34 nmol/L r=1.00 | insignificant effect up to 500 mg/dL of hemoglobin | insignificant effect up to 20 mg/dL of bilirubin | insignificant effect up to 1000 mg/dL of triglycerides | No Biotin in assay formulation | NA | NA |  |  | traceable to an internal standard manufactured using USP (United States Pharmacopeia) material | 10995339_EN Rev. 02, 2019-07 | IFU Document | 0.198 | 6.969 | 2.087 | 1.457 | 0.09 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995553 |  | th04 | 00630414598772 | Global | Therapeutic Drug Monitoring | No | Atellica IM Digitoxin (Dgtn) | Atellica IM Dgtn | Dgtn | 250 |  | Atellica IM CAL Z | Atellica IM Multi‑Diluent 5 | 10 | 28 | Mouse monoclonal anti-digitoxin antibody | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | Up to 8 hrs at RT | 10 μL | No |  | competitive immunoassay using direct chemiluminescent | 1.5–80.0 ng/mL | 2.0–104.6 nmol/L |  |  |  |  | 10–30 ng/mL (13.1–39.2 nmol/L) |  |  | Atellica IM CAL Z | 10995521 (2‑pack) | 28 days recon at 2–8°C 8 hours at RT | Lyophilized |  | 3 for lot calibration, 2 for pack recalibration | 2 |  | 49 | 7 |  |  |  | Yes | Atellica IM Multi‑Diluent 5 | 10995647 |  | 1.5–160.0 ng/mL | 2.0–209.2 nmol/L | No | NA | NA | < 0.01 ng/mL | < 0.1 nmol/L | NA | NA | 2.3 ng/mL | 0.5 ng/mL | NA | 0.7 nmol/L | 3.0 nmol/L | NA | 11.6 ng/mL | 15.1 nmol/L | 2.4 | 4.3 | 18.9 ng/mL | 24.7 nmol/L | 1.7 | 3.7 | 44.4 ng/mL | 58.1 nmol/L | 1.9 | 3.5 | ADVIA Centaur DGTN y = 0.99x + 0.26 ng/mL r=1.00 | ADVIA Centaur DGTN y = 0.99x + 0.34 nmol/L r=1.00 | insignificant effect up to 500 mg/dL of hemoglobin | insignificant effect up to 20 mg/dL of bilirubin | insignificant effect up to 1000 mg/dL of triglycerides | No Biotin in assay formulation | NA | NA |  |  | traceable to an internal standard manufactured using USP (United States Pharmacopeia) material | 10995339_EN Rev. 02, 2019-07 | IFU Document | 0.689 | 6.969 | 2.087 | 4.37 | 0.313 | 17.7 | 5.3 | 11.1
Atellica® IM | 10995530 |  | ca03 | 00630414598574 | Global | Cardiac | No | Atellica IM Creatine Kinase MB (CKMB) | Atellica IM CKMB | CKMB | 100 |  | Atellica IM CKMB CAL (SMN 10995532) 2-pack | Atellica IM CKMB DIL (SMN 10995533) Atellica IM CKMB MCM (SMN 10995534) | 10 | 28 | mouse monoclonal | ;#Serum;#Heparinized Plasma;# | Up to 4 hrs at RT, 2-8C up to 48hr, 1 freeze thaw | 100 μL | No |  | direct chemiluminescent technology | 0.18–300.00 ng/mL | 0.002–3.75 nmol/L |  |  |  |  | CKMB results from 233 apparently healthy individuals gave a median result of 0.78 ng/mL (0.0098 nmol/L) | Serum CKMB results from 167 hospitalized patients with non-cardiac related disorders gave a median result of 1.49 ng/mL (0.0186 nmol/L) | Serum CKMB results from 42 patients with confirmed myocardial injury ranged up to 144 ng/mL (1.80 nmol/L), with a median result of 25.4 ng/mL (0.3175 nmol/L) | Atellica IM CKMB CAL | 10995532 (2‑pack) | 14 days recon at 2–8°C, 4 hrs at RT | Lyophilized | Reconstitued to 2.0 mL/vial | 3 for lot calibration, 2 for pack recalibration | 2 | Human source material | 66 | 28 |  |  |  | Yes | Atellica IM CKMB DIL | 10995533 |  | 0.18–3000.00 ng/mL | 0.0023–37.5 nmol/L | No | NA | NA | 0.08 ng/mL | 0.001 nmol/L | NA | NA | 0.31 ng/mL | 0.13 ng/mL | NA | 0.002 nmol/L | 0.004 nmol/L | NA | 2.18 | 0.027 | 3.8 | 7.3 | 12.48 | 0.156 | 1.1 | 2.4 | 63.31 | 0.791 | 1 | 2.2 | ADVIA Centaur CKMB y = 1.03x + 0.25 ng/mL r=1.00 | ADVIA Centaur CKMB y = 1.03x + 0.003 nmol/L r=1.00 | insignificant effect up to 150 mg/dL of hemoglobin | insignificant effect up to 40 mg/dL of bilirubin | insignificant effect up to 1000 mg/dL of triglycerides | No Biotin in assay formulation | 500 ng/mL | 6.25 nmol/L |  |  | traceable to an internal standard manufactured using highly purified material | 10995316_EN Rev. 02, 2019-07 | IFU Document | 0.205 | 6.969 | 2.087 | 1.457 | 0.093 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995531 |  | ca03 | 00630414598581 | Global | Cardiac | No | Atellica IM Creatine Kinase MB (CKMB) | Atellica IM CKMB | CKMB | 500 |  | Atellica IM CKMB CAL (SMN 10995532) 2-pack | Atellica IM CKMB DIL (SMN 10995533) Atellica IM CKMB MCM (SMN 10995534) | 10 | 28 | Mouse monoclonal anti-CK-MB antibody | ;#Serum;#Heparinized Plasma;# | Up to 4 hrs at RT, 2-8C up to 48hr, 1 freeze thaw | 100 μL | No |  | direct chemiluminescent technology | 0.18–300.00 ng/mL | 0.002–3.75 nmol/L |  |  |  |  | CKMB results from 233 apparently healthy individuals gave a median result of 0.78 ng/mL (0.0098 nmol/L) | Serum CKMB results from 167 hospitalized patients with non-cardiac related disorders gave a median result of 1.49 ng/mL (0.0186 nmol/L) | Serum CKMB results from 42 patients with confirmed myocardial injury ranged up to 144 ng/mL (1.80 nmol/L), with a median result of 25.4 ng/mL (0.3175) | Atellica IM CKMB CAL | 10995532 (2‑pack) | 14 days recon at 2–8°C 4 hrs at RT | Lyophilized | Reconstituted to 2.0 mL/vial | 3 for lot calibration, 2 for pack recalibration | 2 | Human source material | 66 | 28 |  |  |  | Yes | Atellica IM CKMB DIL | 10995533 |  | 0.18–3000.00 ng/mL | 0.0023–37.5 nmol/L | No | NA | NA | 0.08 ng/mL | 0.001 nmol/L | NA | NA | 0.31 ng/mL | 0.13 ng/mL | NA | 0.002 nmol/L | 0.004 nmol/L | NA | 2.18 | 0.027 | 3.8 | 7.3 | 12.48 | 0.156 | 1.1 | 2.4 | 63.31 | 0.791 | 1 | 2.2 | ADVIA Centaur CKMB y = 1.03x + 0.25 ng/mL r=1.00 | ADVIA Centaur CKMB y = 1.03x + 0.003 nmol/L r=1.00 | insignificant effect up to 150 mg/dL of hemoglobin | insignificant effect up to 40 mg/dL of bilirubin | insignificant effect up to 1000 mg/dL of triglycerides | No Biotin in assay formulation | 500 ng/mL | 6.25 nmol/L |  |  | traceable to an internal standard manufactured using highly purified material | 10995316_EN Rev. 02, 2019-07 | IFU Document | 0.818 | 6.969 | 2.087 | 4.37 | 0.372 | 17.7 | 5.3 | 11.1
Atellica® IM | 10995466 |  | ty07 | 00630414597935 | Global | N/AThyroid | No | Atellica IM Anti‑Thyroid Peroxidase (aTPO) | Atellica IM aTPOaTPO | aTPO | 100 | 1 x 100 | Atellica IM CAL O (SMN 10995516) | Atellica IM aTPO QC SMN 10995470 Atellica IM aTPO DIL SMN 10995468 Atellica IM aTPO MCM SMN 10995469 | 14 | 28 | acridinium-ester-labeled- mouse monoclonal anti-TPO antibody | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | 8hr at room temperature, separate cells from serum, 48 hours at 2-8C, >48 hrs freeze at -20C once | 30 | No |  | Competitive immunoassay using direct chemiluminescent technology | 28–1300 U/mL | NA | 1:5 28 - 6500 1:10 28 - 13,000 | 1:5 28 - 6500 1:10 28 - 13,000 | Cutoff 60 U/mL | Cutoff 60 U/mL | NA | NA | NA | Atellica IM CAL O | 10995516 | 8 hours | Liquid | Let stand 5 minutes at room temperature, Gently mix | 3 for lot calibration, 2 for pack recalibration | 2 | Human plasma with antiTPO (human) | 28 | 14 |  |  |  | Yes | Atellica IM aTPO DIL | 10995468 |  | 1:5 28 - 6500 1:10 28 - 13,000 | 1:5 28 - 6500 1:10 28 - 13,000 | Yes | Atellica IM aTPO QC | 10995470 | NA | NA | NA | NA | 28 U/mL | 13 U/mL | NA | 13 U/mL | 28 U/mL | NA | 70 U/mL | 70 U/mL | 4.1 | 7.3 | 490 U/mL | 490 U/mL | 1.3 | 3.1 | 898 U/mL | 898 U/mL | 1.5 | 3.4 | ADVIA Centaur aTPO y = 1.03x - 11 U/mL r = 0.99 | ADVIA Centaur aTPO y = 1.03x - 11 U/mL r = 0.99 | ≤ 5% change up to 500 mg/dL of hemoglobin | ≤ 5% change up to 40 mg/dL of bilirubin | ≤ 5% change up to 1000 mg/dL of lipid | No Biotin in assay formulation | NA | NA | NA | NA | WHO Reference Preparation 66/387 | 10995280_EN Rev. 02, 2019-07 | IFU Document | 0.211 | 6.969 | 2.087 | 1.457 | 0.096 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995467 |  | ty07 | 00630414597942 | Global | N/AThyroid | No | Atellica IM Anti‑Thyroid Peroxidase (aTPO) | Atellica IM aTPOaTPO | aTPO | 500 | 5 x 100 | Atellica IM CAL O (SMN 10995516) | Atellica IM aTPO QC SMN 10995470 Atellica IM aTPO DIL SMN 10995468 Atellica IM aTPO MCM SMN 10995469 | 14 | 28 | acridinium-ester-labeled mouse monoclonal anti-TPO antibody | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | 8hr at room temperature, separate cells from serum, 48 hours at 2-8C, >48 hrs freeze at -20C once | 30 | No |  | Competitive immunoassay using direct chemiluminescent technology | 28–1300 U/mL | NA | 1:5 28 - 6500 1:10 28 - 13,000 | 1:5 28 - 6500 1:10 28 - 13,000 | Cutoff 60 U/mL | Cutoff 60 U/mL | NA | NA | NA | Atellica IM CAL O | 10995516 | 8 hours | Liquid | Let stand 5 minutes at room temperature, Gently mix | 3 for lot calibration, 2 for pack recalibration | 2 | Human plasma with antiTPO (human) | 28 | 14 |  |  |  | Yes | Atellica IM aTPO DIL | 10995468 |  | 1:5 28 - 6500 1:10 28 - 13,000 | 1:5 28 - 6500 1:10 28 - 13,000 | Yes | Atellica IM aTPO QC | 10995470 | NA | NA | NA | NA | 28 U/mL | 13 U/mL | NA | 13 U/mL | 28 U/mL | NA | 70 U/mL | 70 U/mL | 4.1 | 7.3 | 490 U/mL | 490 U/mL | 1.3 | 3.1 | 898 U/mL | 898 U/mL | 1.5 | 3.4 | ADVIA Centaur aTPO y = 1.03x - 11 U/mL r = 0.99 | ADVIA Centaur aTPO y = 1.03x - 11 U/mL r = 0.99 | ≤ 5% change up to 500 mg/dL of hemoglobin | ≤ 5% change up to 40 mg/dL of bilirubin | ≤ 5% change up to 1000 mg/dL of lipid | No Biotin in assay formulation | NA | NA | NA | NA | WHO Reference Preparation 66/387 | 10995280_EN Rev. 02, 2019-07 | IFU Document | 0.849 | 6.969 | 2.087 | 4.37 | 0.386 | 17.7 | 5.3 | 11.1
Atellica® IM | 10995701 |  | id16 | 00630414600147 | Global | ToRCH | No | Atellica IM Toxoplasma IgM (Toxo M) | Atellica IM Toxo M | Toxo M | 50 |  | NA | Atellica IM Toxo M QC | 39 | 19 | Immunoglobulin class capture sandwich immunoassay using direct chemiluminescent technology | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | at 2–8°C for up to 7 days at ≤ -20°C for up to 180 days | 10 μL | No |  | Mouse monoclonal anti-human IgM μ antibody | 0.10–40.00 Index | NA |  |  |  |  | Nonreactive, Reactive, Equivocal |  |  | Atellica IM Toxo M CAL | Provided in Kit | At 2–8°C Until expiration date 8 hrs at RT | Liquid |  | 3 for lot calibration, 2 for pack recalibration | 2 |  | 21 | 14 |  |  |  | No | NA | NA |  | NA | NA | Yes | Atellica IM Toxo M QC | 10995702 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 0.95 Index | NA | 2 | 3.2 | 4.0 Index | NA | 2.2 | 3.1 | 14.98 Index | NA | 4.6 | 5.5 | NA | NA | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 40 mg/dL of conjugated bilirubin up to 20 mg/dL of unconjugated bilirubin | ≤ 10% change up to 1000 mg/dL of triglycerides | No Biotin in assay formulation | NA | NA |  |  | Standardization is based upon the relative clinical agreement with commercially available toxoplasma IgM assays | 10995432_EN Rev. 02, 2019-08 (OUS) 10995433_EN Rev. 02, 2019-06 (US) | IFU Document | 0.257 | 6.969 | 2.48 | 2.913 | 0.117 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995699 |  | id15 | 00630414600123 | Global | ToRCH | No | Atellica IM Toxoplasma IgG (Toxo G) | Atellica IM Toxo G | Toxo G | 100 |  | Atellica IM APW1 (SMN 10995458) | Atellica IM Toxo G QC | 28 | 28 | Mouse monoclonal anti-human IgGFc antibody | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | at 2–8°C for up to 7 days after 7 days at ≤ -20°C | 10 μL | No |  | Immunoglobulin class capture sandwich immunoassay using direct chemiluminescent technology | 0.5–700.0 IU/mL | NA |  |  |  |  | Nonreactive, Reactive, Equivocal |  |  | Atellica IM Toxo G CAL | Provided in Kit | At 2–8°C Until expiration date 8 hrs at RT | Liquid |  | 3 for lot calibration, 2 for pack recalibration | 2 |  | 56 | 14 |  |  |  | No | NA | NA |  | NA | NA | Yes | Atellica IM Toxo G QC | 10995700 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 8.1 IU/mL | NA | 1.8 | 2.8 | 441.5 IU/mL | NA | 0.9 | 1.4 | 690.5 IU/mL | NA | 1.2 | 1.7 | NA | NA | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 60 mg/dL of conjugated bilirubin up to 40 mg/dL of unconjugated bilirubin | ≤ 10% change up to 1000 mg/dL of triglycerides | No Biotin in assay formulation | NA | NA |  |  | Traceable to World Health Organization (WHO) 3rd International standard for human anti-toxoplasma serum | 10995430_EN Rev. 02, 2019-08 (OUS) 10995433_EN Rev. 02, 2019-06 (US) | IFU Document | 0.284 | 6.969 | 2.48 | 2.913 | 0.129 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995675 |  | si01 | 00630414599885 | OUS | Special ID | No | Atellica IM Syphilis (Syph) | Atellica IM Syph | Syph | 200 |  | Atellica IM APW1 (SMN 10995458) | Atellica IM Syph QC | 29 | 60 | T. pallidum recombinant antigens | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#EDTA Plasma;#Citrate Plasma;# | RT not recommended 2–8°C for up to 7 days >7 days store at ≤ -20°C | 100 μL | No |  | Fully automated, antigen sandwich assay, using direct chemiluminescent technology | 0.10-45.00 index | NA |  |  |  |  | Nonreactive, Reactive, Equivocal |  |  | Atellica IM Syph CAL | Provided in Kit | At 2–8°C Until expiration date 8 hrs at RT | Liquid |  | 3 for lot calibration, 2 for pack recalibration | 2 |  | 37 | 21 |  |  |  | No | NA | NA |  | NA | NA | Yes | Atellica IM Syph QC | 10995676 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 0.95 Index (Plasma) | NA | 1.4 | 2.5 | 20.68 Index (Plasma) | NA | 1.2 | 3.5 | 32.22 Index (Plasma) | NA | 1.1 | 3.1 | NA | NA | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 40 mg/dL of conjugated bilirubin up to 40 mg/dL of unconjugated bilirubin | 500 ng/mL 1000 mg/dL of triglycerides | Insignificant effect on the assay up to 1500 ng/mL | NA | NA |  |  | Standardization is based upon relative clinical agreement with commercially available syphilis assays | 10995422_EN Rev. 04, 2019-08 | IFU Document | 0.341 | 6.969 | 2.48 | 2.913 | 0.155 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995668 |  | id12 | 00630414599830 | Global | ToRCH | No | Atellica IM Rubella IgM (Rub M) | Atellica IM Rub M | Rub M | 50 |  | Atellica IM APW1 (SMN 10995458) Atellica IM Multi‑Diluent 2 | Atellica IM Rub M QC | 40 | 28 | Mouse monoclonal anti‑human IgMμ antibody | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | at 2–8°C up to 7 days after 7 days at ≤ -20°C | 20 μL | No |  | Sandwich immunoassay using direct chemiluminescent technology | 0.00–15.00 Index | NA |  |  |  |  | Nonreactive, Reactive, Equivocal |  |  | Atellica IM Rub M CAL | Provided in Kit | At 2–8°C Until expiration date 8 hrs at RT | Liquid |  | 3 for lot calibration, 2 for pack recalibration | 2 |  | 91 | 14 |  |  |  | Yes | Atellica IM Multi‑Diluent 2 | 10995644 |  | ? | ? | Yes | Atellica IM Rub M QC | 10995669 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 0.92 Index | NA | 3.7 | 9.9 | 2.39 Index | NA | 2.9 | 4 | NA | NA |  |  | NA | NA | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 60 mg/dL of conjugated bilirubin up to 40 mg/dL of unconjugated bilirubin | ≤ 10% change up to 1000 mg/dL of triglycerides | No Biotin in assay formulation | NA | NA |  |  | Standardization is based upon relative clinical agreement with commercially available Rubella M assays | 10995419_EN Rev. 02, 2019-08 (OUS) 10995420_EN Rev. 02, 2019-06 (US) | IFU Document | 0.249 | 6.969 | 2.48 | 2.913 | 0.113 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995670 |  | id14 | 00630414600529 | Global | ToRCH | No | Atellica IM Rubella IgG (Rub G) | Atellica IM Rub G | Rub G | 100 |  | Atellica IM APW1 (10995458) | Atellica IM Rub G QC | 19 | 23 | Mouse monoclonal anti-human IgGFc antibody | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | at 2–8°C up to 7 days after 7 days at ≤ -20°C | 10 μL | No |  | Sandwich immunoassay using direct chemiluminescent technology | 0.2–500.0 IU/mL | NA |  |  |  |  | Nonreactive, Reactive, Equivocal |  |  | Atellica IM Rub G CAL | Provided in kit | At 2–8°C Until expiration date 8 hrs at RT | Liquid |  | 3 for lot calibration, 2 for pack recalibration | 2 |  | 53 | 14 |  |  |  | No | NA | NA |  | NA | NA | Yes | Atellica IM Rub G QC | 10995671 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 9.1 IU/mL | NA | 5.4 | 8.3 | 82.2 IU/mL (Control) | NA | 2 | 3.6 | 163.8 IU/mL | NA | 2.7 | 4.8 | NA | NA | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 60 mg/dL of conjugated bilirubin up to 40 mg/dL of unconjugated bilirubin | ≤ 10% change up to 1000 mg/dL of triglycerides | No Biotin in assay formulation | NA | NA |  |  | Traceable to the World Health Organization (WHO) 1st International Standard for human anti-rubella immunoglobulin (RUBI‑1‑94) | 10995417_EN Rev. 02, 2019-08 (OUS) 10995418_EN Rev. 02, 2019-06 (US) | IFU Document | 0.271 | 6.969 | 2.48 | 2.913 | 0.123 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995453 |  | id12 | 00630414597836 | OUS | Hepatitis | No | Atellica IM Hepatitis B surface antibody 2 (aHBs2) | Atellica IM aHBs2aHBs2 | aHBs2 | 200 | 1 x 200 | NA | Atellica IM aHBs2 QC Atellica IM Multi‑Diluent 11 | 14 | 90 | Inactivated human hepatitis B surface antigen | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | 2–8°C for up to 7 days ≤ -20°C for longer storage | 100 μL | No |  | direct sandwich | 3.1–1000.0 mIU/mL | 3.1–1000.0 mIU/mL | NA | NA | NA | NA | nonreactive <8.0 mIU/mL | reactive ≥ 12.0 mIU/mL | retest zone ≥8.0 mIU/mL and <12.0 mIU/mL | Atellica IM aHBs2 CAL | Provided in kit | At 2–8°C Until expiration date, 8 hrs at RT | Liquid | N/A | 3 for lot calibration, 2 for pack recalibration | 2 | processed human plasma | 55 | 42 |  |  |  | Yes | Atellica IM Multi‑Diluent 11 | 10995642 (2‑pack) 10995641 (vial) |  | 3.1 - 10,000 mIU/mL | N/A | Yes | Atellica IM aHBs2 QC | 10995454 | N/A | N/A | N/A | N/A | 3.1 mIU/mL | 1.7 mIU/mL | 9.0 mIU/mL | N/A | N/A | N/A | 10.3 mIU/mL | 10.3 mIU/mL | 5.1 | 5.3 | 131.1 mIU/mL | 131.1 mIU/mL | 1.7 | 3 | 797.3 mIU/mL | 797.3 mIU/mL | 2.3 | 3.7 | N/A | N/A | ≤ 15% change up to 500 mg/dL of hemoglobin | ≤ 15% change up to 40 mg/dL of conjugated and unconjugated bilirubin | ≤ 15% change up to 1000 mg/dL triglycerides | ≤15% change up to 1500 ng/mL of biotin | 200,000 mIU/mL | N/A |  |  | Standardized against the World Health Organization (WHO) 1st International Reference Preparation (1977) | 10995277_EN Rev. 04, 2019-07 | IFU Document | 0.299 | 6.969 | 2.48 | 2.913 | 0.136 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995538 |  | me01 | 00630414598659 | Global | N/AMetabolic | No | Atellica IM Cortisol (Cor) | Atellica IM CorCor | Cor | 50 | 1 x 50 | Atellica IM CAL E | Atellica IM Multi‑Diluent 3 Atellica IM Cor MCM (SMN 10995539) | 10 | 10 | Rabbit anti-cortisol antibody bound to mouse monoclonal anti-rabbit IgG antibody | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | Serum: 8hrs at RT; 48hrs at 2-8C; frozen at -20C | 20 μL | No |  | Competitive immunoassay using direct chemiluminescent | 0.50-75.00 μg/dL | 13.80–2069.25 nmol/L | 0.50–150.00 μg/dL | 13.80–4138.50 nmol/L | 7 - 9; (a.m.): 5.27–22.45 μg/dL 3 - 5 (p.m.) 3.44 - 16.76 μg/dL | 7–9 (a.m): 145.4–619.4 nmol/L 3 - 5 (p.m.): 94.9 - 462.4 nmol/L | NA | NA | NA | Atellica IM CAL E | 10995512 (2‑pack) 10995513 (6‑pack) | 14 days recon at 2–8°C 4 hours recons at RT | Lyophilized | Add 2.0 mL of special reagent water into each vial. Let vials stand for 15-20 minutes at room temperature. | 3 for lot calibration, 2 for pack recalibration | 2 | defibrinated human plasma | 19 | 10 |  |  |  | Yes | Atellica IM Multi‑Diluent 3 | 10995645 |  | 0.50–150.00 μg/dL | 13.80–4138.50 nmol/L | No | Use commercially available controls | NA | NA | NA | NA | NA | ≤ 0.14 μg/dL | ≤ 0.06 μg/dL | ≤ 0.31 μg/dL | 1.66 nmol/L | 3.86 nmol/L | 8.55 nmol/L | 3.11 μg/dL | 85.80 nmol/L | 2.1 | 8.6 | 18.82 μg/dL | 519.24 nmol/L | 1.8 | 7.8 | 31.57 μg/dL | 871.02 nmol/L | 1.5 | 7.6 | ADVIA Centaur Cortisol y = 1.02x + 0.14 μg/dL r=1.00 | ADVIA Centaur Cortisol y = 1.02x + 3.86 nmol/L r=1.00 | insignificant effect up to 500 mg/dL of hemoglobin | insignificant effect up to 20 mg/dL of conjugated bilirubin; insignificant effect up to 20 mg/dL of unconjugated bilirubin | insignificant effect up to 1170 mg/dL of triglycerides | No Biotin in assay formulation | NA | NA | NA | NA | Internal standards manufactured analytically which are traceable to gas chromatography-mass spectrometry (GCMS) | 10995336_EN Rev. 03, 2020-03 | IFU Document | 0.196 | 6.969 | 2.087 | 1.457 | 0.089 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995537 |  | me01 | 00630414600413 | Global | N/AMetabolic | No | Atellica IM Cortisol (Cor) | Atellica IM CorCor | Cor | 250 | 5 x 50 | Atellica IM CAL E | Atellica IM Multi‑Diluent 3 Atellica IM Cor MCM (SMN 10995539) | 10 | 10 | Rabbit anti-cortisol antibody bound to mouse monoclonal anti-rabbit IgG antibody | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | RT for up to 8 hrs >8hrs at 2–8°C >48hrs at ≤ -20°C | 20 μL | No |  | Competitive immunoassay using direct chemiluminescent | 0.50-75.00 μg/dL | 13.80–2069.25 nmol/L | 0.50–150.00 μg/dL | 13.80–4138.50 nmol/L | 7 - 9; (a.m.): 5.27–22.45 μg/dL 3 - 5 (p.m.) 3.44 - 16.76 μg/dL | 7–9 (a.m): 145.4–619.4 nmol/L 3 - 5 (p.m.): 94.9 - 462.4 nmol/L | NA | NA | NA | Atellica IM CAL E | 10995512 (2‑pack) 10995513 (6‑pack) | 14 days recon at 2–8°C, 4 hours recons at RT | Lyophilized | Add 2.0 mL of special reagent water into each vial. Let vials stand for 15-20 minutes at room temperature. | 3 for lot calibration, 2 for pack recalibration | 2 | defibrinated human plasma | 19 | 10 |  |  |  | Yes | Atellica IM Multi‑Diluent 3 | 10995645 |  | 0.50–150.00 μg/dL | 13.80–4138.50 nmol/L | No | Use commercially available controls | NA | NA | NA | NA | NA | ≤ 0.14 μg/dL | ≤ 0.06 μg/dL | ≤ 0.31 μg/dL | ≤1.66 nmol/L | ≤3.86 nmol/L | ≤8.55 nmol/L | 3.11 μg/dL | 85.80 nmol/L | 2.1 | 8.6 | 18.82 μg/dL | 519.24 nmol/L | 1.8 | 7.8 | 31.57 μg/dL | 871.02 nmol/L | 1.5 | 7.6 | ADVIA Centaur Cortisol y = 1.02x + 0.14 μg/dL r=1.00 | ADVIA Centaur Cortisol y = 1.02x + 3.86 nmol/L r=1.00 | insignificant effect up to 500 mg/dL of hemoglobin | insignificant effect up to 20 mg/dL of bilirubin; insignificant effect up to 20 mg/dL of unconjugated bilirubin | insignificant effect up to 1170 mg/dL of triglycerides | No Biotin in assay formulation | NA | NA | NA | NA | Internal standards manufactured analytically which are traceable to gas chromatography-mass spectrometry (GCMS) | 10995336_EN Rev. 03, 2020-03 | IFU Document | 0.682 | 6.969 | 2.087 | 4.37 | 0.31 | 17.7 | 5.3 | 11.1
Atellica® IM | 11200588 |  | ca05 | 00630414228945 | OUS | Cardiac | No | Atellica IM NT‑proBNP (PBNP) | Atellica IM PBNP | PBNP | 100 | 1 X 100 | NA | Atellica IM Multi‑Diluent 1 (SMN 10995637) Atellica IM PBNP MCM (SMN 11200590) | 14 | 28 | sheep monoclonal anti-human antibody | ;#Serum;#Lithium Heparin Plasma;#Potassium EDTA Plasma;# | Serum: 3 days at RT 8 days at 2–8°C Lithium heparin: 2 days at RT 8 days at 2–8°C Dipotassium EDTA: 4 days at RT 8 days at 2–8°C All samples: 12 months at ≤ -20°C | 20 μL | No |  | direct chemiluminescent technology | 35–35,000 pg/mL | 4.13–4130 pmol/L |  |  |  |  | 125 pg/mL for patients aged <75 years | 450 pg/mL for patients aged >75 years |  | IM PBNP CAL | NA | Lyophilized at 2–8°C until expiration date Recon at 2–8°C 24 hrs Recon at ≤ -20°C 30 days | Lyophilized | Reconstituted to 2.0 mL/vial | 3 for lot calibration, 2 for pack recalibration | 2 |  | 41 | 28 |  |  |  | Yes | Atellica IM Multi‑Diluent 1 | 10995637 (2‑pack) 10995638 (6‑pack) 10995639 (vial) |  | 35-350,000 pg/mL | 4.13-41,300 pmol/L | No | N | NA | NA | NA | NA | NA | 20 pg/mL | 13 pg/mL | 18 pg/mL | 1.53 pmol/L | 2.36 pmol/L | 2.12 pmol/L | 131 pg/mL (serum) | 16 pmol/L | 2.3 | 3.9 | 3202 pg/mL (serum) | 378 pmol/L | 1.2 | 3.2 | 31,794 pg/mL (serum) | 3,752 pmol/L | 1.8 | 3.7 | ADVIA Centaur PBNP y = 1.07x + 10 pg/mL r=1.00 | ADVIA Centaur PBNP y = 1.07x + 1.18 pmol/L r=1.00 | ≤ 10% change up to 1000 mg/dL of hemoglobin | ≤ 10% change up to 60 mg/dL of conjugated bilirubin up to 60 mg/dL of unconjugated bilirubin | ≤ 10% change up to 3000 mg/dL of Lipemia | ≤ 10% change in results up to 75 ng/mL. | 300,000 pg/mL | 35,400 pmol/L |  |  | Traceable to an internal standard manufactured using synthetic human NT‑proBNP | 11200584_EN Rev. 04, 2019-05 | IFU Document | 0.286 | 6.875 | 2.125 | 2.438 | 0.13 | 17.463 | 5.398 | 6.193
Atellica® IM | 11200589 |  | ca05 | 00630414228952 | OUS | Cardiac | No | Atellica IM NT‑proBNP (PBNP) | Atellica IM PBNP | PBNP | 500 | 5 X 100 | NA | Atellica IM Multi‑Diluent 1 (SMN 10995637) Atellica IM PBNP MCM (SMN 11200590) | 14 | 28 | sheep monoclonal anti-human antibody | ;#Serum;#Lithium Heparin Plasma;#Potassium EDTA Plasma;# | Serum: 3 days at RT 8 days at 2–8°C Lithium heparin: 2 days at RT 8 days at 2–8°C Dipotassium EDTA: 4 days at RT 8 days at 2–8°C All samples: 12 months at ≤ -20°C | 20 μL | No |  | direct chemiluminescent technology | 35–35,000 pg/mL | 4.13–4130 pmol/L |  |  |  |  | 125 pg/mL for patients aged <75 years | 450 pg/mL for patients aged >75 years |  | IM PBNP CAL | NA | Lyophilized at 2–8°C until expiration date Recon at 2–8°C 24 hrs Recon at ≤ -20°C 30 days | Lyophilized | Reconstitute 2.0 mL/vial | 3 for lot calibration, 2 for pack recalibration | 2 |  | 41 | 28 |  |  |  | Yes | Atellica IM Multi‑Diluent 1 | 10995637 (2‑pack) 10995638 (6‑pack) 10995639 (vial) |  | 35-350,000 pg/mL | 4.13 - 41,300 pmol/L | No | N | NA | NA | NA | NA | NA | 20 pg/mL | 13 pg/mL | 18 pg/mL | 1.53 pmol/L | 2.36 pmol/L | 2.12 pmol/L | 131 pg/mL (serum) | 16 pmol/L | 2.3 | 3.9 | 3,202 pg/mL (serum) | 378 pmol/L | 1.2 | 3.2 | 31,794 pg/mL (serum) | 3,752 pmol/L | 1.8 | 3.7 | ADVIA Centaur PBNP y = 1.07x + 10 pg/mL r=1.00 | ADVIA Centaur PBNP y = 1.07x + 1.18 pmol/L r=1.00 | ≤ 10% change up to 1000 mg/dL of hemoglobin | ≤ 10% change up to 60 mg/dL of conjugated bilirubin up to 60 mg/dL of unconjugated bilirubin | ≤ 10% change up to 3000 mg/dL of Lipemia |  | 300,000 pg/mL | 35,400 pmol/L |  |  | Traceable to an internal standard manufactured using synthetic human NT‑proBNP | 11200584_EN Rev. 04, 2019-05 | IFU Document | 0.953 | 6.875 | 2.125 | 6.75 | 0.433 | 17.463 | 5.398 | 17.145
Atellica® CH | 11097643 |  | pl19 | 00630414210827 | Global | Specific Proteins | No | Atellica CH Haptoglobin | HaptHapt | 4038 | 300 | 2 x 150 | N/A | N/A | 10 | 30 days per pack | Anti-human haptoglobin antibody (goat) | ;#Serum;# | Frozen at -20C up to 14 days | 3.4 μL | No |  | PEG-enhanced immunoturbidimetric | 1–340 mg/dL | 0.01–3.40 g/L | N/A | N/A | 40–280 mg/dL for adults | 0.40–2.80 g/L for adults | N/A | N/A | N/A | Atellica CH LSP CAL | 11099434 | 28 | Liquid | N/A | 3 | 6 | Proteins derived from human serum | 10 | 30 |  | No |  | No | Atellica CH Diluent | 11099300 |  | N/A | N/A | No | N/A | N/A | 1 mg/dL | 0 mg/dL | N/A | N/A | 1 mg/dL | 0 mg/dL | N/A | 0.00 g/L | 0.01 g/L | N/A | 80 (Control) | 0.80 (Control) | 0.7 | 2.2 | 125 (Control) | 1.25 (Control) | 0.7 | 2.2 | 214 (Control) | 2.14 (Control) | 1.2 | 2.9 | ADVIA Chemistry 1800 y = 1.06x + 0 mg/dL r=0.997 | ADVIA Chemistry 1800 y = 1.06x + 0.00 g/L r=0.997 | N/A | N/A | N/A | No Biotin in assay formulation | 600 mg/dL | 6.00 g/L | N/A | N/A | (IRMM) CRM-470 from IFCC | 11111580_EN Rev. 03, 2020-03 | IFU Document | 0.381 | 5.625 | 3.5 | 2.75 | 0.173 | 14.287 | 8.89 | 6.985
Atellica® CH | 10470440 |  | dr12 | 00842768034114 | Global | Drugs of Abuse/Toxicology | Yes | Emit II Plus 6-Acetylmorphine Assay (Alliance Application) | 6AM | 8053 | 200 | 2 x 100 |  |  | 6 | 28 | Mouse Mono-clonal Antibody to 6-AM | ;#Urine;# | ​7 days at 25°C, samples should be frozen after 7 days | 15 μL |  |  | EMIT | 10 ng/mL; Semi Quant 2.5 - 20.0 ng/mL | n/a | n/a | n/a | qualitative/semi-quantitative assay | n/a |  |  |  | Emit® Calibrator/Control Level 0 Emit® II Plus 6-AM/Ecstasy Calibrator/Control Level 1, 2, 3, 4 | 10445406 10470441 10470442 10470443 10470444 | expiration date printed on the vial label | Liquid |  | 3 | 4 |  | 60 | 6 |  | No |  | No | [12 empty] | n/a | n/a | n/a | n/a | n/a | n/a | 11.9 | n/a | 1.5 | 2.8 | 10.6 | n/a | 1.7 | 3.4 | 7.9 | n/a | 1.9 | 4.3 | Viva-E Analyzer ± 50% of 10 ng/mL cutoff |  | n/a | n/a | n/a | No Biotin in assay formulation |  |  | n/a | n/a | Traceable to Syva Emit Cal/Cont and Emit II Plus 6-AM Ecstasy which are gravimetrically prepared standards- qualified by GC/MS within ± 10% of nominal. | 9R022UL_US_E | IFU Document
Atellica® CH | 10445371 |  | th16 | 00842768001413 | Global | Therapeutic Drug Monitoring (TDM) | Yes | Emit Caffeine Assay (Alliance Application) | Caff | 8050 | 225 | 3 x 75 |  |  | 7 | 30 | Sheep antibodies reactive to caffeine | ;#Serum;# | ​Store at 2–8°C | 7 μL |  |  | EMIT | 1–30 µg/mL | 5.2–155 µmol/L | n/a | n/a | n/a | n/a |  |  |  | Emit Caffeine Calibrators (included in reagent kit) 0, 1, 3, 7, 15, 30 | Included | 84 | Lyophilized | Reconstitute 1.0 mL, 1 Hour | 3 | 6 | Human Serum | 60 | 8 |  | No |  | No | [12 empty] | n/a | n/a | n/a | n/a | n/a | n/a | 2.17 | n/a | 2.5 | 9.4 | 8.62 | n/a | 1.4 | 6 | 15.72 | n/a | 1.7 | 5.4 | Viva-E Analzyer, y = 0.96x + 0.9 μg/mL n = 102 r = 0.982 |  | n/a | n/a | n/a | No Biotin in assay formulation |  |  | n/a | n/a |  | 6P594UL.12DS_D | IFU Document
Atellica® CH | 10445389 |  | dr18 | 00842768014314 | Global | Drugs of Abuse/Toxicology | Yes | EMIT Tox Serum Tricyclic Antidepressants Assay (Alliance Application) | TCA | 8054 | 225 | 3 x 75 |  |  | 7 | 25 | Sheet Antibodies | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | ​Store at 2–8°C | 11 μL |  |  | EMIT | 300 ng/mL | n/a | [7 empty] | Emit tox Serum Tricyclic Antidepressants Cal/Control | 10445390 | 12 weeks | Lyophilized | Add 3 mL of distilled or deionized water to vial |  |  | Human serum | 25 | 8 |  | No |  | No | [18 empty] | 196.2 mAu/min |  | 0.3 | 0.9 | 229.6 mAu/min |  | 0.5 | 1.4 | 307.2 mAu/min |  | 0.4 | 1.2 | Viva-E Analyzer ± 50% of 300 ng/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation | [5 empty] | 7C244UL.2DS_F | IFU Document
Atellica® CH | 10445391 |  | dr17 | 00842768001604 | Global | Drugs of Abuse/Toxicology | Yes | EMIT Tox Serum Barbiturate Assay (Alliance Application) | sBrb | 8051 | 225 | 3 x 75 |  |  | 7 | 30 | Sheep Antibodies | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | ​Store at 2–8°C | 11 μL |  |  | EMIT | 3 µg/mL, Semi-quantitive 0.45-6 µg/mL | [8 empty] | Emit tox Serum Calibrators | 10445387 | 12 weeks | Lyophilized | Add 3 mL of distilled or deionized water to the vial. |  | 3 | Human Serum | 30 | 8 |  | No |  | No | [18 empty] | 0.92 µg/mL | 163.3 mAu/min | 0.4 | 1.2 | 3.19 µg/mL | 205.1 mAu/min | 0.5 | 1.1 | 5.83 µg/mL | 253.8 mAu/min | 0.6 | 1.2 | Viva-E Analyzer ± 50% of 3 µg/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation | [5 empty] | 7D164UL.12DS_E | IFU Document
Atellica® CH | 10445388 |  | dr16 | 00842768001581 | Global | Drugs of Abuse/Toxicology | Yes | EMIT Tox Serum Benzodiazepine Assay (Alliance Application) | sBnz | 8052 | 225 | 3 x 75 |  |  | 7 | 30 | Sheep Antibodies | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | ​​Store at 2–8°C | 11 μL |  |  | EMIT | 0.3 µg/mL, Semi-quantitive 0.07-2 µg/mL | [8 empty] | Emit tox Serum Calibrators | 10445387 | 12 weeks | Lyophilized | Add 3 mL of distilled or deionized water to the vial. |  |  | Human Serum | 30 | 8 |  | No |  | No | [18 empty] | 0.04 µg/mL | 155.8 mAu/min | 0.6 | 0.9 | 0.29 µg/mL | 176.3 mAu/min | 0.5 | 1.1 | 1.89 µg/mL | 322.3 mAu/min | 0.7 | 1.2 | Viva-E Analyzer ± 50% of 0.3 µg/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation | [5 empty] | 7B164UL.13DS_F | IFU Document
Atellica® CH | 10720048 |  | dr13 | 00630414601120 | Global | Drugs of Abuse/Toxicology | Yes | Emit II Plus Buprenorphine Assay (Alliance Application) | Bupn | 8049 | 200 | 2 x 100 |  |  | 6 | 28 | Mouse Monoclonal Antibodies | ;#Urine;# | 5 days at 25°C, samples should be frozen after 5 days | 15 μL |  |  | EMIT | 5 ng/mL; Semi Quant 2.5-25 ng/mL | [8 empty] | Emit® Calibrator/Control Level 0 Emit® II Plus Specialty Drug Calibrator/Control Level 1, 2, 3, 4 | 10445406 10720049 10720050 10720051 10720052 | expiration date printed on the vial | Liquid |  |  | 5 | Human Urine | 28 | 8 |  | No |  | No | [18 empty] | 2.93 ng/mL |  | 2.6 | 3.5 | 4.55 ng/mL |  | 1.4 | 2.8 | 6.83 ng/mL |  | 1.2 | 1.7 | Viva-E Analyzer ± 50% of 5 ng/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation | [5 empty] | 10872256_B | IFU Document
Atellica® CH | 10709280 |  | th17 | 00858724003056 | Global | Therapeutic Drug Monitoring (TDM) | Yes | ARK Methotrexate Assay (Alliance Application) | Mtx | 8048 | 110 | 1 x 110 |  |  | 7 | 62 | Rabbit Polyclonal Antibody to Methotrexate | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#Potassium EDTA Plasma;# | ​Clarified samples at 2-8°C up to 2 weeks, is testing delayed then store at less than -10°C | 20 μL |  |  | Homogenous Enzyme Immunoassay | 0.04 - 1.20 μmol/L | 0.02 - 0.55 μg/mL |  |  | ​See IFU |  |  |  |  | ARK Methotrexate Calibrators | 10709281 |  | Liquid |  | 2 | 6 |  | 62 | 15 |  | No |  | No |  | 10709285 | [10 empty] | 0.02 μmol/L | 0.00 μmol/L | 0.04 μmol/L |  |  |  | 0.07 |  | 6.5 | 8.8 | 0.42 |  | 1.7 | 3.5 | 0.81 |  | 2.6 | 3.8 | Roche/Hitachi 917 y = 0.95x + 0 μmol/L r = 0.97 |  | N/A | N/A | N/A | No Biotin in assay formulation | [5 empty] | 1600-0213-00 Rev. 07 | IFU Document
Atellica® IM | 10995541 |  | di04 | 630414600413 | Global | DiabetesDiabetes | No | Atellica IM C-peptide (CpS) | Atellica IM CpSCpS | CpS | 100 | 1 x 100 | Atellica IM CpS CAL SMN 10995542 | Atellica IM Multi-Diluent 10 SMN 10995640 Atellica IM CpS MCM SMN 10995543 | 14 | 28 | Monoclonal mouse anti-C-peptide antibody | ;#Serum;# | Serum : Up to 8 hrs at RT, Up to 24 hrs at 2-8°C, Frozen at <-20 °C | 50 μL | No |  | 2-site sandwich immunoassay using direct chemiluminescent technology | 0.05 - 30.00 ng/mL | 0.05 - 30.00 ng/mL | 1:2 0.05 - 60.00ng/mL 1:5 0.05 - 150 .0 ng/mL | 1:2 0.05 - 60.00ng/mL 1:5 0.05 - 150 .0 ng/mL | 0.81 - 3.85 ng/mL | 0.81 - 3.85 ng/mL | NA | NA | NA | Atellica IM CpS CAL | 10995542 | 4 hours at RT | Liquid | Let the calibrators stand for 5 minutes at RT, Gently mix | 3 | 2 | citric acid buffer with casein and preservatives | 50 | 14 |  |  |  | Yes | Multi-Diluent 10 | 10995640 |  | 1:2 0.05 - 60.00ng/mL 1:5 0.05 - 150 .0 ng/mL | 1:2 0.05 - 60.00ng/mL 1:5 0.05 - 150 .0 ng/mL | No | Use comercially available control | NA | 0.03 ng/mL | 0.03 ng/mL | NA | NA | 0.11 ng/mL | 0.09 ng/mL | NA | 0.09 ng/mL | 0.11 ng/mL | NA | 1.08 ng/mL | 1.08 ng/mL | 2 | 2.9 | 13.81 ng/mL | 13.81 ng/mL | 1.5 | 3.3 | 24.12 ng/mL | 24.12 ng/mL | 3.2 | 4.9 | ADVIA Centaur CpS; y = 0.98x - 0.03 ng/mL r = 1.00 | ADVIA Centaur CpS; y = 0.98x - 0.03 ng/mL r = 1.00 | ≤ 10% change up to 250 mg/dL of hemoglobin | ≤ 10% change up to 20 mg/dL of bilirubin | ≤ 10% change up to 1000 mg/dL of lipid | No Biotin in assay formulation | 125 ng/mL | 125 ng/mL | 125 ng/mL | 125 ng/mL | WHO International Standard 84/510 | 10997742_EN Rev. 01, 2017-10 | IFU Document | 0.227 | 0 | 0 | 0 | 0.103 | 17.7 | 3.7 | 5.3
Atellica® IM | 10995628 |  | di05 | 630414600444 | Global | DiabetesDiabetes | No | Atellica IM Insulin (IRI) | Atellica IM IRIIRI | IRI | 100 | 1 x 100 | Atellica IM IRI CAL SMN 10995629 | Atellica IM IRI MCM SMN 10995632 | 14 | 21 | mouse monoclonal anti-insulin antibodies | ;#Serum;# | Up to 8 hrs at RT, Up to 24 hrs at 2-8°C, Freeze at -20 °C after 24 hrs | 25 μL | No |  | 2-site sandwich chemiluminescent with acridinium-ester technology | 0.5 - 300.0 mU/L | 0.5 - 300.0 mU/L | 1:2 05 - 600 mU/L 1:5 0.5 - 1500 mU/L | 1:2; 05 - 600 mU/L; 1:5 0.5 - 1500 mU/L | 3.0 - 25 mU/L | 3.0 - 25 mU/L | NA | NA | NA | Atellica IM IRI CAL | 10995629 | 8 hrs at RT | Liquid | Let the calibrators stand for 5 minutes at RT, Gently mix | 3 | 2 | Buffered saline with casein, potassium thiocyanate, sodium azide, and preservatives | 91 | 14 |  |  |  | Yes | Atellica IM IRI DIL | 10995630 (2 x 10 mL) 10995631 (1 x 20 mL) |  | 1:2 0.5 - 600 mU/L 1:5 0.5 - 1500 mU/L | 1:2 0.5 - 600 mU/L 1:5 0.5 - 1500 mU/L | No | Use Commercially Available Controls | NA | 0.3 mU/L | 0.3 mU/L | NA | NA | 0.8 mU/L | 0.6 mU/L | NA | 0.6 mU/L | 0.8 mU/L | NA | 3.7 mU/L | 3.7 mU/L | 1.8 | 3.5 | 52.0 mU/L | 52.0 mU/L | 1.6 | 1.6 | 255.7 mU/L | 255.7 mU/L | 1.4 | 2.8 | ADVIA Centaur Insulin y = 0.99x + 0.06 mU/L r = 0.99 | ADVIA Centaur Insulin y = 0.99x + 0.06 mU/L r = 0.99 | ≤ 6% change up to 125 mg/dL of hemoglobin | ≤ 6% change up to 20 mg/dL of bilirubin | ≤ 6% change up to 1000 mg/dL of lipid | No Biotin in assay formulation | 3000 mU/L | 3000 mU/L | 3000 mU/L | 3000 mU/L | WHO 1st IRP 66/3024 | 10997752_EN Rev. 04, 2020-08 | IFU Document | 0.213 | 6.969 | 1.451 | 2.087 | 0.097 | 17.7 | 3.7 | 5.3
Atellica® CH | 10317369 |  | pl16 | 08058056680730 | Global | Specific Proteins | Yes | Ceruloplasmin (Sentinel Alliance Application) | CERULCERUL | 7054 | 230 | 2 x 115 | Atellica CH Empty Reagent Packs (SMN 11097534) | N/A | 9.7 | 35 | anti-ceruloplasmin polyclonal antiserum (goat) | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | 3 days at 2-8 °C | 4 µL | No |  | Turbidimetric | 5 - 90 mg/dL | 5 - 90 mg/dL | N/A | N/A | 20 - 60 mg/dL | 20 - 60 mg/dL | N/A | N/A | N/A | Plasma proteins Cal 3X | 10330537 | 30 days at 2-8 °C | Liquid | No Preparation required | 3 | 5 | Plasmatic plasma proteins in human based matrix | 35 | 35 |  | No |  | No | N/A | 11099300 |  | N/A | N/A | No | N/A | N/A | N/A | N/A | N/A | N/A | 2 mg/dL | N/A | N/A | N/A | 2 mg/dL | N/A | 8.28 mg/dL | 8.28 mg/dL | 1.5 | 3.1 | 33.80 mg/dL | 33.80 mg/dL | 1 | 1.8 | 73.01 mg/dL | 73.01 mg/dL | 0.5 | 1.7 | ADVIA Chemistry 2400 y = 0.960 x - 0.690 mg/dL r = 0.977 n = 143 | ADVIA Chemistry 2400 y = 0.960 x - 0.690 mg/dL r = 0.977 n = 143 | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 30 mg/dL of bilirubin | ≤ 10% change up to 1000 mg/dL of Trig. | N/A | no prozone effect was observed up to a concentration of 1500 mg/dL | no prozone effect was observed up to a concentration of 1500 mg/dL | no prozone effect was observed up to a concentration of 1500 mg/dL | no prozone effect was observed up to a concentration of 1500 mg/dL | CAP RPSP-3 (Reference Preparation for Serum Proteins Lot-3) / RPPHS (Reference Preparation for Protein in Human Serum) | IFU: 11065N-2.0/ 02 2015/09/18 and App. Sheet: F57 2019/12/17 | IFU Document | 0.463 | 4.37 | 2.42 | 4.69 | 0.21 | 11.1 | 6.15 | 11.9
Atellica® CH | 11097647 |  | pl18 | 00630414236919 | Global | Specific Proteins | No | Atellica CH Cystatin C_2 (CYSC_2) | CYSC_2CYSC_2 | 4051 | 800 | 4 x 200 | N/A | N/A | 10 | 40 | anti-cystatin C antibody (rabbit) | ;#Serum;#Lithium Heparin Plasma;#Potassium EDTA Plasma;# | 7 days at 2-8C; Frozen at -20C or below at least 3 months if they are frozen within 24 hours after collection and if repeated freeze-thaw cycles are avoided. | 5.2 μL | No |  | Latex-enhanced immunoturbidimetric | 0.25–8.93 mg/L | 0.25–8.93 mg/L | 0.25–26.79 mg/L | 0.25–26.79 mg/L | 0.64–1.23 mg/L for adults | 0.64–1.23 mg/L for adults | N/A | N/A | N/A | Atellica® CH Cystatin C_2 Calibrators (CYSC_2 CAL) | 11097648 | 60 | Liquid | N/A | 3 | 6 | Human serum; recombinant human cystatin C | 61 | 40 |  | No |  | Yes | Atellica CH Diluent | 11099300 |  | 0.25–26.79 mg/L | 0.25–26.79 mg/L | No | N/A | N/A | 0.12 mg/L | 0.12 mg/L | N/A | N/A | 0.12 mg/L | 0.08 mg/L | 0.25 mg/L | 0.08 mg/L | 0.12 mg/L | 0.25 mg/L | 0.55 mg/L | 0.55 mg/L | 3.6 | 3.7 | 0.86 mg/L | 0.86 mg/L | 2.1 | 2.2 | 3.86 mg/L | 3.86 mg/L | 0.6 | 0.8 | ADVIA Chemistry 1800 CYSC_2 y = 1.01x + 0.04 mg/L r = 0.999 n=100 | ADVIA Chemistry 1800 CYSC_2 y = 1.01x + 0.04 mg/L r = 0.999 n=100 | ≤ 10% change up to 1000 mg/dL of hemoglobin. | ≤ 10% change up to 60 mg/dL of conjugated bilirubin and 60 mg/dL of unconjugated bilirubin. | ≤ 10% change up to 1000 mg/dL of Intralipid. | No Biotin in assay formulation | 56.00 mg/L | 56.00 mg/L | 56.00 mg/L | 56.00 mg/L | IFCC Standard ERM–DA471 | 11111654_EN Rev. 04, 2020-08 | IFU Document | 0.679 | 5.625 | 3.5 | 2.75 | 0.308 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097645 |  | pl17 | 00630414220703 | OUS | Specific Proteins | No | Atellica CH Wide Range C-Reactive Protein (wrCRP) | wrCRPwrCRP | 4042 | 2000 | 4 x 500 | N/A | N/A | 8 | 51 | CRP antibody synthetic latex (rabbit) | ;#Serum;#Lithium Heparin Plasma;# | 13 days at 2-8C; Frozen at -20C up to 6 months | 10 μL | No |  | Latex-enhanced immunoturbidimetric assay | 0.050–15.600 mg/dL | 0.50–156.00 mg/L | 0.050– 62.400 mg/dL | 0.50–624.00 mg/L | 0.000–0.500 mg/dL for adults | 0.00–5.00 mg/L for adults | N/A | N/A | N/A | Atellica CH Wide Range C-Reactive Protein Calibrators (wrCRP CAL) | 11097646 | 28 | Liquid | Ready to Use | 3 | 6 | Recombinant human C-reactive protein (CRP) in a stabilized protein | 61 | 51 |  | No |  | Yes | NA | 11099300 |  | 0.050– 62.400 mg/dL | 0.50–624.00 mg/L | No | NA | NA | 0.030 mg/dL | 0.30 mg/L | N/A | N/A | 0.030 mg/dL | 0.000 mg/dL | 0.050 mg/dL | 0.00 mg/L | 0.30 mg/L | 0.50 mg/L | 0.562 mg/dL | 5.62 mg/L | 6 | 8 | 5.318 mg/dL | 53.18 mg/L | 4 | 6 | 12.762 mg/dL | 127.62 mg/L | 3 | 5 | BN ProSpec CardioPhase hsCRP y = 1.08x - 0.030 mg/dL r=0.997 n=127 | BN ProSpec CardioPhase hsCRP y = 1.08x - 0.30 mg/L r=0.997 n=127 | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 30 mg/dL of conjugated bilirubin and 30 mg/dL unconjugated bilirubin. | ≤ 10% change up to 500 mg/dL of Intralipid | No Biotin in assay formulation | 95.000 mg/dL | 950.00 mg/L | 95.000 mg/dL | 950.00 mg/L | IRMM reference material ERM-DA474/IFCC | 11111647_EN Rev. 02, 2019-10 | IFU Document | 0.974 | 5.625 | 3.5 | 2.75 | 0.442 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097644 |  | ge05 | 00630414220697 | Global | General Chemistry | No | Atellica CH Calcium (Arsenazo) Serum/plasma | CA_2CA_2 | 4041 | 8200 | 4 x 2050 | N/A | N/A | 3 | 63 | Sodium Acetate, Arsenazo III, Non-reactive stabilizers | ;#Serum;#Lithium Heparin Plasma;# | 8hr at RT; 2 days at 4C; frozen at -20C at least 6 months | 4 μL |  |  | Arsenazo - Colorimetric | 1.0–16.0 mg/dL | 0.25–4.00 mmol/L | 1.0–32.0 mg/dL | 0.25–8.00 mmol/L | 8.7–10.4 mg/dL | 2.18–2.60 mmol/L | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 48 hours | Lyophilized | Add 3.0 mL of deionized water into the vial | 3 | 2 | Bovine Serum | 63 | 63 |  | No |  | Yes | N/A | 11099300 |  | 1.0–32.0 mg/dL | 0.25–8.00 mmol/L | No | N/A | N/A | 0.1 mg/dL | 0.03 mmol/L | N/A | N/A | 0.1 mg/dL | 0.0 mg/dL | N/A | 0.00 mmol/L | 0.03 mmol/L | N/A | 6.5 mg/dL | 1.63 mmol/L | 1.1 | 1.5 | 10.2 mg/dL | 2.55 mmol/L | 0.7 | 0.7 | 14.1 mg/dL | 3.53 mmol/L | 0.6 | 0.7 | ADVIA Chemistry 1800 CA_2 y = 1.02x - 0.1 mg/dL r=1.000 n=100 | ADVIA Chemistry 1800 CA_2 y = 1.02x - 0.03 mmol/L r=1.000 n=100 | ≤ 10% change up to 1000 mg/dL of hemoglobin. | ≤ 10% change up to 50 mg/dL of conjugated bilirubin and 50 mg/dL of unconjugated bilirubin. | ≤ 10% change up to 1000 mg/dL of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | Inductively Coupled Plasma Atomic Emission/NIST | 11111651_EN Rev. 03, 2019-03 | IFU Document | 0.604 | 2.875 | 3.5 | 2.75 | 0.274 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097649 |  | ge19 | 00630414238357 | Global | General Chemistry | No | Atellica CH Amylase_2 | AMY_2AMY_2 | 4052 | 1050 | 3 x 350 | N/A | N/A | 9 | 31 | α-Glucosidase (microorganism); Ethylidene-4-NP-G7 | ;#Serum;#Lithium Heparin Plasma;#Urine;# | Serum and Plasma: 8 days at RT; 31 days at 2–8°C; at least 1 year at -20°C Urine: 10 days at RT; 31 days at 2–8°C | 16 µL | No |  | Ethylidene Blocked-pNPG7 | 20 – 1500 U/L | 20 – 1500 U/L | 20 – 4500 U/L | 20 – 4500 U/L | Serum and plasma: 30–118 U/L Urine: ≤ 650 U/L | Serum and plasma: 30–118 U/L Urine: ≤ 650 U/L | N/A | N/A | N/A | Atellica CH SPCL CHEM CAL | 11099438 | 7 | Lyophilized | 5.0 mL of distilled water; Stand for 30 minutes at room temperature; Invert to mix Prior to use mix the contents by gently inverting the vials | 3 | 2 | Human serum | 62 | 31 |  | No |  | Yes | N/A | 11099300 |  | 20 – 4500 U/L | 20 – 4500 U/L | No | N/A | N/A | 7 U/L | 7 U/L | N/A | N/A | 7 U/L | 1 U/L | 20 U/L | 1 U/L | 7 U/L | 20 U/L | 52 U/L | 52 U/L | 1.1 | 1.4 | 187 U/L | 187 U/L | 0.4 | 0.6 | 1128 U/L | 1128 U/L | 0.2 | 0.4 | Roche cobas AMYL2 (serum) y = 1.09x + 0 U/L r = 0.996 n= 118 Roche cobas AMYL2 (urine) y = 1.11x - 1 U/L r = 0.998 n= 114 | Roche cobas AMYL2 (serum) y = 1.09x + 0 U/L r = 0.996 n= 118 Roche cobas AMYL2 (urine) y = 1.11x - 1 U/L r = 0.998 n= 114 | < 10% change up to 500 mg/dL of hemoglobin. | < 10% change up to 30 mg/dL of bilirubin. | < 10% change up to 650 mg/dL of Intralipid® | No Biotin in assay formulation | N/A | N/A | N/A | N/A | Traceable to the IRMM/IFCC-456 reference material and commutable to the IFCC Alpha-Amylase Primary Reference Procedure | 11111707_EN Rev. 02, 2019-08 | IFU Document | 0.6 | 4.187 | 3.5 | 2.75 | 0.272 | 10.636 | 8.89 | 6.985
Atellica® CH | 11097579 |  | ge19 | 00630414238364 | OUS | General Chemistry | No | Atellica CH Pancreatic Amylase_2 (serum and plasma) | PAMY_2PAMY_2 | 4053 | 1720 | 4 x 430 | N/A | N/A | 9 | 31 | monoclonal mouse antibodies | ;#Serum;#Lithium Heparin Plasma;# | 8 days at RT; 31 days at 2-8C; at least 1 year frozen at -20C | 16 | No |  | Enzymatic reaction | 20–1500 U/L | 20–1500 U/L | 20–7500 U/L | 20–7500 U/L | Serum and Plasma: 13–53 U/L | Serum and Plasma: 13-53 U/L | N/A | N/A | N/A | Atellica CH SPCL CHEM CAL | 11097579 | 7 | Lyophilized | 5.0 mL of distilled water; Stand for 30 minutes at room temperature; Invert to mix Prior to use mix the contents by gently inverting the vials | 3 | 2 | Human serum | 62 | 31 |  | No |  | Yes | N/A | 11099300 |  | 20–7500 U/L | 20–7500 U/L | No | N/A | N/A | 9 U/L | 9 U/L | N/A | N/A | 9 U/L | 1 U/L | 20 U/L | 1 U/L | 9 U/L | 20 U/L | 54 U/L | 54 U/L | 0.9 | 1.2 | 186 U/L | 186 U/L | 0.4 | 0.5 | 1129 U/L | 1129 U/L | 0.2 | 0.3 | Roche cobas AMY-P (serum) y = 1.07x + 1.2 U/L r=1.000 n=106 | Roche cobas AMY-P (serum) y = 1.07x + 1.2 U/L r=1.000 n=106 | ≤ 10% change up to 500 mg/dL of hemoglobin. | 10% change up to 30 mg/dL of bilirubin. | ≤ 10% change up to 1000 mg/dL of Intralipid |  | N/A | N/A | N/A | N/A | IRMM/IFCC-456 Pancreatic Alpha- Amylase Reference Material and traceable to JCCLS-SOP through the CRM-001c reference material. | 11111708_EN Rev. 03, 2020-01 | IFU Document | 0.886 | 5.625 | 3.5 | 2.75 | 0.402 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097644 |  | ge05 | 00630414220697 | Global | General Chemistry | No | Atellica CH Calcium (Arsenazo) urine | CA_2CA_2 | 4041 | 8200 | 4 x 2050 | N/A | N/A | 3 | 63 | Sodium Acetate, Arsenazo III, Non-reactive stabilizers | ;#Urine;# | Urine specimens should be collected in a bottle containing 10 mL of 6 M HCl per 24–hour specimen to prevent calcium salt precipitation. | 4 μL | No |  | Arsenazo - Colorimetric | 1.0–32.0 mg/dL | 0.25–8.00 mmol/L | 1.0–160.0 mg/dL | 0.25–40.00 mmol/L | Urine: 100–300 mg/day | Urine: 2.50–7.50 mmol/day | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 48 hours | Lyophilized | Add 3.0 mL of deionized water into the vial | 3 | 2 | Bovine serum | 63 | 63 |  | No |  | Yes | N/A | 11099300 |  | 1.0–160.0 mg/dL | 0.25–40.00 mmol/L | No | N/A | N/A | 0.2 mg/dL | 0.50 mmol/L | N/A | N/A | 0.2 mg/dL | 0.1 mg/dL | N/A | 0.03 mmol/L | 0.50 mmol/L | N/A | 6.2 mg/dL | 1.55 mmol/L | 2 | 2.3 | 12.1 mg/dL | 3.03 mmol/L | 0.6 | 1 | 24.0 mg/dL | 6.00 mmol/L | 0.4 | 0.5 | ADVIA Chemistry 1800 CA_2 (urine) y = 1.03x - 0.2 mg/dL r=0.997 n=107 | ADVIA Chemistry 1800 CA_2 (urine) y = 1.03x - 0.05 mmol/L r=0.997 n=107 | N/A | N/A | N/A | No Biotin in assay formulation | N/A | N/A | N/A | N/A | Inductively Coupled Plasma Atomic Emission/NIST | 11111651_EN Rev. 03, 2019-03 | IFU Document | 0.604 | 2.875 | 3.5 | 2.75 | 0.274 | 7.302 | 8.89 | 6.985
Atellica® CH | 10482437 |  | pl28 | 00842768031410 | OUS | Specific Proteins | Yes | N LATEX FLC KAPPA (BN Alliance Application) | FLCKFLCK | 7040 | 72 | 3 x 1.7 mL | N FLC SUPPLEMENTARY REAGENT (10482439) | N/A | 9.26 | 14 | Suspension of polystyrene particles coated with monoclonal antibodies (mouse) to human FLC kappa | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | S,P: 5 days at 2-8C; frozen at or below -20C up to 6 months | 2 μL | No |  | Particle-enhanced turbidimetry | 3.91 to 60.00 mg/L | 3.91 to 60.00 mg/L | 1.17 to 600.00 mg/L | 1.17 to 600.00 mg/L | S,P: Kappa: 6.70–22.4 mg/L FLC κ/λ ratio: 0.86 (0.31 to 1.56) - median, (minimum – maximum) | S,P: Kappa: 6.70–22.4 mg/L FLC κ/λ ratio: 0.86 (0.31 to 1.56) - median, (minimum – maximum) | N/A | N/A | N/A | N FLC STANDARD SL | 10482440 | 42 | Liquid | N/A | 3 | 6 | Human free light chain proteins, human serum albumin and protease inhibitors | 42 | 14 |  | No |  | Yes | N/A | 11099300 |  | 1.17 to 600.00 mg/L | 1.17 to 600.00 mg/L | Yes | N FLC CONTROL SL1 and SL2 | 10482441 and 10482442 | N/A | N/A | N/A | N/A | N/A | N/A | 1.037 mg/L | N/A | N/A | 1.037 mg/L | 7.82 mg/L | 7.82 mg/L | 1.02 | 1.72 | 26.61 mg/L | 26.61 mg/L | 0.68 | 1.33 | 50.08 mg/L | 50.08 mg/L | 0.72 | 1.41 | BN ProSpec® System N Latex FLC kappa y=0.942 x –0.075 mg/L r=0.982 n=167 BN ProSpec® System FLC κ/λ ratio y=1.03 x –0.013 Ratio r=0.971 n=163 | BN ProSpec® System N Latex FLC kappa y=0.942 x –0.075 mg/L r=0.982 n=167 BN ProSpec® System FLC κ/λ ratio y=1.03 x –0.013 Ratio r=0.971 n=163 | ≤ 10% change up to 1100 mg/dL of hemoglobin. | ≤ 10% change up to 65 mg/dL of bilirubin. | ≤ 10% change up to 1115 mg/dL of Trig. | N/A | No high-dose hook effect up to an analyte concentration of 30 000 mg/L | No high-dose hook effect up to an analyte concentration of 30 000 mg/L | No high-dose hook effect up to an analyte concentration of 30 000 mg/L | No high-dose hook effect up to an analyte concentration of 30 000 mg/L | Internal Reference Plasma-Pool | OPJAG03C11 Rev. 08 – en | IFU Document | 0.22 | 3.15 | 1.38 | 2.36 | 0.01 | 8 | 3.5 | 6
Atellica® CH | 10482438 |  | pl29 | 00842768031427 | OUS | Specific Proteins | Yes | N LATEX FLC LAMBDA (BN Alliance Application) | FLCLFLCL | 7041 | 83 | 3 x 2.1 mL | N FLC SUPPLEMENTARY REAGENT (10482439) | N/A | 9.26 | 14 | Suspension of polystyrene particles coated with monoclonal antibodies (mouse) to human FLC lambda | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | S,P: 5 days at 2-8C; frozen at or below -20C up to 6 months | 5 μL | No |  | Particleenhanced turbidimetry | 5.47 to 70.00 mg/L | 5.47 to 70.00 mg/L | 1.64 to 700.00 mg/L | 1.64 to 700.00 mg/L | S,P: Lambda: 8.30–27.0 mg/L FLC κ/λ ratio: 0.86 (0.31 to 1.56) - median, (minimum – maximum) | S,P: Lambda: 8.30–27.0 mg/L FLC κ/λ ratio: 0.86 (0.31 to 1.56) - median, (minimum – maximum) | N/A | N/A | N/A | N FLC STANDARD SL | 10482440 | 42 | Liquid | N/A | 3 | 6 | Human free light chain proteins, human serum albumin and protease inhibitors | 42 | 14 |  | No |  | Yes | N/A | 11099300 |  | 1.64 to 700.00 mg/L | 1.64 to 700.00 mg/L | Yes | N FLC CONTROL SL1 and SL2 | 10482441 and 10482442 | N/A | N/A | N/A | N/A | N/A | N/A | 1.661 mg/L | N/A | N/A | 1.661 mg/L | 8.04 mg/L | 8.04 mg/L | 2.27 | 3.63 | 32.89 mg/L | 32.89 mg/L | 1.12 | 3.22 | 60.64 mg/L | 60.64 mg/L | 1.72 | 4.1 | BN ProSpec® System N Latex FLC lambda 0.927 x –0.067 mg/L r=0.977 n=170 BN ProSpec® System FLC κ/λ ratio y=1.03 x –0.013 Ratio r=0.971 n=163 | BN ProSpec® System N Latex FLC lambda 0.927 x –0.067 mg/L r=0.977 n=170 BN ProSpec® System FLC κ/λ ratio y=1.03 x –0.013 Ratio r=0.971 n=163 | ≤ 10% change up to 820 mg/dL of hemoglobin. | ≤ 10% change up to 65 mg/dL of bilirubin. | ≤ 10% change up to 734 mg/dL of Trig. | N/A | No high-dose hook effect up to an analyte concentration of 30 000 mg/L | No high-dose hook effect up to an analyte concentration of 30 000 mg/L | No high-dose hook effect up to an analyte concentration of 30 000 mg/L | No high-dose hook effect up to an analyte concentration of 30 000 mg/L | Internal Reference Plasma-Pool | OPJAG03C11 Rev. 08 – en | IFU Document | 0.37 | 3.15 | 1.38 | 2.36 | 0.168 | 8 | 3.5 | 6
Atellica® CH | 11097542 |  | is04 | 00630414238340 | Global | Therapeutic Drug Monitoring (TDM) | No | Atellica CH Tacrolimus (TACR) | TACR | 4048 | 400 | 2 x 200 |  |  | 6 | 31 | Mouse monoclonal antibody reactive to tacrolimus | ;#Whole blood;# | 8 hours at 18–25°C or for at least 1 week at 2–8°C or stored frozen for at least 6 months at -20°C | 25 μL |  |  | EMIT | 2.0 - 30.0 ng/mL | n/a | 2.0 - 80.0 ng/mL | n/a | No firm therapeutic range exists for tacrolimus in whole blood. | n/a |  |  |  | Emit 2000 Tacrolimus Calibrators (0, 2.5, 5, 10, 20, 30 ng/mL) | 10445399 | 56 | Liquid | Thaw at a room temperature of 18–25°C for approximately 1 hour or at a refrigerated temperature of 2–8°C overnight. Do not vortex | 3 | 6 | human whole blood | 64 | 10 |  | No |  | No |  | 10445399 | [10 empty] | 1.1 ng/mL | 0.6 ng/mL | 2.8 ng/mL | n/a | n/a | n/a | 4.7 ng/mL | n/a | 8.4 | 13.7 | 10.5 ng/mL | n/a | 3.9 | 6.7 | 15.1 ng/mL | n/a | 3 | 6.2 | Emit 2000 Tacrolimus on the Viva-ProE System (Liver and Kidney) y = 1.00x + 0.1 ng/mL r = 0.989 n = 125 |  | N/A | ≤ 10% interference up to 45 mg/dL bilirubin, conjugated (Tacrolimus concentration 9.1 and 15.2 ng/mL); ≤ 10% interference up to 60 mg/dL bilirubin, uncconjugated (Tacrolimus concentration 8.7 and 13.6 ng/mL); | ≤ 10% interference up to 1125 mg/dL Lipemia (Intralipid®) (Tacrolimus concentration 8.9 and 16.6 ng/mL) |  |  |  | N/A |  | Purified tacrolimus | 11111631_EN Rev. 01, 2018-09 | IFU Document |  |  |  |  | 0.135 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097540 |  | is02 | 00630414238081 | OUS | Therapeutic Drug Monitoring (TDM) | No | Atellica CH Mycophenolic Acid (MPA) | MPA | 4043 | 600 | 2 x 300 |  |  | 9 | 33 | Mouse monoclonal antibody to mycophenolic acid | ;#Lithium Heparin Plasma;#Potassium EDTA Plasma;# | 7 days at 2–8°C or stored frozen for at least 11 months at -20°C or colder. Samples should not be exposed to more than 3 freeze-thaw cycles | 5.7 μL |  |  | EMIT | 0.10 - 15.00 μg/mL | n/a | 0.10 - 30.00 μg/mL | n/a | No established therapeutic range exists for mycophenolic acid in plasma. | n/a |  |  |  | Emit 2000 Mycophenolic Acid Calibrators (0, 0.5, 2, 5, 10, 15 μg/mL) | 10445380 | expiration on bottle | Liquid | n/a | 3 | 6 | Bovine Serum Albumin | 63 | 6 |  | No |  | No |  | 10445380 | [10 empty] | 0.06 μg/mL | 0.03 μg/mL | n/a | n/a | n/a | n/a | 1.09 μg/mL | n/a | 2 | 5.2 | 7.33 μg/mL | n/a | 2.3 | 5.7 | 11.44 μg/mL | n/a | 3.3 | 7.2 | Syva Emit 2000 Mycophenolic Acid on Viva-ProE y = 0.95x - 0.04 μg/mL r = 0.991 n = 102 |  | ≤ 10% change up to 500 mg/dL of hemoglobin at analyte concentration of 1.11 and 4.20 μg/mL | ≤ 10% change up to 20 mg/dL of Bilirubin, conjugated at analyte concentration of 1.00 and 3.89 μg/mL; ≤ 10% change up to 20 mg/dL of Bilirubin, unconjugated at analyte concentration of 0.97 and 3.84 μg/mL | ≤ 10% change up to 937.5 mg/dL of Lipemia (Intralipid®) at analyte concentration of 1.07 and 3.85 μg/mL; ≤ 10% change up to 1250 mg/dL of Lipemia (Triglycerides concentrate) at analyte concentration of 0.99 and 3.82 μg/mL |  |  |  | N/A |  | Purified mycophenolic acid powder | 11111627_EN Rev. 01, 2018-09 | IFU Document |  |  |  |  | 0.172 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097541 |  | dr19 | 00630414238098 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Oxycodone (OXY) | OXY100 | 7046 | 1520 | 4 x 380 |  |  | 7 | 40 | Mouse monoclonal antibody to oxycodone | ;#Urine;# | Refrigerated or unrefrigerated for up to 7 days or stored frozen after 7 days at -20°C for up to 3 years. | 12 μL |  |  | EMIT | 100 ng/mL, Semi-quantitative 50 - 400 ng/mL | n/a | n/a | n/a | n/a | n/a |  |  |  | Emit® Calibrator/Control Level 0 Emit® II Plus Specialty Multi Drug Calibrators/Controls Level 1, 2, 3 | 10445406 10720807 10720813 10720814 | Once opened, the Emit® Calibrator/Control level 0 is stable for 5 weeks when recapped and stored at 2–8°C when not in use. For Emit® II Plus Specialty Multi Drug Level 1, 2 and 3 Expiration date on bottle | Liquid | n/a | 3 | 4 | human urine | 62 | 28 |  | No |  | No | [12 empty] | n/a | n/a | n/a |  |  |  | n/a | n/a |  |  | n/a | n/a |  |  | n/a | n/a |  |  | Atellica CH compared to LC-MS/MS n=100 100% agreement to positives (51/51) - Semiquantitative and Qualitative 84% agreement to negatives (41/49) - Semiquantitative and Qualitative |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  |  |  | >95% pure oxycodone solution confirmed by LC/MS/MS or GC/MS. | 11111624_EN Rev. 01, 2018-09 | IFU Document |  |  |  |  | 0.374 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097541 |  | dr19 | 00630414238098 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Oxycodone (OXY) | OXY300 | 7055 | 1520 | 4 x 380 |  |  | 7 | 40 | Mouse monoclonal antibody to oxycodone | ;#Urine;# | Refrigerated or unrefrigerated for up to 7 days or stored frozen after 7 days at -20°C for up to 3 years. | 4 μL |  |  | EMIT | 300 ng/mL, Semi-quantitative 75 - 1000 ng/mL | n/a | n/a | n/a | n//a | n/a |  |  |  | Emit® Calibrator/Control Level 0 Emit® II Plus Specialty Multi Drug Calibrators/Controls Level 1, 2, 3, 4 | 10445406 10720807 10720813 10720814 10720815 | Once opened, the Emit® Calibrator/Control level 0 is stable for 5 weeks when recapped and stored at 2–8°C when not in use. For Emit® II Plus Specialty Multi Drug Level 1, 2, 3 and 4 Expiration date on the bottle | Liquid | N/A | 3 | 5 | human urine | 62 | 28 |  | No |  | No | [12 empty] | n/a | n/a | n/a |  |  |  | n/a | n/a |  |  | n/a | n/a |  |  | n/a | n/a |  |  | Atellica CH compared to LC-MS/MS n=100 100% agreement to positives (37/37) - Semiqualitative and Qualitative 89% agreement to negatives (56/63) - Semiqualitative and Qualitative |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  |  |  | >95% pure oxycodone solution confirmed by LC/MS/MS or GC/MS. | 11111624_EN Rev. 01, 2018-09 | IFU Document |  |  |  |  | 0.374 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097539 |  | th07 | 00630414220680 | Global | Therapeutic Drug Monitoring (TDM) | No | Atellica CH Lidocaine (LIDO) | LIDO | 4044 | 200 | 2 x 100 |  |  | 8 | 31 | Mouse monoclonal antibody | ;#Serum;#Sodium Heparin Plasma;#Potassium EDTA Plasma;# | Serum specimens may be stable for at least 6 hours at room temperature.6 Serum and plasma specimens may be stable for at least 7 days at 2–8°C.8 pH neutral samples are stable for at least 1–2 years if stored at -20°C. | 7.5 μL |  |  | PETINIA | 0.5 - 12.0 μg/mL | 2.1–51.2 μmol/L | 0.5 - 24.0 μg/mL | 2.1–102.5 μmol/L | 1.5–5.0 μg/mL | 6.4–21.3 μmol/L |  |  |  | Atellica CH DRUG CAL II | 11099405 | 30 | Liquid | Allow to equilibrate to room temperature and mix thoroughly before use. | 3 | 5 | Bovine serum | 61 | 18 |  | No |  | Yes | [12 empty] | 0.3 μg/mL | 0.2 μg/mL | n/a | 0.9 μmol/L | 1.3 μmol/L | n/a | 1.5 μg/mL | 6.4 μmol/L | 2.5 | 3.1 | 7.5 μg/mL | 32.0 μmol/L | 1.2 | 1.6 | 10.4 μg/mL | 44.4 μmol/L | 1 | 1.3 | LIDO on Dimension EXL with LOCI Module y = 0.98x + 0.1 μg/mL (y = 0.98x + 0.4 μmol/L) n =110 r = 0.998 |  | ≤ 10% interference up to 375 mg/dL of hemoglobin at analyte concentration 5.1 and 8.3 μg/mL | ≤ 10% interference up to 60 mg/dL of Bilirubin, conjugated at analyte concentration 5.0 and 8.0 μg/mL; ≤ 10% interference up to 60 mg/dL of Bilirubin, unconjugated at analyte concentration 5.0 and 8.1 μg/mL; | ≤ 10% interference up to 1000 mg/dL of Lipemia (Intralipid®) at analyte concentration 5.0 and 8.0 μg/mL |  |  |  | N/A |  | USP Lidocaine standard | 11111635_EN Rev. 01, 2018-09 | IFU Document |  |  |  |  | 0.114 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097537 |  | th08 | 00630414220659 | Global | Therapeutic Drug Monitoring (TDM) | No | Atellica CH N-Acetylprocainamide (NAPA) | NAPA | 4045 | 200 | 2 x 100 |  |  | 7 | 31 | Mouse monoclonal antibody | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#Potassium EDTA Plasma;# | Separated serum specimens may be stored for at least 8 hours at room temperature. Serum and plasma specimens may be stored for at least 2 weeks at 2–8°C8 or stored frozen for at least 6 months at -20°C | 2.5 μL |  |  | PETINIA | 0.5–30.0 μg/mL | 1.8–108.3 μmol/L | 0.5–60.0 μg/mL | 1.8–216.6 μmol/L | The combined procainamide and N‑acetylprocainamide therapeutic range is 10–30 μg/mL. |  |  |  |  | Atellica CH DRUG CAL II | 11099405 | 30 | Liquid | Allow to equilibrate to room temperature and mix thoroughly before use. | 3 | 5 | Bovine Serum | 61 | 12 |  | No |  | Yes | [12 empty] | 0.3 μg/mL | 0.2 μg/mL | n/a | 0.7 μmol/L | 1.1 μmol/L | n/a | 2.2 μg/mL | 7.9 μmol/L | 2.8 | 4.5 | 13.4 μg/mL | 48.4 μmol/L | 0.7 | 1.4 | 23.6 μg/mL | 85.2 μmol/L | 0.5 | 1.1 | NAPA on Dimension EXL with LOCI Module y = 1.04x - 0.1 μg/mL (y = 1.04x - 0.4 μmol/L) n = 109 r = 0.997 |  | ≤ 10% interference from up to 500 mg/dL hemoglobin at analyte concentration of 9.8 and 14.4 μg/mL | ≤ 10% interference from up to 60 mg/dL Bilirubin, conjugated at analyte concentration of 9.4 and 14.3 μg/mL; ≤ 10% interference from up to 60 mg/dL Bilirubin, unconjugated at analyte concentration of 9.8 and 14.5 μg/mL; | ≤ 10% interference from up to 1000 mg/dL Lipemia (Intralipid®) at analyte concentration of 9.7 and 14.4 μg/mL |  |  |  | N/A |  | Traceable to N-acetylprocainamide of high purity. | 11111639_EN Rev. 01, 2018-09 | IFU Document |  |  |  |  | 0.114 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097538 |  | th10 | 00630414220673 | Global | Therapeutic Drug Monitoring (TDM) | No | Atellica CH Procainamide (PROC) | PROC | 4046 | 200 | 2 x 100 |  |  | 7 | 31 | Mouse monoclonal antibody | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#Potassium EDTA Plasma;# | Separated serum specimens may be stored for at least 8 hours at room temperature. Serum and plasma specimens may be stored for at least 2 weeks at 2–8°C or stored frozen for at least 6 months at -20°C. | 2.5 μL |  |  | PETINIA | 0.5–20.0 μg/mL | 2.1–85.0 μmol/L | 0.5–40.0 μg/mL | 2.1–170.0 μmol/L | 4.0–12.0 μg/mL | 17.0–51.0 μmol/L |  |  |  | Atellica CH DRUG CAL II | 11099405 | 30 | Liquid | Allow to equilibrate to room temperature and mix thoroughly before use. | 3 | 5 | Bovine serum | 61 | 18 |  | No |  | Yes | [12 empty] | 0.4 μg/mL | 0.2 μg/mL | n/a | 0.9 μmol/L | 1.7 μmol/L | n/a | 3.5 μg/mL | 14.9 μmol/L | 1.6 | 2.3 | 7.2 μg/mL | 30.6 μmol/L | 1.3 | 1.6 | 17.7 μg/mL | 75.2 μmol/L | 0.7 | 1 | PROC on Dimension EXL with LOCI Module y = 1.01x + 0.1 μg/mL (y = 1.01x + 0.4 μmol/L) n = 101 r = 0.998 |  | ≤ 10% interference from up to 500 mg/dL hemoglobin at analyte concentration 5.8 and 9.7 μg/mL | ≤ 10% interference from up to 60 mg/dL Bilirubin, conjugated at analyte concentration 5.6 and 9.5 μg/mL; ≤ 10% interference from up to 60 mg/dL Bilirubin, unconjugated at analyte concentration 5.8 and 9.5 μg/mL | ≤ 10% interference from up to 1000 mg/dL Lipemia (Intralipid®) at analyte concentration 5.7 and 9.5 μg/mL |  |  |  | N/A |  | Traceable to USP Procainamide standard | 11111643_EN Rev. 01, 2018-09 | IFU Document |  |  |  |  | 0.114 | 7.302 | 8.89 | 6.985
Atellica® CH | 11354454 |  | dr19 | 00858724003469 | OUS | Drugs of Abuse/Toxicology | Yes | ARK Ethyl Glucuronide (EtG) Assay (Alliance Application) | ETG500 | 7044 | 154 | 1 x 154 |  |  | 7.1 | 60 | Sheep monoclonal antibody to ethyl glucuronide | ;#Urine;# | Store at 2-8°C (36–46°F) and assay within 7 days after collection. After 7 days store the urine sample frozen. | 6 μL |  |  | Homogeneous enzyme immunoassay technique | 100 - 2000 ng/mL | n/a | n/a | n/a | n/a | n/a |  |  |  | ARK Ethyl Glucuronide Calibrators | 11354457 | Use prior to the expiration date. | Liquid | Calibrators are ready to use. Mix each level by gentle inversion before dispensing | 3 | 5 | Human Urine | 60 | 8 |  | No |  | No |  | 11354458 | [10 empty] | n/a | n/a | 50.0 ng/mL |  |  |  | 384.1 ng/mL | n/a | 2 | 2.7 | 505.6 ng/mL | n/a | 2.3 | 3 | 630.5 ng/mL | n/a | 1.4 | 2.3 | Atellica CH compared to Beckman AU680 100% agreement to positives (60/60) 100% agreement to negatives (40/40) n = 100 |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | There is no internationally recognized standard for ethyl glucuronide. A certified solution of ethyl glucuronide is traceable to LC/MS. | 1600-0609-00 Rev. 02 | IFU Document |  |  |  |  | 0.109 | 12.7 | 7.9 | 7.6
Atellica® CH | 11354454 |  | dr19 | 00858724003469 | OUS | Drugs of Abuse/Toxicology | Yes | ARK Ethyl Glucuronide (EtG) Assay (Alliance Application) | ETG1K | 7056 | 154 | 1 x 154 |  |  | 7.1 | 60 | Sheep monoclonal antibody to ethyl glucuronide | ;#Urine;# | Store at 2-8°C (36–46°F) and assay within 7 days after collection. After 7 days store the urine sample frozen. | 6 μL |  |  | Homogeneous enzyme immunoassay technique | 100 - 2000 ng/mL | n/a | n/a | n/a | n/a | n/a |  |  |  | ARK Ethyl Glucuronide Calibrators | 11354457 | Use prior to the expiration date. | Liquid | Calibrators are ready to use. Mix each level by gentle inversion before dispensing | 3 | 5 | Human Urine | 60 | 8 |  | No |  | No |  | 11354458 | [10 empty] | n/a | n/a | 50.0 ng/mL | n/a | n/a | n/a | 751.2 ng/mL | n/a | 1.8 | 2.3 | 1000.6 ng/mL | n/a | 1.3 | 2.4 | 1257.4 ng/mL | n/a | 1.8 | 2.5 | Atellica CH compared to Beckman AU680 98% agreement to positives (41/42) 100% agreement to negatives (58/58) n = 100 |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | There is no internationally recognized standard for ethyl glucuronide. A certified solution of ethyl glucuronide is traceable to LC/MS. | 1600-0609-00 Rev. 02 | IFU Document |  |  |  |  | 0.109 | 12.7 | 7.9 | 7.6
Atellica® CH | 10445383 |  | th14 | 00842768001543 | Global | Therapeutic Drug Monitoring (TDM) | No | Emit® Amikacin Assay (Alliance Application) | Amik | 7052 | 450 | 6 x 75 |  |  | 8.5 | 30 | Sheep antibodies reactive to amikacin | ;#Serum;#EDTA Plasma;#Citrate Plasma;# | Store the serum or plasma refrigerated at 2–8°C. | 6 µL |  |  | EMIT | 2.5-50 μg/mL | 4.3–86 μmol/L | 2.5 - 150 μg/mL | 4.3–258 μmol/L | 20–25 μg/mL for moderately severe infections and 25–30 μg/mL for life threatening infections or less sensitive pathogens. Trough amikacin oncentrations of 1–4 μg/mL for less severe infections and 5–8 μg/mL for life-threatening infections usually ensure that the concentration is above the minimum inhibitory concentration of the pathogen and that drug elimination is adequate. | 34–43 μmol/L for moderately severe infections and 43–51 μmol/L for life threatening infections or less sensitive pathogens. Trough amikacin concentrations of 1.7–6.8 μmol/L for less severe infections and 8.6–14 μmol/L for life-threatening infections usually ensure that the concentration is above the minimum inhibitory concentration of the pathogen and that drug elimination is adequate. |  |  |  | Emit Amikacin Calibrators (included in reagent kit) (0, 2.5, 5, 10, 20, 50) | 10445383 | 84 days | Lyophilized | Reconstitute with 1 mL of distilled or deionized water. After reconstitution, allow the calibrators to equilibrate at a room temperature for 2 hours | 3 | 6 | Human Serum | 19 | 19 |  | No |  | No |  | 10445383 | [10 empty] | n/a | n/a | n/a | n/a | n/a | n/a | 5.5 μg/mL | 9.4 μmol/L | 2.2 | 6.7 | 15.2 μg/mL | 26.0 μmol/L | 1.2 | 5 | 21.3 μg/mL | 36.4 μmol/L | 1.3 | 6.5 | Syva Amikacin assay on the Viva-ProE y=0.98x + 0.2 μg/mL n = 50 |  | No clinically significant interference has been found in samples with up to 800 mg/dL hemoglobin | No clinically significant interference has been found in samples with up to 30 mg/dL bilirubin | No clinically significant interference has been found in samples with up to 1000 mg/dL triglycerides |  |  |  | n/a | n/a | n/a | 6X314UL.12DS_US_E | IFU Document |  |  |  |  | 0.23
Atellica® CH | 10445455 |  | dr20 | 00842768020315 | Global | Drugs of Abuse/Toxicology | No | Emit II Lysergic Acid Diethylamide (LSD) Assay (Alliance Application) | LSD | 7045 | 240 | 1 x 240 |  |  | 9.7 | 30 | Mouse monoclonal antibodies reactive to lysergic acid diethylamide | ;#Urine;# | Unrefrigerated for up to 7 days following collection. After 7 days, specimens should be stored frozen. LSD degrades with prolonged exposure to light. Store specimens in opaque containers or wrap them in opaque paper or aluminum foil and store frozen. | 22.8 µL |  |  | EMIT | 0.5 ng/mL, Semi-quantitive 0–2.5 ng/mL | n/a | n/a | n/a | n/a | n/a |  |  |  | Emit® LSD 0 ng/mL Calibrator Emit® LSD 0.5 ng/mL Calibrator Emit® LSD 1.5 ng/mL Calibrator Emit® LSD 2.5 ng/mL Calibrator | 10445454 10445456 10445457 10445458 | Expiration date on label. | Liquid | n/a | 3 | 4 | Nonsterile human urine | 63 | 7 |  | No |  | No |  | n/a | [10 empty] | n/a | n/a | n/a | n/a | n/a | n/a | 119.22 mAU/min |  | 0.3 | 0.7 | 127.02 mAU/min |  | 0.3 | 0.6 | 138.76 mAU/min |  | 0.2 | 0.8 | Atellica CH compared to Viva-ProE System 100% agreement to positives (54/54) 100% agreement to negatives (52/52) n = 106 |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | n/a | 9L004UL.10DS_F | IFU Document |  |  |  |  | 0.2 | 14 | 9 | 8.9
Atellica® CH | 10445474 |  | dr21 | 00842768014437 | Global | Drugs of Abuse/Toxicology | No | Emit® II Plus Methaqualone (Alliance Application) | Mthaq | 7053 | 212 | 1 x 212 |  |  | 6.7 | 30 | Sheep polyclonal antibodies to methaqualone | ;#Urine;# | Stored unrefrigerated for up to 7 days, stored refrigerated for 30 days before analysis. After 7 days unrefrigerated or 30 days refrigerated, samples should be stored frozen. | 14 µL |  |  | EMIT | 300 ng/mL, Semi-quantitative 150–900 ng/mL | n/a | n/a | n/a | n/a | n/a |  |  |  | Emit® Calibrator/Control Level 0, 2, 3, 4, 5 | 10445406 10445408 10445409 10445410 10445411 | Once opened, the Emit® Calibrators/Controls are stable for 5 weeks when recapped and stored at 2–8°C when not in use | Liquid | n/a | 3 | 5 | Human Urine | 62 | 30 |  | No |  | No |  | n/a | [10 empty] | n/a | n/a | n/a | n/a | n/a | n/a | 147.94 mAU/min |  | 0.5 | 0.7 | 204.10 mAU/min |  | 0.3 | 0.5 | 176.81 mAU/min |  | 0.4 | 0.6 | Atellica CH compared to Viva-ProE System 100% agreement to positives (50/50) 100% agreement to negatives (50/50) n = 100 |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | n/a | 9Q022UL.7DS_E | IFU Document |  |  |  |  | 0.165 | 14 | 8.9 | 8.9
Atellica® CH | 10445287 |  | dr23 | 00842768017247 | Global | Drugs of Abuse/Toxicology | Yes | Syva® Creatinine Validity Test (Alliance Application) | SVTCR | 7047 | 462 | 3 x 154 |  |  | 6.3 | 7 | Reagent 1: NaOH with stabilizer; Reagent 2: 25 mM picric acid | ;#Urine;# | Stable for 1–2 days when stored at 2–10°C. | 8 µL |  |  | Colormetric | 2-20 mg/dL | n/a | n/a | n/a | SAMHSA classification of the specimen is based on creatinine and specific gravity results. See IFU Table 1. For countries outside the U.S., refer to country-specific guidelines. | n/a |  |  |  | Syva® Creatinine Validity Calibrator 2 Syva® Creatinine Validity Calibrator 20 | 10445273 10445269 | 21 days | Liquid | n/a | 3 | 2 | aqueous solution of creatinine | 80 | 2 |  | No |  | No |  | n/a | [10 empty] | n/a | n/a | n/a | n/a | n/a | n/a | 1.92 mg/dL |  | 1.3 | 5.8 | 19.31 mg/dL |  | 0.3 | 4.5 | 96.78 mg/dL |  | 0.3 | 4.5 | Syva® Creatinine Validity Test on the Viva-ProE analyzer y = 1.0x + 1.0 mg/dL n = 100 |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | n/a | 3T252UL.7DS_C | IFU Document |  |  |  |  | 0.265 | 14 | 9 | 11.5
Atellica® CH | 10445288 |  | vt1 | 00842768017254 | Global | Drugs of Abuse/Toxicology | Yes | Syva® Nitrite Validity Test (Alliance Application) | SVTNT | 7048 | 308 | 2 x 154 |  |  | 7.3 | 30 | Sulfanilamide in 0.3 M H3PO4 with surfactant | ;#Urine;# | Room temperature for 7 days, after 7 days store sample frozen. | 6.8 µL |  |  | Colormetric | 5-3500 μg/mL | n/a | n/a | n/a | A urine specimen is reported adulterated when the nitrite concentration is ≥500 μg/mL. A urine specimen is reported as invalid when the nitrite concentration is ≥200 μg/mL. | n/a |  |  |  | Syva® Validity Negative Calibrator/ Control Syva® Nitrite Validity Calibrator 500 | 10445268 10445275 | 21 days | Liquid | n/a | 3 | 2 | aqueous solution of NaNO2, NaCl, and preservative | 62 | 30 |  | No |  | No |  | n/a | [10 empty] | n/a | n/a | n/a | n/a | n/a | n/a | 222 µg/mL |  | 0.4 | 1.9 | 580 µg/mL |  | 0.3 | 1 | 493 µg/mL |  | 0.3 | 1 | Syva® Nitrite Validity Test on the Viva-ProE 100% Agreement to Positives 50/50 100% Agreement to Negatives 50/50 n = 100 |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | n/a | 3T162UL.6DS_US_C | IFU Document |  |  |  |  | 0.22 | 15.4 | 10.5 | 10.3
Atellica® CH | 10445300 |  | vt1 | 00842768017353 | Global | Drugs of Abuse/Toxicology | Yes | Syva® Specific Gravity Validity Test (Alliance Application) | SVTSG | 7051 | 648 | 4 x 162 |  |  | 5.1 | 30 | Poly (methyl vinyl ether-alt-maleic acid), isopropyl alcohol, Brij-35, sodium dodecyl sulfate | ;#Urine;# | Store at room temperature for 7 days. After 7 days, samples should be stored frozen. | 2 µL |  |  | Colormetric | 1.0010–1.0250 | n/a | n/a | n/a | Any specimen with a specific gravity result that is less than 1.0030 or greater than 1.0200 or has a creatinine value < 20 mg/dL should be tested on a refractometer in order to determine suitability for drug screening | n/a |  |  |  | Syva® Specific Gravity Validity Calibrator 1.0030 Syva® Specific Gravity Validity Calibrator 1.0200 | 10445290 10445291 | 21 | Liquid | n/a | 3 | 2 | Synthetic urine matrix | 52 | 2 |  | No |  | No |  | n/a | [10 empty] | n/a | n/a | n/a | n/a | n/a | n/a | 1.0027 |  | 0.03 | 0.1 | 1.0030 |  | 0.02 | 0.1 | 1.0200 |  | 0.02 | 0.1 | Atellica CH compared to Viva-ProE System (1.0030) 97% agreement to positives (76/80) 100% agreement to negatives (50/50) n = 130 Atellica CH compared to Viva-ProE System (1.0200) 100% agreement to positives (50/50) 100% agreement to negatives (80/80) n = 130 |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | n/a | 3T062UL.3DS_C | IFU Document |  |  |  |  | 0.226 | 14 | 9 | 11.7
Atellica® CH | 10445299 |  | vt1 | 00842768017346 | Global | Drugs of Abuse/Toxicology | Yes | Syva® pH Validity Test (Alliance Application) | SVTPH | 7050 | 810 | 5 x 162 |  |  | 4.5 | 30 | Aqueous solution of thymol blue | ;#Urine;# | Frozen samples must be thawed and mixed thoroughly before analysis. | 6.4 µL |  |  | Colormetric | 2.0 - 12.0 | n/a | n/a | n/a | The specimen will be characterized as adulterated if the pH is <3 or ≥11. The specimen will be characterized as invalid if the pH is ≥3.0 and <4.5 or ≥9.0 and <11.0. | n/a |  |  |  | Syva® pH Validity Calibrator 2.0, 3.0, 4.5, 9.0, 11.0, 12.0 | 10445279 10445280 10445286 10445285 10445283 10445284 | 90 | Liquid | n/a | 3 | 6 | Synthetic urine | 50 | 30 |  | No |  | No |  | n/a | [10 empty] | n/a | n/a | n/a | n/a | n/a | n/a | 3.20 |  | 0.5 | 4.9 | 4.46 |  | 0.5 | 2.1 | 9.00 |  | 1.8 | 2.5 | Atellica CH compared to Viva-ProE System (3.0 cutoff) 100% agreement to positives (52/52) 100% agreement to negatives (50/50) n=102 Atellica CH compared to Viva-ProE System (4.5 cutoff) 100% agreement to positives (52/52) 100% agreement to negatives (50/50) n=102 Atellica CH compared to Viva-ProE System (9.0 cutoff) 100% agreement to positives (52/52) 100% agreement to negatives (50/50) n=102 Atellica CH compared to Viva-ProE System (11.0 cutoff) 100% agreement to positives (50/50) 100% agreement to negatives (52/52) n=102 |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | n/a | 3T072UL.5DS_US_B | IFU Document |  |  |  |  | 0.218 | 14 | 11.5 | 9
Atellica® CH | 10445289 |  | vt1 | 00842768017261 | Global | Drugs of Abuse/Toxicology | Yes | Syva® Oxidant Validity Test (Alliance Application) | SVTOX | 7049 | 800 | 5 x 160 |  |  | 4.5 | 30 | 0.03% 2,2’-azino-bis(3-ethylbenzthiazoline-6-sulfonic acid) diammonium salt | ;#Urine;# | Store at room temperature for 7 days. After 7 days, samples should be stored frozen. | 6 µL |  |  | Colormetric | 1.0-50 μg/mL | n/a | n/a | n/a | An oxidant concentration assayed at >50 μg/mL is not normally observed in human urine; such samples should be suspected of adulteration. | n/a |  |  |  | Syva® Validity Negative Calibrator/Control Syva® Chromium (VI) Validity Calibrator 50 | 10445268 10445277 | 28 days | Liquid | n/a | 3 | 2 | Aqueous solution of 50 μg/mL Chromium | 62 | 21 |  | No |  | No |  | n/a | [10 empty] | n/a | n/a | n/a | n/a | n/a | n/a | 39.8 µg/mL |  | 0.3 | 1.4 | 72.1 µg/mL |  | 0.6 | 1.2 | 50.9 µg/mL |  | 0.3 | 1.5 | Atellica CH compared to Viva-ProE System 100% agreement to positives (50/50) 100% agreement to negatives (50/50) n = 100 |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | n/a | 3T674UL.5DS_US_C | IFU Document |  |  |  |  | 0.207 | 14.1 | 9 | 11.5
Atellica® IM | 10733004 |  | on10 | 00630414245751 | OUS | Oncology | No | Atellica IM Calcitonin (CALCT) | Atellica IM CALCT | CALCT | 100 |  | Atellica IM PW3 (SMN 10995666) (50.0 mL/pack ) | Atellica IM Multi-Diluent 13 | 38 | 28 | Recombinant anti-calcitonin antibody | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | 4 hours at RT, 24 hours at 2–8°C, up to 3 weeks at ≤ -20°C, up to 11 weeks at ≤ -70°C only 1 freeze/thaw cycle | 100 μL | No |  | 2-step sandwich immunoassay | 1.89–2000.00 pg/mL | 0.55–585.20 pmol/L |  |  |  |  | Males: ≤ 12.69 pg/mL ( ≤ 3.71 pmol/L) | Females: ≤ 5.89 pg/mL ( ≤ 1.72 pmol/L) |  | Atellica IM CALCT CAL | Included in the kit | Recon at RT 4 hours, recon at 2–8°C 3 days, recon ≤ -20°C 6 months thaw only once | Lyophilized | Recon with 2.0 mL of special reagent water. Invert vials 5 times, let stand for 30 minutes at RT | 3 for lot calibration, 2 for pack recalibration | 2 | Synthetic calcitonin antigen; casein perservative | 28 | 14 |  |  |  | Yes | Atellica IM Multi-Diluent 13 | 10995643 |  | 1.89 - 200,000.00 pg/mL | 0.55 - 58,520.00 pmol/L | No | [6 empty] | 1.89 pg/mL | 1.60 pg/mL | 1.89 pg/mL | 0.47 pmol/L | 0.55 pmol/L | 0.55 pmol/L | 5.47 pg/mL | 1.60 pmol/L | 3.8 | 5.7 | 346.65 pg/mL | 101.43 pmol/L | 2.2 | 3.2 | 1435.74 pg/mL | 420.10 pmol/L | 2.1 | 3.6 | ADVIA Centaur CALCT y = 0.95x + 0.74 pg/mL r= 0.99 | ADVIA Centaur CALCT y = 0.95x + 0.22 pmol/L r= 0.99 | ≤ 10% interference up to 500 mg/dL | ≤ 10% interference up to 40 mg/dL of conjugated bilirubin and up to 60 mg/dL of unconjugated bilirubin | ≤ 10% interference up to 3000 mg/dL (Intralipid) | ≤ 10% up to 40 ng/mL of biotin | 2,400,000 pg/mL | 702,240 pmol/L |  |  | WHO 2nd International Reference Preparation for Calcitonin (Human); NIBSC code: 89/620 | RPBL1331R01_EN Rev. 03, 2020-06 | IFU Document | 0.275 | 6.969 | 2.48 | 2.913 | 0.125 | 17.7 | 6.3 | 7.4
Atellica® IM | 10733001 |  | an05 | 00630414245720 | Global | Anemia | No | Atellica® IM Active-B12 | Atellica® IM AB12AB12 | AB12 | 100 | 1 x 100 | Atellica IM APW1 (probe wash) | Atellica IM AB12 MCM (master curve material) | 38 | 28 | Anti-transcobalamin (TC) monoclonal antibody and anti-holoTC monoclonal antibody | ;#Serum;# | 16 hours at RT, 3 days at 2–8°C. For longer storage, specimens may be frozen for up to 3 months at ≤ -20°C | 50 µL | No |  | 2-step sandwich immunoassay using direct | 4.25 - 146.00 pmol/L | 4.25 - 146.00 pmol/L | 4.25 - 292.00 pmol/L | 4.25 - 292.00 pmol/L | 27.24–169.62 pmol/L | 27.24–169.62 pmol/L |  |  |  | Atellica IM AB12 CAL | N/A (Kitted) | 8 hours at RT and 30 days at 2-8°C | Liquid | N/A | 3 for lot calibration, 2 for pack recalibration | 2 | Recombinant holotranscobalamin; BSA; sodium azide | 28 | 28 |  |  |  | Yes | Atellica IM Multi-Diluent 13 | 10995643 |  | 4.25 - 292.00 pmol/L | 4.25 - 292.00 pmol/L | Yes | Atellica IM AB12 QC | 10733002 | N/A | N/A | N/A | N/A | 0.83 pmol/L | 0.46 pmol/L | 4.25 pmol/L | 0.46 pmol/L | 0.83 pmol/L | 4.25 pmol/L | 19.18 pmol/L | 19.18 pmol/L | 2.7 | 4.9 | 57.37 pmol/L | 57.37 pmol/L | 2.5 | 4.8 | 108.29 pmol/L | 108.29 pmol/L | 2.1 | 4.5 | Atellica IM AB12 assay was compared to the ADVIA Centaur AB12 assay y = 1.05x - 1.21pmol/L r = 0.98 | Atellica IM AB12 assay was compared to the ADVIA Centaur AB12 assay y = 1.05x - 1.21pmol/L r = 0.98 | ≤10% change up to 500 mg/dL of hemoglobin | ≤10% change up to 40 mg/dL of bilirubin | ≤10% change up to 1000 mg/dL of triglycerides | ≤10% change up to 1,000,000 ng/mL of biotin | 91,200 pmol/L | 91,200 pmol/L | 91,200 pmol/L | 91,200 pmol/L | WHO NIBSC 03/178 | RPBL1359R2_EN Rev. 02, 2019-10 | IFU Document | 0.18 | 6.97 | 2.91 | 2.48 | 0.08 | 17.7 | 7.4 | 6.3
Atellica® IM | 10733006 |  | an04 | 00630414245775 | OUS | Anemia | No | Atellica® IM Erythropoietin | Atellica® IM EPOEPO | EPO | 100 | 1 x 100 | N/A | N/A | 19 | 28 | Monoclonal mouse anti-EPO antibody and Anti-EPO mouse monoclonal antibody | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#EDTA Plasma;# | 8 hours at RT, 7 days at 2–8°C. For longer storage, specimens may be frozen for up to 3 months at ≤ -20°C. | 100 | No |  | One-step sandwich immunoassay using direct | 0.98 - 750 mIU/mL | 0.98 - 750 mIU/mL | 0.98 - 7500 mIU/mL | 0.98 - 7500 mIU/mL | 5.45–28.35 mIU/mL | 5.45–28.35 mIU/mL |  |  |  | Atellica IM EPO CAL | N/A (Kitted) | 24 hours at RT or 90 days at 2-8°C | Liquid | N/A | 3 for lot calibration, 2 for pack recalibration | 2 | Recombinant human EPO; calf serum | 28 | 14 |  |  |  | Yes | Atellica IM Multi-Diluent 13 | 10995643 |  | 0.98 - 7500 mIU/mL | 0.98 - 7500 mIU/mL | Yes | Atellica IM EPO QC | 10733008 | N/A | N/A | N/A | N/A | 0.98 mIU/mL | 0.69 mIU/mL | 0.98 mIU/mL | N/A | N/A | N/A | 2.30 mIU/mL | 2.30 mIU/mL | 4.6 | 5.7 | 29.74 mIU/mL | 29.74 mIU/mL | 1.4 | 3.1 | 612.92 | 612.92 | 1.4 | 2.7 | Atellica IM EPO was compared to ADVIA Centaur EPO: y = 0.94x + 0.58 mIU/mL r = 1.00 | Atellica IM EPO was compared to ADVIA Centaur EPO: y = 0.94x + 0.58 mIU/mL r = 1.00 | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 40 mg/dL of bilirubin | ≤ 10% change up to 3000 mg/dL (intralipid) | ≤ 10% change up to 1,000,000 ng/mL of biotin | 192,334 mIU/mL | 192,334 mIU/mL | 192,334 mIU/mL | 192,334 mIU/mL | WHO 2nd IRP 67/343 and WHO 3rd IS 11/170 | RPBL1337R03_EN Rev. 03, 2019-04 | IFU Document | 0.28 | 6.97 | 2.91 | 2.48 | 0.126 | 17.7 | 7.4 | 6.3
Atellica® IM | 10732998 |  | au01 |  | OUS | Autoimmune | No | Atellica IM anti-CCP IgG (aCCP) | Atellica IM aCCP | aCCP | 100 | 1X100 |  | Atellica IM aCCP MCM SMN 10733000 | 33 | 35 | Monoclonal anti-human IgG (0.1 μg/mL) | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | 22 hrs at RT, 7 days at 2–8°C, > 7days store at -20°C or colder, avoid > 2 freeze/thaw cycles | 15 μL | No |  | 2-step sandwich immunoassay | 0.54–200.00 U/mL | NA |  |  |  |  | Healthly individuals median 0.74 U/mL | Samples ≤ 5.00 U/mL indicate that anti-CCP IgG antibodies were not detected | Samples ≥5.00 U/mL indicate that anti-CCP IgG antibodies were detected | Atellica IM aCCP CAL | Included in the kit | 2–8°C 90 days after opening, At room temperature 8 hours | Liquid | NA | 3 for lot calibration, 2 for pack recalibration | 2 | Processed human plasma positive for anti-CCP antibodies | 35 | 14 |  |  |  | Yes | Atellica IM Multi-Diluent 1 | 10995639 |  | 0.54–2000.00 U/mL | NA | Yes | Atellica IM aCCP QC | 10732999 | NA | NA | NA | NA | 0.54 U/mL | 0.20 U/mL | NA | NA | NA | NA | 4.44U/mL | NA | 4.4 | 4.7 | 33.18 U/mL | NA | 3.6 | 4.2 | 107.14 U/mL | NA | 4.5 | 4.7 | NA | NA | ≤ 10% change up to 1000 mg/dL of hemoglobin | ≤ 10% change up to 40 mg/dL of conjugated bilirubin and unconjugated bilirubin | ≤ 10% change up to 1500 mg/dL of Intralipid | ≤ 10% change up to 1500 ng/dL | 2145 U/mL | NA |  |  | Traceable to an internal standard | RPBL1372R02_EN Rev. 04, 2019-10 | IFU Document
Atellica® IM | 11202207 |  | id17 | 00630414246499 | OUS | ToRCH | No | Atellica IM CMV IgG (CMV IgG) | Atellica IM CMV IgG | CMVIgG | 100 | 1 x 100 | Atellica IM APW1 (10995458) | Atellica IM CMV IgG QC (11202209) | 40 min | 30 | anti‑human IgG mouse monoclonal antibody | ;#Serum;#Lithium Heparin Plasma;#Potassium EDTA Plasma;# | On the clot for up to 14 days at 2–8°C. Separated specimens: 48 hrs RT, 7 days 2–8°C. Frozen 2 months at -20°C to -70°C. Do not store in a frost-free freezer. | 20 μL | No |  | 2‑step sandwich immunoassay using indirect chemiluminescent technology | 0.05–30.00 Index | [5 empty] | Nonreactive / Reactive / No equivocal zone |  |  | Atellica IM CMV IgG CAL | Included | 60 days after opening at 2–8°C | Liquid | N/A | 3 for lot calibration, 2 for pack recalibration | 2 |  | 14 | 14 |  |  |  | No | [5 empty] | Yes | Atellica IM CMV IgG QC (quality control | 11202209 | Initial relative sensitivity 99.5% | Initial relative specificity 98.5% | Resoloved relative sensitivity 100.0% | Resoloved relative specificity 99.8% | N/A | N/A | N/A | N/A | N/A | N/A | 0.01 |  | 1.3 | 2.9 | 1.18 |  | 1.5 | 3.4 | 9.17 |  | 1.3 | 3 |  |  | 500 mg/dL of hemoglobin | 20 mg/dL of unconjugated bilirubin | 3000 mg/dL of intralipids (Triglycerides) | 4500 ng/mL of biotin |  |  |  |  | The Atellica IM CMV IgG assay standardization is based upon clinical agreement with a commercially available CMV IgG assay. | 11200619_EN Rev. 03, 2019-08 | IFU Document
Atellica® CH | 11097536 |  | di02 | 00630414220505 | Global | Diabetes | No | Atellica CH Enzymatic Hemoglobin A1c (A1c_E) | A1c_EA1c_E | 4039 | 600 | 2 x 300 | N/A | N/A | 10 | 63 | Peroxidase (Horseradish); fructosyl peptide oxidase (E. coli, recombinant); ofloxacin | ;#Whole blood;#Hemolysate;# | 48 hr at RT, 7 days at 2–8°C, or stored frozen for up to 21 months (with one freeze-thaw) at -70°C. | 4.5 μL | No |  | Enzymatic | 3.80% - 14.00% HbA1c | 18.03–129.50 mmol/mol HbA1c | N/A | N/A | < 5.7%: Normal 5.7–6.4%: Prediabetes ≥ 6.5%: Diabetic | < 39 mmol/mol: Normal 39–47 mmol/mol: Prediabetes ≥ 48 mmol/mol: Diabetic | N/A | N/A | N/A | Atellica CH A1c_E CAL | 11099338 | 11 days | Lyophilized | Level 1 is liquid and ready to use. Level 2 and 3: Add 1.0 mL of deionized water into each vial, let the vials stand for 30 minutes at RT to allow the lyophilized material to dissolve. Gently swirl and invert the vials. | 3 | 3 | Level 1: Horseradish peroxidase. Levels 2 and 3: human whole blood | 180 | 63 |  | No |  | No | N/A | 11099300 |  | N/A | N/A | No | N/A | N/A | NA | NA | NA | NA | 3.53% | 3.11% | N/A | 61.38 | 65.54 | N/A | 5.28 % | N/A | 0.5 | 1.6 | 6.49 % | N/A | 0.4 | 1.4 | 11.79 % | N/A | 0.3 | 1.2 | NGSP Reference Method y = 0.986x - 0.030 (%HbA1c) r=0.995 n=172 | NGSP Reference Method y = 0.986x - 0.664 (mmol/mol) r=0.995 n=172 | N/A | < 5% change up to 10 mg/dL of Bilirubin, conjugated and unconjugated | < 5% change up to 1000 mg/dL of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | IFCC reference calibrators | 11111592_EN Rev. 01, 2019-01 | IFU Document | 0.6 | 4.134 | 4.134 | 2.756 | 0.272 | 10.5 | 10.5 | 7
Atellica® IM | 10733019 |  | re11 | 00630414245904 | OUS | Reproductive Endocrinology | No | Atellica IM Pregnancy-Associated Plasma Protein-A (PAPP-A) | Atellica IM PAPP-A | PAPPA | 100 | 1 x 100 | NA | Atellica IM Multi-Diluent 13 | 33 | 28 | Anti-PAPP-A mouse monoclonal antibody labeled with acridinium ester | ;#Serum;# | 8 hours at room temperature (RT), 72 hours at 2–8°C, up to 6 months at -20°C Avoid >1 freeze/thaw cycle | 20 μL | No |  | 2-step sandwich immunoassay using direct chemiluminescent technology | 0.01–10.00 IU/L | 0.01–10.00 mIU/mL |  |  |  |  | < 0.01 IU/L (mIU/mL) healthy non-pregnant females | NA | NA | Atellica IM PAPP-A CAL | NA | 8 hrs at room temperature (RT), 30 days at 2–8°C | Liquid | NA | 3 for lot calibration, 2 for pack recalibration | 2 | Human serum; bovine serum | 28 | 14 |  |  |  | Yes | Atellica IM Multi-Diluent 13 | 10995643 |  | 0.01–100.00 IU/L | 0.01–100.00 mIU/mL | No |  |  | NA | NA | NA | NA | 0.005 IU/L | 0.004 IU/L | 0.01 IU/L | 0.004 mIU/mL | 0.005 mIU/mL | 0.01 mIU/mL | 0.32 IU/L | 0.32 mIU/mL | 2.7 | 3.4 | 4.17 IU/L | 4.17 mIU/mL | 2.9 | 4.1 | 9.03 IU/L | 9.03 mIU/mL | 3.6 | 4.4 | ADVIA Centaur PAPP-A y = 0.96x + 0.00 IU/L r = 1.00 | ADVIA Centaur PAPP-A y = 0.96x + 0.00 mIU/mL r = 1.00 | ≤ 10% interference up to 500 mg/dL (0.31 mmol/L) | ≤ 10% interference up to 40 mg/dL (684 μmol/L) conjugated and 60 mg/dL (1026 μmol/L) unconjugated | ≤ 10% interference up to 1000 mg/dL (11.3 mmol/L) | Not used in assay structure | 157 IU/L | 157 mIU/mL |  |  | Standardized to an internal standard | RPBL1308R03_EN Rev. 03, 2018-09 | IFU Document
Atellica® IM | 10733020 |  | re11 | 00630414245911 | OUS | Reproductive Endocrinology | No | Atellica IM Pregnancy-Associated Plasma Protein-A (PAPP-A) | Atellica IM PAPP-A | PAPPA | 500 | 5 X 100 |  | Atellica IM Multi-Diluent 13 | 33 | 28 | Anti-PAPP-A mouse monoclonal antibody labeled with acridinium ester | ;#Serum;# | 8 hours at room temperature (RT), 72 hours at 2–8°C, up to 6 months at -20°C Avoid >1 freeze/thaw cycle | 20 μL | No |  | 2-step sandwich immunoassay using direct chemiluminescent technology | 0.01–10.00 IU/L | 0.01–10.00 mIU/mL |  |  |  |  | < 0.01 IU/L (mIU/mL) healthy non-pregnant females | NA | NA | Atellica IM PAPP-A CAL | NA | 8 hrs at room temperature (RT), 30 days at 2–8°C | Liquid | NA | 3 for lot calibration, 2 for pack recalibration | 2 | Human serum; bovine serum | 28 | 14 |  |  |  | Yes | Atellica IM Multi-Diluent 13 | 10995643 |  | 0.01–100.00 IU/L | 0.01–100.00 mIU/mL | No |  |  | NA | NA | NA | NA | 0.005 IU/L | 0.004 IU/L | 0.01 IU/L | 0.004 mIU/mL | 0.005 mIU/mL | 0.01 mIU/mL | 0.32 IU/L | 0.32 mIU/mL | 2.7 | 3.4 | 4.17 IU/L | 4.17 mIU/mL | 2.9 | 4.1 | 9.03 IU/L | 9.03 mIU/mL | 3.6 | 4.4 | ADVIA Centaur PAPP-A y = 0.96x + 0.00 IU/L r = 1.00 | ADVIA Centaur PAPP-A y = 0.96x + 0.00 mIU/mL r = 1.00 | ≤ 10% interference up to 500 mg/dL (0.31 mmol/L) | ≤ 10% interference up to 40 mg/dL (684 μmol/L) conjugated and 60 mg/dL (1026 μmol/L) unconjugated | ≤ 10% interference up to 1000 mg/dL (11.3 mmol/L) | Not used in assay structure | 157 IU/L | 157 mIU/mL |  |  | Standardized to an internal standard | RPBL1308R03_EN Rev. 03, 2018-09 | IFU Document
Atellica® IM | 10733009 |  | re03 | 00630414245805 | OUS | Reproductive Endocrinology | No | Atellica IM Free Beta Human Chorionic Gonadotropin (FBHCG) | Atellica IM FBHCG | FBHCG | 100 | 1 X 100 | NA | Atellica IM Multi-Diluent 13 | 40 | 28 | Anti-free beta HCG mouse monoclonal antibody labeled | ;#Serum;# | 24 hours at room temperature, up to 6 days at 2–8°C, frozen for up to 6 months at -20°C. Avoid >1 freeze/thaw cycle | 20 μL | No |  | 2-step sandwich immunoassay using direct chemiluminescent technology | 0.14–200.00 IU/L | 0.14–200.00 ng/mL |  |  |  |  | < 0.28 IU/L (ng/mL) healthy non-pregnant females | NA | NA | Atellica IM FBHCG CAL | Included in the kit | Reconstituted at room temperature 8 hours, reconstituted at 2–8°C 30 days | Lyophilized | 2.0 mL of special reagent water. Invert 5 times and let stand for 30 minutes at room temperature | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, human free beta HCG; BSA | 28 | 14 |  |  |  | Yes | Atellica IM Multi-Diluent 13 | 10995643 |  | 0.14–2000.00 IU/L | 0.14–2000.00 ng/mL | No |  |  | NA | NA | NA | NA | 0.11 IU/L | 0.07 IU/L | 0.14 IU/L | 0.07 ng/mL | 0.11 ng/mL | 0.14 ng/mL | 8.31 IU/L | 8.31 ng/mL | 1.5 | 2.3 | 81.17 IU/L | 81.17 ng/mL | 1.3 | 2.6 | 177.70 IU/L | 177.70 ng/mL | 1.3 | 2.3 | ADVIA Centaur FBHCG y = 0.95x - 0.98 IU/L r = 1.00 | ADVIA Centaur FBHCG y = 0.95x - 0.98 ng/mL r = 1.00 | ≤ 10% interference up to 500 mg/dL (0.31 mmol/L) | ≤ 10% interference up to 40 mg/dL (684 μmol/L) conjugated, 60 mg/dL (1026 μmol/L) unconjugated | ≤ 10% interference up to 1000 mg/dL (11.3 mmol/L) | ≤ 10% interference up to 1500 ng/mL (6140 nmol/L) | 78,940 IU/L | 78,940 ng/mL |  |  | WHO 1st International Reference 75/551 | RPBL1319R04_EN Rev. 04, 2019-12 | IFU Document
Atellica® IM | 10733010 |  | re03 | 00630414245812 | OUS | Reproductive Endocrinology | No | Atellica IM Free Beta Human Chorionic Gonadotropin (FBHCG) | Atellica IM FBHCG | FBHCG | 500 | 5 x 100 | NA | Atellica IM Multi-Diluent 13 | 40 | 28 | Anti-free beta HCG mouse monoclonal antibody labeled | ;#Serum;# | 24 hours at room temperature, up to 6 days at 2–8°C, frozen for up to 6 months at -20°C. Avoid >1 freeze/thaw cycle | 20 μL | No |  | 2-step sandwich immunoassay using direct chemiluminescent technology | 0.14–200.00 IU/L | 0.14–200.00 ng/mL |  |  |  |  | < 0.28 IU/L (ng/mL) healthy non-pregnant females | NA | NA | Atellica IM FBHCG CAL | Included in the kit | Reconstituted at room temperature 8 hours, reconstituted at 2–8°C 30 days | Lyophilized | 2.0 mL of special reagent water. Invert 5 times and let stand for 30 minutes at room temperature | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, human free beta HCG; BSA | 28 | 14 |  |  |  | Yes | Atellica IM Multi-Diluent 13 | 10995643 |  | 0.14–2000.00 IU/L | 0.14–2000.00 ng/mL | No |  |  | NA | NA | NA | NA | 0.11 IU/L | 0.07 IU/L | 0.14 IU/L | 0.07 ng/mL | 0.11 ng/mL | 0.14 ng/mL | 8.31 IU/L | 8.31 ng/mL | 1.5 | 2.3 | 81.17 IU/L | 81.17 ng/mL | 1.3 | 2.6 | 177.70 IU/L | 177.70 ng/mL | 1.3 | 2.3 | ADVIA Centaur FBHCG y = 0.95x - 0.98 IU/L r = 1.00 | ADVIA Centaur FBHCG y = 0.95x - 0.98 ng/mL r = 1.00 | ≤ 10% interference up to 500 mg/dL (0.31 mmol/L) | ≤ 10% interference up to 40 mg/dL (684 μmol/L) conjugated, 60 mg/dL (1026 μmol/L) unconjugated | ≤ 10% interference up to 1000 mg/dL (11.3 mmol/L) | ≤ 10% interference up to 1500 ng/mL (6140 nmol/L) | 78,940 IU/L | 78,940 ng/mL |  |  | WHO 1st International Reference 75/551 | RPBL1319R04_EN Rev. 04, 2019-12 | IFU Document
Atellica® CH | 10329216 |  | ge46 | 08058056680150 | OUS | General Chemistry | Yes | Aldolase (Sentinel Alliance Application) | AldoAldo | 7068 | 668 | 4 x 167 | Atellica CH Empty Reagent Packs (11097534) | N/A | 8.9 | 30 per pack | collidin buffer, mono-iodoacetate, F-1.6-DP, sodium azide, NADH, GDH, TPI, LDH and ammonium sulphate | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | 15 days at 2-8 °C | 25 µL | No |  | Colorimetric | 1.0 - 30.0 U/L. | 1.0 - 30.0 U/L. | N/A | N/A | Serum/Plasma: ≤ 7.6 U/L (37 °C); ≤ 3.1 U/L (25 °C) | Serum/Plasma: ≤ 7.6 U/L (37 °C); ≤ 3.1 U/L (25 °C) | N/A | N/A | N/A | Clin Chem Calibrator (Sentinel) | 10313439 |  | Lyophilized | Add 3.0 mL of distilled water. Leave to stand for 30 minutes. Gently invert to ensure the homogeneity of the solution. | 3 | 2 | human based serum stabilized by adding preservatives | 30 per pack | 30 per pack |  | No |  | Yes | N/A | 11099300 |  | N/A | N/A | No | N/A | N/A | 1.0 U/L | 1.0 U/L | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 7.0 U/L | 7.0 U/L | 4.3 | 4.8 | 16.0 U/L | 16.0 U/L | 2.4 | 2.9 | 30.0 U/L | 30.0 U/L | 2.5 | 2.9 | Sentinel Aldolase on BC AU680 y=0.9641x + 0.0855 U/L r = 0.999 n=126 | Sentinel Aldolase on BC AU680 y=0.9641x + 0.0855 U/L r = 0.999 n=126 | Avoid hemolysed samples | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Internal standard | IFU: 17003-2.0/ 02 2015/07/16 and App. Sheet: 17003-1.0/02 F57 2019/04/24 |  | 0.6614 | 7.99 | 3.23 | 4.69 | 0.3 | 20.3 | 8.2 | 11.9
Atellica® CH | 11419175 |  | ge47 | 08058056682000 | OUS | General Chemistry | Yes | Total Bile Acids Liquid (Sentinel Alliance Application) | TBATBA | 7066 | 334 | 2 x 167 | Atellica CH Empty Reagent Packs (SMN 11097534) Sentinel Total Bile Acids Control Set (SMN 11532786) | N/A | 8.5 | 30 per pack | Thio-NAD⁺, Detergent, Sodium Azide, Good’s Buffer, NADH, 3α-HSD, Buffer, Conjugate Bile Acid | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#Potassium EDTA Plasma;# | Serum: 1 day at 20 °C, 1 weeks at 2-8 °C or 12 months at -20 °C. Plasma: 6 hr at RT, 48 hr at 2-8 °C, 2 months at -20 °C or 2 months at -70 °C. | 12 µL | No |  | Enzymatic Cycling Colorimetric | 1.0 – 180 μmol/L | 1.0 – 180 μmol/L | N/A | N/A | Male (fasting): 1.0 - 6.0 μmol/L; Female (fasting): 1.0 - 6.0 μmol/L | Male (fasting): 1.0 - 6.0 μmol/L; Female (fasting): 1.0 - 6.0 μmol/L | N/A | N/A | N/A | Total Bile Acids Standard | included in the kit | 28 | Liquid | N/A | 3 | 2 | sodium taurodeoxycholate (Sorachim cat # TAU-100) | 30 per pack | 30 per pack |  | No |  | No | N/A |  |  | N/A | N/A | Yes | Total Bile Acids Control Set | 11532786 | NA | NA | NA | NA | 0.19 μmol/L | 0.10 μmol/L | 0.4 μmol/L | 0.10 μmol/L | 0.19 μmol/L | 0.4 μmol/L | 3.0 µmol/L | 3.0 µmol/L | 2.1 | 3.2 | 18.0 µmol/L | 18.0 µmol/L | 0.8 | 1.3 | 180.0 µmol/L | 180.0 µmol/L | 0.7 | 1.2 | Randox TBA on ARCHITECT c8000 y = 1.07 x + 0.3032 µmol/L r = 0.999 n=144 | Randox TBA on ARCHITECT c8000 y = 1.07 x + 0.3032 µmol/L r = 0.999 n=144 | ≤ 10% change (or ± 0.6 μmol/L for samples < 6 μmol/L) up to 1000 mg/dL of hemoglobin | ≤ 10% change (or ± 0.6 μmol/L for samples < 6 μmol/L) up to 22.4 mg/dL of conjugated bilirubin and up to 29.9 mg/dL of unconjugated bilirubin | ≤ 10% change (or ± 0.6 μmol/L for samples < 6 μmol/L) up to 1703 mg/dL of Trig and up to 750 mg/dL of Intralipid |  | N/A | N/A | N/A | N/A | Internal Master Calibrator | IFU: 1700201 - 1.0/ 02 2018/04/13 and App sheet: F57 2019/04/24 |  | 0.833 | 4.37 | 2.42 | 4.68 | 0.378 | 11.1 | 6.15 | 11.9
Atellica® CH | 11533403 |  | ca03 | 08058056680549 | OUS | General Chemistry | Yes | CK-MB Liquid (Sentinel Alliance Application) | eCKMBeCKMB | 7067 | 320 | 1 x 320 | Atellica CH Empty Reagent Packs (SMN 11097534) | N/A | 8.5 | 30 per pack | monoclonal murine antibodies anti-CK-M | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | 3 days at 2-8 °C; 4 weeks at -20°C | 25 µL | No |  | Enzymatic | 3.0 - 1000 U/L | 3.0 - 1000 U/L. | 3.0 - 10000 U/L | 3.0 - 10000 U/L | ≤ 24 U/L | ≤ 24 U/L | N/A | N/A | N/A | CK-MB Calibrator (Sentinel) | 10284952 | 2 | Lyophilized | Add 1.0 mL of distilled water. Replace stopper and gently swirl. Allow reconstituted material to stand for at least 60 minutes. Gently invert just prior to use. Avoid foaming. | 3 | 2 | Human CK-MM and porcine brain CK-BB in bovine albumin matrix. | 30 per pack | 30 per pack |  | No |  | Yes | N/A | 11099300 |  | 3.0 - 10000 U/L | 3.0 - 10000 U/L | No | N/A | N/A | 3.0 U/L | 3.0 U/L | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 20 U/L | 20 U/L | 4.7 | 7 | 300 U/L | 300 U/L | 0.6 | 1.5 | 1000 U/L | 1000 U/L | 0.2 | 1.4 | Architect c16000 CK-MB y = 0.94 x + 2 U/L r = 1.000 n = 120 | Architect c16000 CK-MB y = 0.94 x + 2 U/L r = 1.000 n = 120 | ≤ 10% change up to 50 mg/dL of hemoglobin | ≤ 10% change up to 50 mg/dL of bilirubin | ≤ 10% change up to 600 mg/dL of intralipid | N/A | N/A | N/A | N/A | N/A | Inhibition method according to IFCC mod. recommendations | IFU: 17649H - 2.0/ 02 2015/04/14; App. Sheet: F57 2019/04/24 |  | 0.364 | 4.37 | 2.42 | 4.69 | 0.165 | 11.1 | 6.15 | 11.9
Atellica® CH | 10317329 |  | ge49 | 08058056680013 | OUS | General Chemistry | Yes | Zinc (Sentinel Alliance Application) | ZnZn | 7086 | 308 | 2 x 154 | Atellica CH Empty Reagent Packs (11097534) | N/A | 9.7 | 34 | good's buffer, specific masking agents for iron and copper, ascorbic acid, complexant 5-Br-PAPS | ;#Serum;#Heparinized Plasma;# | 14 days at 2-8 °C | 25 | No |  | Colorimetric | 5.0 – 2000 μg/dL | 5.0 – 2000 μg/dL | N/A | N/A | Adults; 70 - 120 μg/dL | Adults; 70 - 120 μg/dL | N/A | N/A | N/A | Zinc Standard | included in the kit |  | Liquid | N/A | 3 | 2 |  | 41 | 34 |  | No |  | Yes | N/A | 11099300 |  | N/A | N/A | No | N/A | N/A | 5.0 μg/dL | 5.0 μg/dL | NA | NA | N/A | N/A | N/A | N/A | N/A | N/A | 116.55 μg/dL | 116.55 μg/dL | 0.6 | 1.1 | 540.97 μg/dL | 540.97 μg/dL | 0.2 | 0.4 | 1068.23 μg/dL | 1068.23 μg/dL | 0.2 | 0.4 | Randox TBA on ARCHITECT c8000 y = 1.07 x + 0.3032 µmol/L r = 0.999 n=144 | Randox TBA on ARCHITECT c8000 y = 1.07 x + 0.3032 µmol/L r = 0.999 n=144 | ≤ 10% change up to 500 mg/dL of hemoglobin. | ≤ 10% change up to 40 mg/dL of bilirubin. | ≤ 10% change up to 1000 mg/dL of Trig. |  | N/A | N/A | N/A | N/A | NIST (National Institute of Standards and Technology). | IFU: 17640H - 2.0/ 02 2015/04/15. App Sheet: 17640H-1.0/02 F57 2019/07/25 |  | 0.708 | 5.59 | 3.23 | 4.69 | 0.321 | 14.2 | 8.2 | 11.9
Atellica® CH | 11533083 |  | ge48 | 08058056680020 | OUS | General Chemistry | Yes | Copper (Sentinel Alliance Application) | CopCop | 8011 | 567 | 3 x 189 | Atellica CH Empty Reagent Packs (11097534) | N/A | 9.7 | 30 | malate dehydrogenase (MDH) (mammalian); PEP: microorganism (PEPC); | ;#Serum;#Heparinized Plasma;# | 14 days at 2-8 °C | 20 | No |  | Colorimetric | 3.0 – 500 μg/dL | 3.0 – 500 μg/dL | N/A | N/A | Male: 70 - 140 μg/dL; Female : 80 - 155 μg/dL | Male: 70 - 140 μg/dL; Female : 80 - 155 μg/dL | N/A | N/A | N/A | Copper Standard | included in the kit |  | Liquid | N/A | 3 | 2 |  | 41 | 30 |  | No |  | Yes | N/A | 11099300 |  | N/A | N/A | No | N/A | N/A | 3.0 μg/dL | 3.0 μg/dL | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 97.59 μg/dL | 97.59 μg/dL | 1.4 | 2.4 | 317.70 μg/dL | 317.70 μg/dL | 0.5 | 0.7 | 414.32 μg/dL | 414.32 μg/dL | 0.3 | 0.6 | Beckman AU5800 y = 1.030x -5.466 μg/dL r = 0.999 n = 110 | Beckman AU5800 y = 1.030x -5.466 μg/dL r = 0.999 n = 110 | ≤ 10% change up to 300 mg/dL of hemoglobin | ≤ 10% change up to 40 mg/dL of bilirubin | ≤ 10% change up to 1500 mg/dL of lipids | N/A | N/A | N/A | N/A | N/A | NIST (National Institute of Standards and Technology) | IFU: 17638H - 2.0/ 02 2015/04/21. App. Sheet: 17638H-1.0/02 F57 2019/07/25 | IFU Document | 0.745 | 5.59 | 3.23 | 4.69 | 0.34 | 14.2 | 8.2 | 11.9
Atellica® CH | 11309910 |  | ge50 | 08058056681751 | Global | General Chemistry | Yes | Angiotensin Converting Enzyme (Sentinel Alliance Application) | ACEACE | 7069 | 308 | 2 x 154 | Atellica CH Empty Reagent Packs ( SMN 11097534) | N/A | 9.7 | 30 | Human specimens | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;# | 7 days at 2-8 °C or 1 year at -20 °C | 100 | No |  | Enzymatic | 8.4 - 120 U/L | 8.4 - 120 U/L | 8.4 - 1200 U/L | 8.4 - 1200 U/L | 13.3 – 63.9 U/L | 13.3 – 63.9 U/L | N/A | N/A | N/A | Sentinel ACE Calibrator | 11309911 | 7 | Lyophilized | Add 1.0 mL of distilled or deionized water. Leave to stand for 5 minutes. Swirl each vial just prior to use. | 3 | 2 | Human serum containing porcine ACE | 134 | 30 |  | No |  | Yes | N/A | N/A |  | 8.4 - 1200 U/L | 8.4 - 1200 U/L | Yes | Sentinel ACE Control Set | 11309912 | N/A | N/A | N/A | N/A | 4.5 U/L | 2.4 U/L | 8.4 U/L | 2.4 U/L | 4.5 U/L | 8.4 U/L | 20 U/L | 20 U/L | 7.1 | 8.6 | 75 U/L | 75 U/L | 1.4 | 4.7 | 120 U/L | 120 U/L | 1.1 | 4.4 | Sentinel ACE on Atellica CH vs. Beckman AU680 y = 0.91x + 4.27 r = 0.984 n = 115 | Sentinel ACE on Atellica CH vs. Beckman AU680 y = 0.91x + 4.27 r = 0.984 n = 115 | ≤ 10% change up to 100 mg/dL of hemoglobin | ≤ 10% change up to 13 mg/dL of unconjugated bilirubin; ≤ 10% change up to 26 mg/dL of conjugated bilirubin | ≤ 10% change up to 200 mg/dL of Trig | N/A | N/A | N/A | N/A | N/A | Internal master calibrator material. | IFU: 1767001-3.0/ 02 2020/03/04. App Sheet: 1767001-2.0/02 F57 2020/05/20 |  | 0.3 | 4.37 | 2.42 | 4.68 | 0.135 | 11.1 | 6.15 | 11.9
Atellica® IM | 10995648 |  | ca07 | 00630414599663 | Global | Cardiac | No | Atellica IM Myoglobin (MYO) | Atellica IM MYO | MYO | 250 | 1 x 100 | Atellica IM CAL U (SMN 10995519) | Atellica IM Multi-Diluent 10 (SMN 10995640) Atellica IM MYO MCM (SMN 10995650) | 10 | 28 | Goat polyclonal anti-myoglobin antibody/mouse monoclonal anti-myoglobin antibody | ;#Serum;#Heparinized Plasma;# | Up to 8 hrs at RT. 48hr 2-8C, freeze 1 time | 10 uL | No |  | direct chemiluminescent technology | 3.00 - 1000.00 ng/mL | 3.00 - 1000.00 ug/L |  |  |  |  | A myoglobin value of >110 ng/mL (ug/L) is highly suggestive of myocardial infarct |  |  | Atellica IM CAL U | 10995519 (2-pack) | 4 hrs at RT, 24 hrs recon at 2-8C, 60 days recon at -80 to -60C | Lyophilized | reconstitute 2.0 mL/Vial | 3 for lot calibration, 2 for pack recalibration | 2 |  | 33 | 14 |  |  |  | Yes | Atellica IM Multi-Diluent 10 | 10995640 |  | 3.00 - 20,000 ng/mL | 3.00 - 20,000.00 ug/L | No |  |  | 2.53 ng/mL | 2.53 ug/L | NA | NA | 3.57 ng/mL | 2.56 ng/mL | NA | 2.56 ug/L | 3.57 ng/mL | NA | 155.95 mg/dL (Control) | 155.95 ug/L (Control) | 2 | 2.9 | 361.51 mg/dL (Control) | 361.51 ug/L (Control) | 2.1 | 3 | 832.09 mg/dL | 832.09 ug/L | 2.2 | 2.9 | ADVIA Centaur Myoglobin y=0.96x + 2.80 ng/mL r=1.00 | ADVIA Centaur Myoglobin y=0.06x + 2.80 ug/L r=1.00 | < 5% change up to 1000 mg/dL of hemoglobin | < 5% change up to 40 mg/dL conjugated bilirubin 40 mg/dL of unconjugated bilirubin | <5% change up to 1000 mg/dL of triglycerides | No Biotin in assay formulation | 200,000 ng/mL | 200,000 ug/L |  |  | Traceable to an internal standard manufactured using highly purified material | 10995411_EN Rev. 02, 2019-07 | IFU Document | 0.176 | 6.969 | 2.087 | 1.457 | 0.08 | 17.7 | 5.3 | 3.7
Atellica® IM | 11203564 |  | id14 | 00630414285641 | OUS | ToRCH | No | Atellica IM Rubella G II (RbGII) | RbGII | RbGII | 100 |  | 10995458 Atellica IM APW1 (probe wash) | 11203576 Atellica IM RbGII QC (quality control material) 10995644 Atellica IM Multi‑Diluent 2 (diluent) 11203585 4 x 1.0 mL levels of master curve material | 42 | 60 | Mouse anti-human IgG monoclonal antibody | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#Potassium EDTA Plasma;# | 5 days at room temperature, and for up to 14 days at 2–8°C. Specimens devoid of red blood cells may be frozen for up to 7 months ≤ -20°C. Do not store in a frost-free freezer. | 15 μl | No |  | 2-step sandwich immunoassay that uses direct, chemiluminescent technology | 0.68–400.00 IU/mL | [5 empty] | Nonreactive: < 10.00 IU/mL; Reactive: ≥ 10.00 IU/mL |  |  | 1 vial Atellica IM RbGII CAL low calibrator 1 vial Atellica IM RbGII CAL high calibrator Atellica IM RbGII CAL lot-specific value sheet | Included in kit | 65 days 2–8°C, 8 hours room temperature | Liquid | N/A | 3 for lot calibration, 2 for pack recalibration | 2 | Processed human plasma negative and positive for anti-rubella IgG antibodies; preservatives | 70 | 60 |  |  |  | No | Atellica IM Multi‑Diluent 2 ReadyPack ancillary reagent pack | Included |  |  |  | Yes | Atellica IM RbGII QC (quality control material) | 11203576 | [10 empty] | 23.18 |  | 2.5 | 5.1 |  |  |  |  | 221.66 |  | 4.3 | 6.5 |  |  | ≤ 10% change in results up to 500 mg/dL of hemoglobin | ≤ 10% change in results up to 60 mg/dL of conjugated bilirubin, 40 mg/dL of unconjugated bilirubin | ≤ 10% change in results up to 1000 mg/dL of lipemia (Intralipid) | [5 empty] | Traceable to WHO 1st International Standard for human anti-rubella immunoglobulin (RUBI‑1‑94) | 11203565_EN Rev. 01, 2019-10 | IFU Document
Atellica® IM | 11200558 |  | re12 |  | Global | Reproductive Endocrinology | No | Atellica IM Androstenedione (ANDRO) | ANRO | ANDRO | 50 | 1x50 | NA | ANDRO Master Curve Material (SMN11200557) Atellica IM Multi-Diluent 1 | 14 | 30 | Biotinylated sheep monoclonal anti-androstenedione antibody | ;#Serum;#Lithium Heparin Plasma;#Potassium EDTA Plasma;# | up to 24 hours at room temperature, up to 31 days at 2–8°C, up to 12 months at ≤ -20°C. | 20 μL | No |  | 1-step, competitive immunoassay | 0.30 – 9.00 ng/mL | 1.05 – 31.42 nmol/L |  |  |  |  | Adult males < 0.30–1.65 ng/mL (1.05–5.76 nmol/L) (See IFU for pediatric ranges) | Females 22–39 years 0.37–2.25 ng/mL (1.29–7.86 nmol/L) | Females 40–54 years < 0.30–1.90 ng/mL (1.05–6.63 nmol/L) | ANDRO low and high calibrators | Included in the kit | 24 hrs recon at room temperature, 28 days recon at 2–8°C, 60 days recon at ≤ -20°C | Lyophilized | 2.0 mL of special reagent water, stand for 15–20 minutes at room temperature | 3 for lot calibration, 2 for pack recalibration | 2 | low and high levels of androstenedione | 90 | 30 |  |  |  | Yes | Atellica IM Multi‑Diluent 1 | 10995637 (2-pack),10995638 (6-pack),10995639 (vial) |  | 0.30 – 45.00 ng/mL | 1.05 – 157.10 nmol/L | No |  |  | NA | NA | NA | NA | 0.16 ng/mL | 0.13 ng/mL | 0.17 ng/mL | 0.45 nmol/L | 0.56 nmol/L | 0.59 nmol/L | 0.48 ng/mL | NA | 3.5 | 4.6 | 1.06 ng/mL | NA | 2.5 | 3.1 | 7.34 ng/mL | NA | 2.8 | 4 | ADVIA Centaur ANDRO assay (x) y = 0.93x + 0.03 ng/mL r= 0.992 | ADVIA Centaur ANDRO assay (x) y = 0.93x + 0.10 nmol/L r= 0.992 | ≤ 10% change up to 2 g/L of hemoglobin | ≤ 10% change up to .2 g/L of conjugated and unconjugated bilirubin, | ≤ 10% change up to 1000 mg/dL (intralipid) | no change up to > 500 ng/mL | NA | NA |  |  | traceable to an internal standard manufactured using highly purified material | 11203697_EN Rev.02, 2020-04 | IFU Document
Atellica® IM | 11200643 |  | id10 | 00630414246963 | OUS | Hepatitis | No | Atellica IM HBsAgII Quant | Atellica IM QHBsQHBs | QHBs | 200 | 1x200 | Atellica IM PW3 Atellica IM Multi-Diluent 14 | QHBs QC MCMs | 26 | 90 | monoclonal mouse anti-HBsAg | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#EDTA Plasma;# | 8 hrs at room temp,separated specimens. 7 days at 2-8C in primary tubes, 14 days at 2-8C in separated specimens, -20C for longer storage | 100 uL | No |  | sandwich, direct chemiluminescent | 0.017-125.0 IU/mL undiluted | 0.017-125.0 IU/mL undiluted | NA | NA | NA | NA | see IFU | N/A | N/A | Atellica IM QHBs CAL | provided in kit | 60 days at 2-8C | Liquid | N/A | 3 for lot calibration, 2 for pack recalibration | 2 | purified human HBsAg in buffer with sodium azide | 21 | 21 |  |  |  | Yes | Atellica IM Multi-Diluent 14 | 11200662 |  | 0.017-312,500 IU/mL | 0.017-312,500 IU/mL | Yes | Atellica IM QC | 11200644 | N/A | N/A | N/A | N/A | 0.014 IU/mL | 0.007 IU/mL | 0.017 IU/mL | 0.007 IU/mL | 0.014 IU/mL | 0.017 IU/mL | 0.659 | 0.659 | 1.4 | 2.6 | 305.6 | 305.6 | 2.7 | 3.6 | 1506 | 1506 | 3.6 | 5.7 | N/A | N/A | ≤ 10% interference up to 500 mg/dL of hemoglobin | ≤ 10% interference up to 40 mg/dL of bilirubin | ≤ 10% up to 1000 mg/dL of triglycerides | ≤ 10% up to 10 ng/mL of biotin | 3.5 mg/mL | 3.5 mg/mL |  |  | (WHO) 3rd International Standard for HBsAg, NIBSC code 12/226 | 11200655_EN Rev. 03, 2019-02 | IFU Document |  |  |  |  | 0.179 | 17.5 | 5.5 | 6.3
Atellica® CH | 11097543 |  | ge09 | 00630414279206 | Global | General Chemistry | No | Total Protein, CSF/Urine | UCFPUCFP | 4049 | 1480 | 4 x 370 | N/A | N/A | 5 | 30 | Pyrogallol red, methanol, sodium molybdate | ;#Urine;#Cerebrospinal fluid (CSF);# | Urine: 3 days when stored at 4°C with no additives; After 24 hours, store at 2–4°C for < 72 hours or frozen at -20°C for up to 1 year; CSF: Analyze fresh specimens or store at 4°C8 for less than 72 hours. Frozen specimens are stable at -20°C for 6 months. | 6.3 | No |  | pyrogallol red-molybdate | 6.0 – 250.0 mg/dL | 60 – 2500 mg/L | 6.0 – 2500 mg/dL | 60 – 25,000 mg/L | Urine: 1 -14 mg/dL CSF: 15 -45 mg/dL | Urine:10 – 140 mg/L CSF: 150 – 450 mg/L | N/A | N/A | N/A | Urinary/Cerebrospinal Fluid Protein Calibrator (UCFP CAL) | 11099339 | 60 | Liquid | No preparation required | 3 | 5 | Human serum albumin and IgG | 60 | 7 |  | No |  | Yes | N/A | 11099300 |  | 6.0 – 2500 mg/dL | 60 – 25,000 mg/L | No | N/A | N/A | 6.0 mg/dL | 60 mg/L | N/A | N/A | 6.0 mg/dL | 3.0 mg/dL | 6.0 mg/dL | 30 mg/L | 60 mg/L | 60 mg/L | 23.3 mg/dL | 233 mg/L | 2.4 | 3.4 | 70.8 mg/dL | 708 mg/L | 0.8 | 1.4 | 149.4 mg/dL | 1494 mg/L | 0.6 | 1.2 | Urine Dimension UCFP method y = 1.01x - 1.0 mg/dL r=0.999 n=109 CSF Dimension UCFP method y = 1.05x - 1.5 mg/dL r=0.999 n=118 | Urine Dimension UCFP method y = 1.01x - 10 mg/L r=0.999 n=109 CSF Dimension UCFP method y = 1.04x - 15 mg/L r=0.999 n=118 | N/A | N/A | N/A | No Biotin in assay formulation | N/A | N/A | N/A | N/A | Traceable to the NIST-SRM 927 | 11111705_EN Rev. 01, 2020-01 | IFU Document | 0.266 | 7.302 | 8.89 | 6.985 | 0.121 | 18.55 | 22.58 | 17.74
Atellica® IM | 11204647 |  | si02 |  | Global | Special ID | No | Atellica IM Zika Ab (ZikaAb) - Must be run with Atellica IM Zika IgM (ZikaM) | ZikaAb | ZikaAb | 100 |  | Atellica IM Zika IgM (ZikaM) 11204645 Atellica IM ZikaAb QC 11204648 Atellica IM ZikaM QC 11204646 |  | 39 | 28 | NS1 antigen | ;#Serum;#Lithium Heparin Plasma;#Potassium EDTA Plasma;# | After centrifugation, specimens in the primary collection device in contact with the cells or clot are stable for up to 24 hours at room temperature, and for up to 5 days at 2–8°C. • Separated samples are stable for up to 24 hours at room temperature, and for up to 5 days at 2–8°C. • Separated samples, devoid of red blood cells, are stable at ≤ -20°C for up to 150 days.Avoid more than 3 freeze-thaw cycles. Do not store in a frost-free freezer. | 15 μL | No |  | Antibody capture immunoassay using a 2-pass format | 0.00–9.00 Index | [5 empty] | < 0.80 Index (nonreactive), ≥ 0.80 Index (equivocal) using the Atellica IM ZikaAb, assay must be reflex tested using the Atellica IM ZikaM assay |  |  | Included |  | Opened at 2–8°C for 60 days | N/A | N/A | 3 for lot calibration, 2 for pack recalibration | 2 | Processed human plasma negative or positive for Zika antibodies; | 56 | 14 |  |  |  | No | NA |  |  |  |  | Yes | Atellica IM ZikaAb QC | 11204648 | [24 empty] | ≤ 10% change up to 1000 mg/dL of hemoglobin | ≤ 10% change up to 40 mg/dL of conjugated bilirubin | ≤ 10% change up to 3000 mg/dL of triglycerides | ≤ 10% change up to 3500 ng/mL of biotin |  |  |  |  | Based on agreement with known Zika positive samples. | 11204673_EN Rev. 01, 2019-11, 11205700_EN Rev. 01, 2019-11 | IFU Document
Atellica® IM | 11204645 |  | si02 |  | Global | Special ID | No | Atellica IM Zika IgM (ZikaM) - Must be run with Atellica IM Zika Ab (ZikaAb) | ZikaM | ZikaM | 50 | 50 tests | Atellica IM Zika Ab (ZikaAb) 11204647 Atellica IM ZikaAb QC 11204648 Atellica IM ZikaM QC 11204646 |  | 39 | 28 | Anti-human IgM monoclonal antibody | ;#Serum;#Lithium Heparin Plasma;#Potassium EDTA Plasma;# | • After centrifugation, specimens in the primary collection device in contact with the cells or clot are stable for up to 24 hours at room temperature, and for up to 5 days at 2–8°C. • Separated samples are stable for up to 24 hours at room temperature, and for up to 5 days at 2–8°C. • Separated samples, devoid of red blood cells, are stable at ≤ -20°C for up to 150 days. Avoid more than 3 freeze-thaw cycles. Do not store in a frost-free freezer. | 15 μL | No |  | IgM capture immunoassay using a 2-pass format | 0.00–9.00 Index | [5 empty] | Must be reflexed from ZikaAb. < 1.00 Index (nonreactive), ≥ 1.00 Index (reactive) using the Atellica IM ZikaM assay must be retested in duplicate (see IFU for more details). |  |  | Included |  | Opened at 2–8°C 60 days | N/A |  | 3 for lot calibration, 2 for pack recalibration | 2 | Processed human plasma negative or positive for Zika IgM antibodies | 49 | 14 |  |  |  | No | NA |  |  |  |  | Yes | Atellica IM ZikaM QC | 11204646 | [24 empty] | ≤ 10% change up to 1000 mg/dL of hemoglobin | ≤ 10% change up to 40 mg/dL of conjugated bilirubin | ≤ 10% change up to 3000 mg/dL of triglycerides | ≤ 10% change up to 3500 ng/mL of biotin |  |  |  |  | Based on agreement with known Zika positive samples. | 11205700_EN Rev. 01, 2019-11, 11204673_EN Rev. 01, 2019-11 | IFU Document
Atellica® IM | 11206923 |  | si03 | 00630414605586 | OUS | Special ID | No | Atellica IM SARS-CoV-2 Total | Atellica IM COV2T | COV2T | 500 | 5 x 100 tests | Atellica IM COV2T QC SMN 11206712 | NA | 10 | 28 | Streptavidin-coated paramagnetic microparticles preformed with biotinylated S1 RBD antigen, bovine serum albumin, goat serum | ;#Serum;#Lithium Heparin Plasma;#Potassium EDTA Plasma;# | stable onboard for 24 hrs. separated samples- 3 days at RT, 5 days at 2-8 C or freeze at -20 C for longer storage | 50 | No |  | 1-step antigen sandwich immunoassay using acridinium ester chemiluminescent technology | 0.60-75.0 Index | 0.60-75.00 Index |  |  |  |  | Nonreactive: <1.0 Index | Reactive: ≥ 1.0 Index | NA | Atellica IM COV2T CAL | included in reagent kit | 8 hrs at RT, 60 days opened at 2-8C | Liquid | Ready to use | 3 for lot calibration, 2 for pack recalibration | 2 | Processed human plasma nonreactive for antibodies to SARS-CoV-2 and processed human plasma spiked with antiboies to SARS-CoV-2 | 28 | 14 |  |  |  | Yes | ATellica IM Mult-Diluent 2 | 10995644 |  | 0.60 -375.00 Index | 0.60 -375.00 Index | Yes | Atellica IM COV2T | 11206712 | NA | NA | NA | NA | 0.5 | 0.4 | 0.6 | 0.4 | 0.5 | 0.6 | 2.72 Index | 2.72 Index | 2.1 | 7.6 | 10.84 Index | 19.84 Index | 2.6 | 9.2 | 65.28 Index | 65.78 Index | 2.7 | 10.1 | N/A | N/A | <10% change up to 1000 mg/dL of hemoglobin | <10% change up to 40 mg/dL of conjugated or unconjugated bilirubin | <10% change up to 2000 mg/dL of triglycerides (Intralipid) | <10% change up to 3500 ng/dL of biotin | NA | NA |  |  | Based on agreement with known positive and negative SARS-CoV-2 samples | 11206928_EN Rev. 02, 2020-10 | IFU Document |  |  |  |  | 0.5 | 17.7 | 5.5 | 17.5
Atellica® IM | 11206711 |  | si03 | 00630414605593 | OUS | Special ID | No | Atellica IM SARS-CoV-2 Total | Atellica IM COV2T | COV2T | 100 | 1 x 100 tests | Atellica IM COV2T QC SMN 11206712 | NA | 10 | 28 | Steptavidin-coated paramagnetic microparticles preformed with biotinylated S1 RBD antigen, bovine serum albumin, goat serum | ;#Serum;#Lithium Heparin Plasma;#Potassium EDTA Plasma;# | stable onboard for 24 hrs, separated samples - 3 days at RT, 5 days at 2-8C or freeze at -20C for longer storage | 50 | No |  | 1-step antigen sandwich immunoassay using acridinium ester chemiluminescent technology | 0.60-75.00 Index | 0.60-75.00Index |  |  |  |  | Nonreactive :<1.0 Index | Reactive: ≥1.0 Index | NA | Atellica IM COV2T CAL | included in reagent kit | 8 hrs at RT, 60 days opened at 2 - 8 C | Liquid | Ready to use | 3 for lot calibration, 2 for pack recalibration | 2 | Processed human plasma nin reactive for antibodies to SARS-CoV-2 and processed plasma spiked with antibodies to SARS-CoV-2 | 28 | 14 |  |  |  | Yes | Atellica IM Mult-Diluent 2 | 10995644 |  | 0.60-375 Index | 0.60-375 Index | Yes | Atellica IM COV2T QC | 11206712 | NA | NA | NA | NA | 0.5 | 0.4 | 0.6 | 0.4 | 0.5 | 0.6 | 2.72 Index | 2.72 Index | 2.1 | 7.6 | 10.84 Index | 10.84 Index | 2.6 | 9.2 | 65.28 | 65.28 | 2.7 | 10.1 | N/A | N/A | ≤ 10% change up to 1000 mg/dL of hemoglobin | ≤ 10% change up to 40 mg/dL of conjugated or unconjugated bilirubin | ≤ 10% change up to 2000 mg/dL of triglycerides (Intralipid) | ≤ 10% change up to 3500 ng/dL of biotin | NA | NA |  |  | Based on agreement with known postive and negative SARS-CoV-2 samples | 11206928_EN Rev.02, 2020-10 | IFU Document |  |  |  |  | 0.13 | 17.9 | 6.7 | 7.4
Atellica® CH | 11532568 |  | ge35 | 00630414287270 | Global | General Chemistry | No | Atellica CH Lactate_2 (Lac_2) (Plasma) | Lac_2Lac_2 | 8010 | 380 | 4 x 95 | N/A | N/A | 4 | 30 days per pack | PIPES buffer, TOOS, Peroxidase (plant), lactate oxidase and ascorbate oxidase (Microorganism);4-aminoantipyrine | ;#Lithium Heparin Plasma;#Potassium EDTA Plasma;#Sodium fluoride/potassium oxalate Plasma;# | Plasma: 1 day at 2–8°C; 30 days at -20°C. | 4 μL | No |  | lactate oxidase to pyruvate and hydrogen peroxide | Plasma: 0.6–110.0 mg/dL | Plasma: 0.07–12.21 mmol/L | Plasma: 0.6 - 1100.0 mg/dL | Plasma: 0.07 - 122.10 mmol/L | Plasma:4.5–19.8 mg/dL | Plasma: 0.50–2.20 mmol/L | N/A | N/A | N/A | Atellica CH SPCL CHEM CAL | 11099438 | 7 | Lyophilized | 5.0 mL of distilled water; Stand for 30 minutes at room temperature; Invert to mix Prior to use mix the contents by gently inverting the vials | 3 | 2 | Human serum | 62 | 30 |  | No |  | Yes | N/A | 11099300 |  | Plasma: 0.6 - 1100.0 mg/dL | Plasma: 0.07 - 122.10 mmol/L | No | N/A | N/A | 0.6 mg/dL | 0.07 mmol/L | N/A | N/A | 0.6 mg/dL | 0.4 mg/dL | N/A | 0.04 mmol/L | 0.07 mmol/L | N/A | 12.5 mg/dL | 1.39 mmol/L | 1.9 | 4.9 | 48.7 mg/dL | 5.41 mmol/L | 0.9 | 4.8 | 99.0 mg/dL | 10.99 mmol/L | 0.8 | 4.6 | ADVIA Chemistry 1800 Lac (plasma) y = 1.01x - 0.2 mg/dL r=0.995 n =101 | ADVIA Chemistry 1800 Lac (plasma) y = 1.01x - 0.02 mmol/L r=0.995 n =101 | ≤ 10% change up to 1000 mg/dL of hemoglobin | <10% change up to 9.2 mg/dL of conjugated bilirubin; <10% change up to 7.5 dL mg/dL of unconjugated bilirubin | < 10% change up to 1000 mg/dL of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | Internal standard using highly purified material | 11417138_EN 11417138_EN Rev. 02, 2020-09 | IFU Document | 0.442 | 5.625 | 3.5 | 2.75 | 0.201 | 14.287 | 8.89 | 6.985
Atellica® IM | 10733011 |  |  |  | OUS | Inflamation | No | Atellica IM Interleukin-6 (IL6) | Atellica IM IL6 | IL6 | 100 | 1 x 100 | NA | Atellica IM IL6 QC Atellica IM Multi-Diluent 13 | 14 | 42 | Monoclonal mouse anti-IL-6 antibody | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | 5 hours at room temperature, 24 hours at 2–8°C,frozen for 1 month at ≤ -20°C. Avoid more than 3 freeze/thaw cycles. | 50 μL | No |  | one-step direct immunoassay using chemiluminescent technology | 2.7 – 5500.0 pg/mL | NA |  |  |  |  | up to 4.4 pg/mL | NA | NA | Atellica IM IL6 CAL | Included in the kit | Reconstituted at 2–8°C 30 days, at room temperature 8 hours | Lyophilized | 2.0 mL of special reagent water. Invert vials 5 times, let stand for 30 minutes at room temperature. | 3 for lot calibration, 2 for pack recalibration | 2 | recombinant human IL-6; bovine serum albumin | 42 | 14 |  |  |  | Yes | Atellica IM Multi‐Diluent 13 | 10995643 |  | 2.7 – 55,000 pg/mL | NA | Yes | Atellica IM IL6 | 10733014 | NA | NA | NA | NA | 1.0 pg/mL | 0.6 pg/mL | 2.7 pg/mL | NA | NA | NA | 5.1 pg/mL | NA | 7.9 | 9.6 | 249.3 pg/mL | NA | 2.7 | 3.1 | 3981.6 pg/mL | NA | 2.1 | 2.5 | ADVIA Centaur IL6 (x) y = 0.97x + 0.08 pg/mL r= 1.00 | NA | ≤ 10% interference up to 1000 mg/dL (0.62 mmol/L) | ≤ 10% interference up to 25 mg/dL (427.5 μmol/L) conjugated, 60 mg/dL (1026 μmol/L) unconjugated | ≤ 10% interference up to 1500 mg/dL (17.0 mmol/L) | NA | 299,586 pg/mL | NA |  |  | WHO 1st IS (IL-6, Human rDNA derived), NIBSC code 89/548 | RPBL1348R02_EN_Rev. 02, 2020-07 | IFU Document
Atellica® IM | 10733015 |  |  |  | OUS | Inflamation | No | Atellica IM Lipopolysaccharide Binding Protein (LBP) | Atellica IM LBP | LBP | 50 | 1 x 50 | NA | Atellica IM LBP Quality Control Material Atellica IM Multi‐Diluent 1 | 14 | 28 | Monoclonal mouse anti-LBP antibody | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | up to 12 hours at room temperature, for up to 5 days at 2–8°C, 1 month at ≤ -20°C. Avoid more than 1 freeze/thaw cycle | 20 μL | No |  | one-step sandwich immunoassay using direct chemiluminescent technology | 1.1 – 120.0 μg/mL | NA |  |  |  |  | 2.3 – 18.3 μg/mL | NA | NA | Atellica IM LBP CAL | Included in the kit | Opened at 2–8°C 28 days, At room temperature 4 hours | Liquid | NA | 3 for lot calibration, 2 for pack recalibration | 2 | Recombinant LBP antigen; BSA; preservatives | 28 | 14 |  |  |  | No | Atellica IM Multi-Diluent 1 | 10995639 |  | 1.1 – 240.0 μg/mL |  | Yes | Atellica IM LBP QC | 10733018 | NA | NA | NA | NA | 1.1 μg/mL | 0.7 μg/mL | 1.1 μg/mL | NA | NA | NA | 14.4 μg/mL | NA | 4.4 | 7.1 | 60.4 μg/mL | NA | 6.2 | 8.1 | 103.2 μg/mL | NA | 5.3 | 7.2 | ADVIA Centaur® (X) Atellica IM (y) = 0.90 (x) + 0.70 μg/mL (intercept) r= 0.98 | NA | ≤ 10% interference up to 1000 mg/dL (0.62 mmol/L) | ≤ 10% interference up to 60 mg/dL (1026 μmol/L) Conjugated, and Unconjugated | ≤ 10% interference up to 1500 mg/dL (17.0 mmol/L) | NA | 1093 μg/mL | NA |  |  | traceable to an internal standard. | RPBL1320R03_EN Rev. 03, 2019-10 | IFU Document
Atellica® CH | 11097579 |  | ge19 | 00630414238364 | OUS | General Chemistry | No | Atellica CH Pancreatic Amylase_2 (Urine) | PAMY_2PAMY_2 | 4053 | 1720 | 4 x 430 | N/A | N/A | 9 | 31 | monoclonal mouse antibodies | ;#Urine;# | 10 days at RT; 31 days at 2-8C; Urine amylase is unstable in acidic urine. Adjust urine to a pH ≥ 7 before storage | 16 | No |  | Enzymatic reaction | 20–1500 U/L | 20–1500 U/L | 20–7500 U/L | 20–7500 U/L | Urine: Spontaneously voided urine: Male 7–356 U/L Female 13–319 U/L | Urine: Spontaneously voided urine: Male 7–356 U/L Female 13–319 U/L | N/A | N/A | N/A | Atellica CH SPCL CHEM CAL | 11097579 | 7 | Lyophilized | 5.0 mL of distilled water; Stand for 30 minutes at room temperature; Invert to mix Prior to use mix the contents by gently inverting the vials | 3 | 2 | Human serum | 62 | 31 |  | No |  | Yes | N/A | 11099300 |  | 20–7500 U/L | 20–7500 U/L | No | N/A | N/A | 5 U/L | 5 U/L | N/A | N/A | 5 U/L | 2 U/L | 17 U/L | 2 U/L | 5 U/L | 17 U/L | 48 U/L | 48 U/L | 3.1 | 3.3 | 105 U/L | 105 U/L | 0.9 | 1.1 | 1137 U/L | 1137 U/L | 0.2 | 0.5 | Roche cobas AMY-P (urine) y = 1.13x + 2.3 U/L r=0.994 n=106 | Roche cobas AMY-P (urine) y = 1.13x + 2.3 U/L r=0.994 n=106 | ≤ 10% change up to 500 mg/dL of hemoglobin. | ≤ 10% change up to 30 mg/dL of bilirubin | ≤ 10% change up to 1000 mg/dL of Intralipid |  | N/A | N/A | N/A | N/A | IRMM/IFCC-456 Pancreatic Alpha- Amylase Reference Material and traceable to JCCLS-SOP through the CRM-001c referen | 11111708_EN Rev. 03, 2020-01 | IFU Document | 0.886 | 5.625 | 3.5 | 2.75 | 0.402 | 14.287 | 8.89 | 6.985
Atellica® CH | 10470528 |  |  | 8058056680747 | OUS | General Chemistry | Yes | Glucose-6-Phosphate Dehydrogenase (Sentinel Alliance Application) | G6PDHG6PDH | 4082 | 448 | 1 x 448 | Atellica CH Empty Reagent Packs (11097534) | N/A | 8.5 | 30 | Non Biological G6P-DH; triethanolamine buffer; EDTA; NADP; G6PDH | ;#Whole blood;# | Unprepared: 8 hours at 2-8 °C Prepared sample: 2 hours at 2-8 °C | 4 | Yes |  | Enzymatic / Rate Reaction | 39 to 3700 U/L | 39 to 3700 U/L | N/A | N/A | In erythrocytes: 290 - 412 mU/109 erythrocytes (37 °C) 10.10 - 14.19 U/g Hb (37 °C) | In erythrocytes: 290 - 412 mU/109 erythrocytes (37 °C) 10.10 - 14.19 U/g Hb (37 °C) | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 62 | 30 |  | No |  | Yes | N/A | 11099300 |  | N/A | N/A | No | Sentinel G6P-DH Control Set | 11532787 | 11.9 U/L | 11.9 U/L | N/A | N/A | 11.9 U/L | 6.5 U/L | 35.1 U/L | 6.5 U/L | 11.9 U/L | 35.1 U/L | 218.9 | 218.9 | 7.2 | 8.3 | 990.3 | 990.3 | 1.9 | 2.5 | 3639.7 | 3639.7 | 0.9 | 1.6 | Abbott Alinity c8000: y = 18.23 + 1.097x U/L r-0.996 U/L | Abbott Alinity c8000: y = 18.23 + 1.097x U/L r-0.996 U/L | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | IFU: 17005-4.0/02 https://intranet.sentinel.it/sentineldoc</a> | IFU Document
Atellica® CH | 11532568 |  |  | 00630414287270 | Global | General ChemistryPlease select a value... | No | Atellica CH Lactate_2 (Lac_2) CSF | Lac_2Lac_2 | 8010 | 380 | 4 x 95 | N/A | N/A | 4 | 30 days per pack | PIPES buffer, TOOS, Peroxidase (plant), lactate oxidase and ascorbate oxidase (Microorganism);4-aminoantipyrine | ;#Cerebrospinal fluid (CSF);# | CSF: 24 hr at 2–8°C; 30 days at -20°C | 4 μL | No |  | lactate oxidase to pyruvate and hydrogen peroxide | CSF: 0.9–70 mg/dL | CSF: 0.10–7.77 mmol/L | CSF: 0.9 –140 mg/dL | CSF: 0.10 - 15.54 mmol/L | CSF: 5.4–19.8 mg/dL | CSF: 0.60–2.20 mmol/L | N/A | N/A | N/A | Atellica CH SPCL CHEM CAL | 11099438 | 7 | Lyophilized | 5.0 mL of distilled water; Stand for 30 minutes at room temperature; Invert to mix Prior to use mix the contents by gently inverting the vials | 3 | 2 | Human serum | 62 | 30 |  |  |  | Yes | N/A | 11099300 |  | CSF: 0.9 –140 mg/dL | CSF: 0.10 - 15.54 mmol/L | No | N/A | N/A | 0.6 mg/dL | 0.07 mmol/L | N/A | N/A | 0.6 mg/dL | 0.1 mg/dL | N/A | 0.04 mmol/L | 0.07 mmol/L | N/A | 12.5 mg/dL | 1.39 mmol/L | 1.9 | 4.9 | 48.7 mg/dL | 5.41 mmol/L | 0.9 | 4.8 | 99.0 mg/dL | 10.99 mmol/L | 0.8 | 4.6 | Dimension Vista LA method (CSF) y = 0.96x - 0.6 mg/dL r=0.998 n =124 | Dimension Vista LA method (CSF) y = 0.96x - 0.07 mmol/L r=0.998 n =124 | ≤ 10% change up to 1000 mg/dL of hemoglobin | <10% change up to 9.2 mg/dL of conjugated bilirubin; <10% change up to 7.5 dL mg/dL of unconjugated bilirubin | < 10% change up to 1000 mg/dL of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | Internal standard using highly purified material | 11417138_EN 11417138_EN Rev. 02, 2020-09 | IFU Document
Atellica® IM | 10995541 |  |  | 630414600413 | Global | DiabetesDiabetes | No | Atellica IM C-peptide (CpS) | Atellica IM CpSCpS | CpS | 100 | 1 x 100 | Atellica IM CpS CAL SMN 10995542 | Atellica IM Multi-Diluent 10 SMN 10995640 Atellica IM CpS MCM SMN 10995543 | 14 | 28 | Monoclonal mouse anti-C-peptide antibody | ;#Urine;# | Store at 2 - 8 C during the collection process. If not run within 2hours of collection completion aliquot and freeze at -20 C | 50 μL | No |  | 2-site sandwich chemiluminescent with acridinium-ester technology | 0.50 -300 ng/mL | 0.50 -300 ng/mL | All urines must be diluted 1:10 | All urines must be diluted 1:10 | <156.46 µg/24 hours | <156.46 µg/24 hours | NA | NA | NA | Atellica IM CpS CAL | 10995542 | 4 hours at RT | Liquid | Let the calibrators stand for 5 minutes at RT, Gently mix | 3 | 2 | citric acid buffer with casein and perservatives | 50 | 14 |  |  |  | No | Multi-Diluent 10 | 10995640 |  | All urines must be diluted 1:10 | All urines must be diluted 1:10 | No | Use comercially available control | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 18.95 ng/mL | 18.95 ng/mL | 2.4 | 5.1 | 97.28 ng/mL | 97.28 ng/mL | 2 | 9.3 | 248.40 ng/mL | 248.40 ng/mL | 3.7 | 8.5 | ADVIA Centaur CpS; y = 1.05x + 0.10 ng/mL r = 1.00 | ADVIA Centaur CpS; y = 1.05x + 0.10 ng/mL r = 1.00 | NA | NA | NA | NA | NA | NA | NA | NA | WHO International Standard 84/510 | 10997742_EN Rev .01, 2017-10 | IFU Document
Atellica® IM | 10995538 |  |  | 00630414598659 | OUS | N/AMetabolic | No | Atellica IM Cortisol (Cor) | Atellica IM CorCor | Cor | 50 | 1 x 50 | Atellica IM CAL E | Cortisol Urine Reconstitution Buffer (SMN 10995540) | 10 | 10 | Rabbit anti-cortisol antibody&nbsp; bound to mouse monoclonal anti-rabbit IgG antibody | ;#Urine;# | >8hrs at 2–8°C >48hrs freeze at ≤ -20°C for up to one month | 20 μL | No |  | Competitive immunoassay using direct chemiluminescent | 0.50- 53.00 μg/dL | 13.80– 1462.27 nmol/L | NA | NA | Direct Urine : 20.9 - 292.3 μg/24 hours Extracted Urine 9.5 - 136.2 μg/24 hours | Direct Urine: 57.7 - 806.8 nmol/24 hours Extracted urine 26.2 - 375.9 nmol/24 hours | NA | NA | NA | Atellica IM CAL E | 10995512 (2‑pack) 10995513 (6‑pack) | 14 days recon at 2–8°C 4 hours recons at RT | Lyophilized | Add 2.0 mL of special reagent water into each vial. Let vials stand for 15-20 minutes at room temperature. | 3 for lot calibration, 2 for pack recalibration | 2 | defibrinated human plasma | 19 | 10 |  |  |  | Yes | Atellica IM Multi‑Diluent 3 | 10995645 |  | NA | NA | No | Use commercially available controls | NA | NA | NA | NA | NA | Direct Urine ≤ 0.45 μg/dL, Extracted urine ≤ 0.44 μg/dL | Direct Urine ≤ 0.19 μg/dL, Extracted urine ≤ 0.18 μg/dL | Direct Urine ≤ 0.48 μg/dL, Extracted urine ≤ 0.44 μg/dL | Direct urine ≤ 5.24 nmol/L Extracted urine ≤4.97 nmol/L | Direct urine ≤ 12.42 nmol/L Extracted urine ≤12.14 nmol/L | Direct urine ≤ 13.24 nmol/L Extracted urine ≤12.14 nmol/L | 7.05 μg/dL | 194.51 nmol/L | 2.6 | 8.9 | 23.09 μg/dL | 637.05 nmol/L | 5.5 | 9.6 | 35.99 μg/dL | 992.96 nmol/L | 4.2 | 8.9 | ADVIA Centaur Cortisol y = 1.10x + 0.59 μg/dL r = 0.98 | ADVIA Centaur Cortisol y = 1.10x + 16.28 nmol/L r = 0.98 | NA | NA | NA | NA | NA | NA | NA | NA | Internal standards manufactured analytically which are traceable to gas chromatography-mass spectrometry (GCMS) | 10995336_EN Rev. 03, 2020-03 | IFU Document
Atellica® IM | 10995537 |  |  | 00630414600413 | OUS | N/AMetabolic | No | Atellica IM Cortisol (Cor) | Atellica IM CorCor | Cor | 250 | 5 x 50 | Atellica IM CAL E | Cortisol Urine Reconstitution Buffer (SMN 10995540) | 10 | 10 | Rabbit anti-cortisol antibody bound to mouse monoclonal anti-rabbit IgG antibody | ;#Urine;# | >8hrs at 2–8°C >48hrs freeze at ≤ -20°C for up to one month | 20 μL | No |  | Competitive immunoassay using direct chemiluminescent | 0.50- 53.00 μg/dL | 13.80– 1462.27 nmol/L | NA | NA | Direct Urine : 20.9 - 292.3 μg/24 hours Extracted Urine 9.5 - 136.2 μg/24 hours | Direct Urine: 57.7 - 806.8 nmol/24 hours Extracted urine 26.2 - 375.9 nmol/24 hours | NA | NA | NA | Atellica IM CAL E | 10995512 (2‑pack) 10995513 (6‑pack) | 14 days recon at 2–8°C, 4 hours recons at RT | Lyophilized | Add 2.0 mL of special reagent water into each vial. Let vials stand for 15-20 minutes at room temperature. | 3 for lot calibration, 2 for pack recalibration | 2 | defibrinated human plasma | 19 | 10 |  |  |  | Yes | Atellica IM Multi‑Diluent 3 | 10995645 |  | NA | NA | No | Use commercially available controls | NA | NA | NA | NA | NA | Direct Urine ≤ 0.45 μg/dL, Extracted urine ≤ 0.44 μg/dL | Direct Urine ≤ 0.19 μg/dL, Extracted urine ≤ 0.18 μg/dL | Direct Urine ≤ 0.48 μg/dL, Extracted urine ≤ 0.44 μg/dL | Direct urine ≤5.24 nmol/L Extracted urine 4.97 nmol/L | Direct urine ≤ 12.42 nmol/L Extracted urine ≤12.14 nmol/L | Direct urine ≤ 13.24 nmol/L Extracted urine ≤12.14 nmol/L | 7.05 μg/dL | 194.51 nmol/L | 2.6 | 8.9 | 23.09 μg/dL | 637.05 nmol/L | 5.5 | 9.6 | 35.99 μg/dL | 992.96 nmol/L | 4.2 | 8.9 | ADVIA Centaur Cortisol y = 1.10x + 0.59 μg/dL r = 0.98 | ADVIA Centaur Cortisol y = 1.10x + 16.28 nmol/L r = 0.98 | NA | NA | NA | NA | NA | NA | NA | NA | Internal standards manufactured analytically which are traceable to gas chromatography-mass spectrometry (GCMS) | 10995336_EN Rev. 03, 2020-03 | IFU Document
Atellica® IM | 11200450 |  |  |  | OUS | N/AThyroid | No | Atellica IM Thyroglobulin (Tg) | Atellica IM TgTg | Tg | 100 | 1 x 100 | NA | Atellica IM Multi-Diluent 11 SMN 109995642 (2-pk) SMN 10995641 (1 x 10 mL) | 18 | 28 | murine monoclonal anti-thyroglobulinantibody | ;#Serum;#EDTA Plasma;#Lithium Heparin Plasma;# | up to 3 days at RT, up to 7 days at 2-8C and freeze at -20 C for up to 12 months | 100 µL | No |  | 1-step sandwich chemiluminescent immunoassay | 0.050 - 150 ng/mL | 0.076 - 227 pmol/L | 1:10 0.50 - 1500 ng/mL, 1:20 0.50 - 3000 ng/mL, 1:50 0.50 - 7500 ng/mL | 1:10 0.076 - 2270pmol/L, 1 :20 0.076 - 4540 pmol/L, 1:50 0.076 - 11,350 pmol/L | 1.82 - 111 ng/mL | 2.76 - 168 pmol/L | NA | NA | NA | Atellica IM Tg CAL | included in reagent kit | aliquot and freeze -20C for up to 2 months | Lyophilized | Reconstitute with 4.0 mL distilled of deionized water. Mix by gentle inversion or swirling. | 2 | 2 | human thyroglobulin | 21 | 21 |  |  |  | Yes | Atellica IM Multi-Diluent 11 | 10995642/10995641 |  | 1:10 0.50 - 1500 ng/mL, 1:20 0.50 - 3000 ng/mL, 1:50 0.50 - 7500 ng/mL | 1:10 0.076 - 2270pmol/L, 1 :20 0.076 - 4540 pmol/L, 1:50 0.076 - 11,350 pmol/L | No | Use commercially available controls | NA | NA | NA | NA | NA | 0.036 ng/mL | 0.026 ng/mL | 0.050 ng/mL | 0.039 pmol/L | 0.036 pmol/L | 0.076 pmol/L | 0.157 ng/mL | 0.238 pmol/L | 5.1 | 8.3 | 78.9 ng/mL | 120 pmol/L | 1.2 | 2.7 | 136 ng/mL | 206 pmol/L | 1.9 | 2.9 | Commercial Tg method : y = 1.26 x - 0.215 ng/mL r = 0.979 | Commercial Tg method: y + 1.26x - 0.326 pmol/L r = 0.979 | ≤ 10% change up to1000 mg/dL of hemoglobin | ≤ 10% change up to 1900 mg/dL of intralipid | ≤ 10% change up to 60 mg/dL of bilirubin | ≤ 10% change up to 3500 ng/mL of biotin | none up to 80,000 ng/mL | none up to 121, 200 pmol/L | none up to 80,000 ng/mL | none up to 121, 200 pmol/L | CRM 457 | 11200496_EN Rev. 01 2020-11 | IFU Document
Atellica® IM | 11201759 |  |  |  | OUS | N/AThyroid | No | Atellica IM Anti‑Thyroglobulin II (aTgII) | Atellica IM aTg IIaTgII | aTgII | 100 | 1 x 100 | NA | Atellica IM aTgII QC SMN 1120752 Atellica IM aTgII MCM SMN 11201751 | 14 | 28 | acridinium-ester-labeled human thyroglobulin | ;#Serum;#Potassium EDTA Plasma;#Lithium Heparin Plasma;#Sodium Heparin Plasma;# | 8hr at room temperature, 7 days at 2-8C, freeze at -20c up to 6 months | 25 | No |  | 1-step analyte-bridging immunoassay using acridinium ester chemiluminescence | 1.3 - 1000 IU/mL | 1.3 - 1000 IU/mL | Not recommended | Not recommended | Cutoff 4.5 IU/mL | Cutoff 4.5 IU/mL | NA | NA | NA | Atellica IM aTgII CAL | included in reagent kit | 8 hrs | Lyophilized | Add 1.0 mL of reagent water . Let stand 20-30 minutes at RT. Gently mix and invert vials to ensure homogeneity | 2 | 2 | monoclonal mouse anti-human thyroglobulin | 90 | 28 |  |  |  | NA | NA | NA |  | Not recommended | Not recommended | Yes | Atellica IM aTgII QC | 11201752 | NA | NA | NA | NA | 0.9 IU/ML | 0.7 IU/mL | 0.9 IU/mL | 0.9 IU/mL | 0.7 IU/mL | 0.9 IU/mL | 7.56 IU/mL | 7.56 IU/mL | 1.9 | 2.3 | 50.9 IU/mL | 50.9 IU/mL | 1.5 | 2.3 | 915 IU/mL | 915 IU/mL | 1.9 | 2.4 | ADVIA Centaur; y =; 0.94x + 0.01 IU/mL; r = 0.997 | ADVIA Centaur; y = 0.94x + 0.01 IU/mL r = 0.997 | ≤ 10% change up to 1000 mg/dL of hemoglobin | ≤ 10% change up to 60 mg/dL of bilirubin | ≤ 10% change up to 3500 mg/dL of intralipid | ≤ 10% change up to 3500 mg/dL of biotin | none up to 50,000 IU/mL | none up to 50,000 IU/mL | none up to 50,000 IU/mL | none up to 50,000 IU/mL | WHO IRP, anti-Thyroglobulin Serum, Human (NIBSC 65/093 | 11204720_EN Rev. 01, 2020-02 | IFU Document
Atellica® IM | 11201761 |  |  |  | OUS | N/AThyroid | No | Atellica IM Anti‑Thyroglobulin II (aTgII) | Atellica IM aTg IIaTgII | aTgII | 500 | 5 x 100 | NA | Atellica IM aTgII QC SMN 1120752 Atellica IM aTgII MCM SMN 11201751 | 14 | 28 | acridinium-ester-labeled human thyroglobulin | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#Potassium EDTA Plasma;# | 8hr at room temperature, 7 days at 2-8C, freeze at -20c up to 6 months | 25 | No |  | 1-step analyte-bridging immunoassay using acridinium ester chemiluminescence | 1.3 - 1000 IU/mL | 1.3 - 1000 IU/mL | Not recommended | Not recommended | Cutoff 4.5 IU/mL | Cutoff 4.5 IU/mL | NA | NA | NA | Atellica IM aTgII CAL | included in reagent kit | 8 hrs | Lyophilized | Add 1.0 mL of reagent water . Let stand 20-30 minutes at RT. Gently mix and invert vials to ensure homogeneity | 2 | 2 | monoclonal mouse anti-human thyroglobulin | 90 | 28 |  |  |  | NA | NA | NA |  | Not recommended | Not recommended | Yes | Atellica IM aTgII QC | 11201752 | NA | NA | NA | NA | 0.9 IU/ML | 0.7 IU/mL | 0.9 IU/mL | 0.9 IU/mL | 0.7 IU/mL | 0.9 IU/mL | 7.56 IU/mL | 7.56 IU/mL | 1.9 | 2.3 | 50.9 IU/mL | 50.9 IU/mL | 1.5 | 2.3 | 915 IU/mL | 915 IU/mL | 1.9 | 2.4 | ADVIA Centaur; y =; 0.94x + 0.01 IU/mL; r = 0.997 | ADVIA Centaur; y = 0.94x + 0.01 IU/mL r = 0.997 | ≤ 10% change up to 1000 mg/dL of hemoglobin | ≤ 10% change up to 60 mg/dL of bilirubin | ≤ 10% change up to 3500 mg/dL of intralipid | ≤ 10% change up to 3500 mg/dL of biotin | none up to 50,000 IU/mL | none up to 50,000 IU/mL | none up to 50,000 IU/mL | none up to 50,000 IU/mL | WHO IRP, anti-Thyroglobulin Serum, Human (NIBSC 65/093 | 11204720_EN Rev. 01, 2020-02 | IFU Document
Atellica® IM | 10995585 |  |  |  | Global | N/AThyroid | No | Atellica IM Free Triiodothyronine (FT3) | Atellica IM FT3FT3 | FT3 | 60 | 1 x 60 | Atellica IM CAL A 2- pack SMN 10995500 Atellica IM CAL A 6- pack SMN 10995501 | Atellica IM FT3 MCM SMN 10995586 | 14 | 28 | acridinium-ester-labeled mouse monoclonal anti-T3 antibodies | ;#Serum;#EDTA Plasma;#Lithium Heparin Plasma;# | 8hr at RT; 48 hr at 2-8C; Freeze samples at or below -20°C if the sample is not assayed within 48 hr | 35 | No |  | competitive immunoassay using direct chemiluminescent technology | 0.20 - 20.00 pg/mL | 0.31 - 30.80 pmol/L | NA | NA | 2.3 - 4.2 pg/mL | 3.5 - 6.5 pmol/L | NA | NA | NA | Atellica IM CAL A | 10995500 (2-pack) 10995501 (6-pack) | 28 days recon at 2–8°C, 4 hours recon at RT | Lyophilized | Add 5.0 mL of reagent water, let the calibrators stand for 15–20 minutes at RT. Mix by gentle inversion until homogenous | 3 for lot calibration. 2 for pack calibration | 2 | human plasma with sodium azide,preservatives and protein stabilizers | 82 | 28 |  |  |  | No | NA | NA |  | NA | NA | No | Use commecially available controls | NA | 0.06 pg/mL | 0.09 pmol/L | NA | NA | 0.02 pg/mL | 0.00 pg/mL | NA | 0.03 pmol/L | 0.00 pmol/L | NA | 2.87 pg/mL | 4.42 pmol/L | 0.94 | 1.78 | 4.96 pg/mL | 7.64 pmol/L | 0.67 | 1.07 | 18.13 pg/mL | 27.92 pmol/L | 1.96 | 3.05 | ADVIA Centaur FT3 y = 0.99 x + 0.007 pg/mL r = 0.99 | ADVIA Centaur FT3 y = 0.99x + 0.011 pmol/L r = 0.99 | <10% change up to 500 mg/dL of hemoglobin | <10% change up to 20 mg/dL of bilirubin | <10% change up to 1000 mg/dL Intralipid | No Biotin in assay formulation | NA | NA | N/A | N/A | Internal standard manufactured using U.S.P.(United States Pharmacopeia) material | 10995347_EN Rev.03, 2020-06 | IFU Document
Atellica® IM | 10995584 |  |  |  | Global | N/AThyroid | No | Atellica IM Free Triiodothyronine (FT3) | Atellica IM FT3FT3 | FT3 | 300 | 5 x 60 | Atellica IM CAL A 2- pack SMN 10995500 Atellica IM CAL A 6- pack SMN 10995501 | Atellica IM FT3 MCM SMN 10995586 | 14 | 28 | acridinium-ester-labeled mouse monoclonal anti-T3 antibodies | ;#Serum;#EDTA Plasma;#Lithium Heparin Plasma;# | 8hr at RT; 48 hr at 2-8C; Freeze samples at or below -20°C if the sample is not assayed within 48 hr | 35 | No |  | competitive immunoassay using direct chemiluminescent technology | 0.20 - 20.00 pg/mL | 0.31 - 30.80 pmol/L | NA | NA | 2.3 - 4.2 pg/mL | 3.5 - 6.5 pmol/L | NA | NA | NA | Atellica IM CAL A | 10995500 (2-pack) 10995501 (6-pack) | 28 days recon at 2–8°C, 4 hours recon at RT | Lyophilized | Add 5.0 mL of reagent water, let the calibrators stand for 15–20 minutes at RT. Mix by gentle inversion until homogenous | 3 for lot calibration. 2 for pack calibration | 2 | human plasma with sodium azide,preservatives and protein stabilizers | 82 | 28 |  |  |  | No | NA | NA |  | 0.02 pg/mL | 0.03 pmol/L | No | Use commecially available controls | NA | 0.06 pg/mL | 0.09 pmol/L | NA | NA | 0.02 pg/mL | 0.00 pg/mL | NA | 0.03 pmol/L | 0.00 pmol/L | NA | 2.87 pg/mL | 4.42 pmol/L | 0.94 | 1.78 | 4.96 pg/mL | 7.64 pmol/L | 0.67 | 1.07 | 18.13 pg/mL | 27.92 pmol/L | 1.96 | 3.05 | ADVIA Centaur FT3 y = 0.99 x + 0.007 pg/mL r = 0.99 | ADVIA Centaur FT3 y = 0.99x + 0.011 pmol/L r = 0.99 | <10% change up to 500 mg/dL of hemoglobin | <10% change up to 20 mg/dL of bilirubin | <10% change up to 1000 mg/dL Intralipid | No Biotin in assay formulation | NA | NA | N/A | N/A | Internal standard manufactured using U.S.P.(United States Pharmacopeia) material | 10995347_EN Rev.03, 2020-06 | IFU Document
Atellica® IM | 10995703 |  |  | 00630414600161 | OUS | N/AThyroid | No | Atellica IM Thyroid Stimulating Hormone 3-Ultra (TSH3‑UL) | Atellica IM TSH3‑ULTSH3UL | TSH3UL | 110 | 1 x 110 | NA | Atellica IM TSH3‑UL MCM SMN10995705 Multi-Diluent 15 SMN 11200798 | 14 | 63 | anti-FITC monoclonal antibody | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | RT 24 hours, 2–8°C 48 hours, serum and EDTA ≤ -20°C; up to 30 days, Lithium heparin ≤ -20°C; up to 14 days; Freeze samples only 1 time | 75 μL | No |  | 3rd generation immunoassay using direct chemiluminescent technology | 0.008–150.000 μIU/mL | 0.008–150.000 mIU/L | NA | NA | 0.55 - 4.78 μIU/mL | 0.55 - 4.78 mIU/L | NA | NA | NA | Atellica IM TSH3-UL CAL | Included in reagent kit | 28 days recon at 2–8°C,&nbsp; 4 hours recon at RT | Lyophilized | Add 2.0 mL of reagent water, let the calibrators stand for 15–20 minutes at RT. Gently swirl and invert the vials until homogeneous | 3 for lot calibration, 2 for pack recalibration | 2 | Equine serum | 49 | 63 |  |  |  | Yes | Multi-Diluent 15 | 11200798 |  | 1:2 0.008 – 300.000 μIU/mL, 1:5 0.008 - 750.000 μIU/mL | 1:2 0.008 – 300.000 mIU/L 1;5 0.008 - 750.000 mIU/L | No | Use commercially available Controls | NA | NP | NP | NP | NP | 0.007 μIU/mL | 0.003 μIU/mL | 0.007 μIU/mL | 0.003 mIU/L | 0.007 mIU/L | 0.007 mIU/L | 0.157 μIU/mL | 0.157 mIU/L | 2.1 | 3.7 | 8.995 μIU/mL | 8.995 mIU/L | 1.5 | 3.4 | 118.735 μIU/mL | 118.735 mIU/L | 2.1 | 3.3 | ADVIA Centaur TSH3‑UL y = 1.05x + 0.001 μIU/mL r = 1.000 | ADVIA Centaur TSH3‑UL y = 1.05x + 0.001 mIU/L r = 1.000 | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 40 mg/dL of conjugated bilirubin; ≤ 10% change up to 40 mg/dL of unconjugated bilirubin | ≤ 10% change up to 1000 mg/dL of Intralipid | No Biotin in assay formulation | 3000 μIU/mL | 3000 mIU/L | 3000 mIU/L | 3000 mIU/L | WHO 3rd International Standard for human TSH (IRP 81/565) | 10995434_EN Rev. 03, 2019-08(OUS) | IFU Document
Atellica® IM | 10995704 |  |  | 00630414600178 | OUS | N/AThyroid | No | Atellica IM Thyroid Stimulating Hormone 3-Ultra (TSH3‑UL) | Atellica IM TSH3‑ULTSH3UL | TSH3UL | 550 | &nbsp;5 x 110 | NA | Atellica IM TSH3‑UL MCM SMN10995705 Multi-Diluent 15 SMN 11200798 | 14 | 63 | anti-FITC monoclonal antibody | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | RT 24 hours, 2–8°C 48 hours, serum and EDTA ≤ -20°C; up to 30 days, Lithium heparin ≤ -20°C; up to 14 days; Freeze samples only 1 time | 75 μL | No |  | 3rd generation immunoassay using direct chemiluminescent technology | 0.008–150.000 μIU/mL | 0.008–150.000 mIU/L | NA | NA | 0.55 - 4.78 μIU/mL | 0.55 - 4.78 mIU/L | NA | NA | NA | Atellica IM TSH3-UL CAL | NA-included in reagent kit | 28 days recon at 2–8°C,&nbsp;&nbsp; 4 hours recon at RT | Lyophilized | Add 2.0 mL of reagent water, let the calibrators stand for 15–20 minutes at RT. Gently swirl and invert the vials until homogeneous | 3 for lot calibration, 2 for pack recalibration | 2 | Equine Serum | 49 | 63 |  |  |  | Yes | Multi-Diluent 15 | 11200798 |  | 1:2 0.008 – 300.000 μIU/mL, 1:5 0.008 - 750.000 μIU/mL | 1:2 0.008 – 300.000 mIU/L 1;5 0.008 - 750.000 mIU/L | No | Use commercially available Controls | NA | NP | NP | NP | NP | 0.007 μIU/mL | 0.003 μIU/mL | 0.007 μIU/mL | 0.003 mIU/L | 0.007 mIU/L | 0.007 mIU/L | 0.157 μIU/mL | 0.157 mIU/L | 2.1 | 3.7 | 8.995 μIU/mL | 8.995 mIU/L | 1.5 | 3.4 | 118.735 μIU/mL | 118.735 mIU/L | 2.1 | 3.3 | ADVIA Centaur TSH3‑UL y = 1.05x + 0.001 μIU/mL r = 1.000 | ADVIA Centaur TSH3‑UL y = 1.05x + 0.001 mIU/L r = 1.000 | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 40 mg/dL of conjugated bilirubin, ≤ 10% change up to 40 mg/dL of unconjugated bilirubin | ≤ 10% change up to 1000 mg/dL of Intralipid | No Biotin in assay formulation | 3000 μIU/mL | 3000 mIU/L | 3000 mIU/L | 3000 mIU/L | WHO 3rd International Standard for human TSH (IRP 81/565) | 10995434_EN Rev. 03, 2019-08(OUS) | IFU Document

## Sheet5 (hidden)
Atellica® CH;Atellica® IM; – OUS;Global;
Discipline/Disease state | TEST | N | Materials Required But Not Provided | Time | OBS | Sample Types | Sample Stability | Sample Volume | Sample Pretreatment | Assay Principle | Analytical Measuring Range (Common Units) | Extended Range (Common Units) | Expected Values (Common Units) | Expected Values #1 | Expected Values #2 | Expected Values #3 | Calibrator Name | Calibrator Open Vial Stability | Lot Calibration Interval (days) | Pack Calibration Interval (days) | Automatic Dilution | Sample Diluent | Dilution Factor2 | Reportable Range With Dilution (Common Units) | Quality Control Name | Hemolysis | Icterus | Lipemia | Biotin | Hook Effect (Common Units) | Standardization Reference Method/Material
Special ID | Syph | 200 | APW1 | 29 | 60 | Serum/Plasma | RT not recommended 2–8°C for up to 7 days >7 days store at ≤ -20°C | 100 μL | No | Fully automated, antigen sandwich assay, using direct chemiluminescent technology | 0.10-45.00 index |  |  | Nonreactive, Reactive, Equivocal |  |  | Atellica IM Syph CAL | At 2–8°C Until expiration date 8 hrs at RT | 37 | 21 | No | NA |  | NA | Atellica IM Syph QC | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 40 mg/dL of conjugated bilirubin up to 40 mg/dL of unconjugated bilirubin | 500 ng/mL 1000 mg/dL of triglycerides | Insignificant effect on the assay up to 1500 ng/mL | NA | Standardization is based upon relative clinical agreement with commercially available syphilis assays
Reproductive Endocrinology | eE2 | 100 |  | 14 | 42 | Serum/Plasma | RT 20 hours, 2–8°C after 20 hours, ≤ -20°C after 48 hours up to 6 months, Freeze only 1 time | 80 μL | No | Competitive assay | 11.80–3000.00 pg/mL |  |  | See IFU |  |  | Atellica IM CAL 30 | Recon at 2–8°C 14 days, Recon at RT 4hrs | 22 | 21 | Yes | Atellica IM eE2 DIL |  | 11.8-15.000 pg/mL | NA | ≤ 10% change up to 250 mg/dL of hemoglobin | ≤ 10% change upto 20 mg/dL of conjugated/unconjugated bilirubin | ≤ 10% change up to 1000 mg/dL of triglycerides | ≤ 10% change up to 1500 ng/mL of biotin | NA | standardized using internal standards manufactured analytically which are traceable to ID-GC/MS
Reproductive Endocrinology | PRL | 50 |  | 14 | 28 | Serum/Plasma | up to 8hrs at RT, 8 hrs at 2–8°C, 48 hrs at ≤ -20°C | 25 μL | No | 2-site sandwich immunoassay | 0.30–200.00 ng/mL |  |  | See IFU |  |  | Atellica IM CAL B | at 2–8°C 28 days | 34 | 28 | Yes | Atellica IM Multi‑Diluent 1 |  | 0.30–1000.00 ng/mL | NA | insignificant effect up to 500 mg/dL of hemoglobin | insignificant effect up to 20 mg/dL of bilirubin | insignificant effect up to 1000 mg/dL of triglycerides | No Biotin in assay formulation | 30,000 ng/mL | traceable to the World Health Organization (WHO) 3rd IRP for human Prolactin (84/500)
Reproductive Endocrinology | PRGE | 90 |  | 14 | 28 | Serum/Plasma | up to 8hrs at RT, 8 hrs at 2–8°C, 48 hrs at ≤ -20°C, Freeze samples only 1 time and mix thoroughly after thawing. | 10 μL | No | competitive immunoassay | 0.21–60.00 ng/mL |  |  | See IFU for Normals, Pregnant and Postmenopausal Females | Males 0.28–1.22 ng/mL 0.89–3.88 nmol/L | See IFU for Pediatric Reference Intervals by Age | Atellica IM CAL E | at 2–8°C 28 days | 90 | 28 | Yes | Atellica IM Multi‑Diluent 3 |  | 0.21–600.00 ng/mL | NA | ≤ 10% change up to 250 mg/dL (0.15 mmol/L) hemoglobin | ≤ 10% change up to 20 mg/dL (341 μmol/L) conjugated bilirubin 20 mg/dL (341 μmol/L) unconjugated bilirubin | ≤ 10% change up to 500 mg/dL (5.65 mmol/L) intralipids | No Biotin in assay formulation | N/A | standardized using internal standards manufactured analytically which are traceable to gas chromatography–mass spectroscopy (GCMS)
Reproductive Endocrinology | TSTII | 100 | APW1 | 14 | 18 | Serum/Plasma | up to 48 hours at RT, 7 days at 2–8ºC, 7 days at ≤ -20°C | 20 μL | No | competitive immunoassay using direct chemiluminescent technology | 7.00–1500.00 ng/dL |  |  | See IFUs |  |  | Atellica IM TSTII CAL | 14 days recon at 2–8°C 6 hrs recon at RT | 91 | 28 | Yes | Atellica IM Multi‑Diluent 3 |  | 7.00–3000 ng/dL | NA | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 15 mg/dL of conjugated bilirubin 20 mg/dL of unconjugated bilirubin | ≤ 10% change up to 1000 mg/dL of triglycerides | ≤ 10% change in results up to 30 ng/mL of biotin | N/A | traceable to the primary testosterone standard National Measurement Institute (NMI) M914
Reproductive Endocrinology | FSH | 190 |  | 14 | 28 | Serum/Plasma | RT up to 8 hours > 8 hours at 2–8°C > 48 hr at ≤ -20°C Freeze 1 time | 50 μL | No | 2-site samdwich immonoassay | 0.30–200.00 mIU/mL |  |  | See IFU |  |  | Atellica IM CAL B | 28 days recon at 2–8°C 4 hours at RT | 36 | 14 | Yes | Atellica IM Multi‑Diluent 1 |  | 0.30–400.00 mIU/mL | NA | ≤ 10% change up to up to 150 mg/dL of hemoglobin | ≤ 10% change up to 20 mg/dL of conjugated bilirubin 20 mg/dL of unconjugated bilirubin | ≤ 10% change up to 500 mg/dL of lipemia | No Biotin in assay formulation | 1000 mIU/mL | Traceable to World Health Organization (WHO) 2nd International Standard for human FSH (IS 94/632)
Reproductive Endocrinology | LH | 110 |  | 14 | 28 | Serum/Plasma | RT up to 8 hours > 8 hours at 2–8°C >48 hr at ≤ -20°C Freeze 1 time | 25 μL | No | 2-site sandwich immunoassay | 0.07–200.00 mIU/mL |  |  | See IFU |  |  | Atellica IM CAL B | 28 days recon at 2–8°C 4 hours at RT | 65 | 28 | Yes | Atellica IM Multi‑Diluent 1 |  | 0.07–400.00 mIU/mL | NA | ≤ 10% change up to up to 500 mg/dL of hemoglobin | ≤ 10% change up to 20 mg/dL of conjugated bilirubin 20 mg/dL of unconjugated bilirubin | ≤ 10% change up to 3000mg/dL of lipemia | No Biotin in assay formulation | 18,000 mIU/mL | Traceable to World Health Organization (WHO) 2nd International Standard for human LH (IS 80/552).
N/AThyroid | FT4 | 50 |  | 14 | 21 | Serum/Plasma | RT 8 hrs, 2–8°C after 8 hours; ≤ -20°C after 48 hours; Freeze only 1 time | 25 μL | No | Competitive immunoasay using direct chemiluminescent technology | 0.1 – 12.0 ng/dL | NA | 0.89- 1.76 ng/dL | NA | NA | NA | Atellica IM CAL A | 28 days recon at 2–8°C, 4 hours recon at RT | 21 | 7 | No | NA |  | NA | Use commercially available Controls | NSI up to 300 mg/dL of hemoglobin | NSI up to 20 mg/dL of bilirubin | NSI up to 1000 mg/dL of triglycerides | NSI up to 3500 ng/mL of biotin | NA | Internal standard manufactured using U.S.P.(United States Pharmacopeia) material
N/AThyroid | T3 | 120 |  | 14 | 28 | Serum/Plasma | up to 8hrs at RT, 8 hrs at 2–8°C, 48hrs at ≤ -20°C | 30 | No | Competitive immunoassay using direct chemiluminescent technology | 0.10–8.00 ng/mL | NA | 0.6 - 1.81 ng/mL | NA | NA | NA | Atellica IM CAL A | 28 days recon at 2–8°C; 4 hrs recon at RT | 56 | 21 | Manual | Atellica IM T3 DIL |  | N/A | Use commercially available controls | ≤ 10% change up to 250 mg/dL (0.15 mmol/L) hemoglobin | ≤ 10% change up to 20 mg/dL (341 μmol/L) conjugated bilirubin 20 mg/dL (341 μmol/L) unconjugated bilirubin | ≤ 10% change up to1000 mg/dL (11.3 mmol/L) intralipids | No Biotin in assay formulation | N/A | Internal standard manufactured using U.S.P.(United States Pharmacopeia) material
N/AThyroid | aTPO | 100 |  | 14 | 28 | Serum/Plasma | 8hr at room temperature, separate cells from serum, 48 hours at 2-8C, >48 hrs freeze at -20C once | 30 | No | Competitive immunoassay using direct chemiluminescent technology | 28–1300 U/mL | 1:5 28 - 6500 1:10 28 - 13,000 | Cutoff 60 U/mL | NA | NA | NA | Atellica IM CAL O | 8 hours | 28 | 14 | Yes | Atellica IM aTPO DIL |  | 1:5 28 - 6500 1:10 28 - 13,000 | Atellica IM aTPO QC | ≤ 5% change up to 500 mg/dL of hemoglobin | ≤ 5% change up to 40 mg/dL of bilirubin | ≤ 5% change up to 1000 mg/dL of lipid | No Biotin in assay formulation | NA | WHO Reference Preparation 66/387
N/AThyroid | aTgII | 100 |  | 14 | 28 | Serum/Plasma | 8hr at room temperature, 7 days at 2-8C, freeze at -20c up to 6 months | 25 | No | 1-step analyte-bridging immunoassay using acridinium ester chemiluminescence | 1.3 - 1000 IU/mL | Not recommended | Cutoff 4.5 IU/mL | NA | NA | NA | Atellica IM aTgII CAL | 8 hrs | 90 | 28 | NA | NA |  | Not recommended | Atellica IM aTgII QC | ≤ 10% change up to 1000 mg/dL of hemoglobin | ≤ 10% change up to 60 mg/dL of bilirubin | ≤ 10% change up to 3500 mg/dL of intralipid | ≤ 10% change up to 3500 mg/dL of biotin | none up to 50,000 IU/mL | WHO IRP, anti-Thyroglobulin Serum, Human (NIBSC 65/093
N/AThyroid | TSH3UL | 110 |  | 14 | 63 | Serum/Plasma | RT 24 hours, 2–8°C 48 hours, serum and EDTA ≤ -20°C; up to 30 days, Lithium heparin ≤ -20°C; up to 14 days; Freeze samples only 1 time | 75 μL | No | 3rd generation immunoassay using direct chemiluminescent technology | 0.008–150.000 μIU/mL | NA | 0.55 - 4.78 μIU/mL | NA | NA | NA | Atellica IM TSH3-UL CAL | 28 days recon at 2–8°C,&nbsp; 4 hours recon at RT | 49 | 63 | Yes | Multi-Diluent 15 |  | 1:2 0.008 – 300.000 μIU/mL, 1:5 0.008 - 750.000 μIU/mL | Use commercially available Controls | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 40 mg/dL of conjugated bilirubin; ≤ 10% change up to 40 mg/dL of unconjugated bilirubin | ≤ 10% change up to 1000 mg/dL of Intralipid | No Biotin in assay formulation | 3000 μIU/mL | WHO 3rd International Standard for human TSH (IRP 81/565)
N/AMetabolic | HCY | 100 |  | 14 | 28 | Serum/Plasma | 48 hrs at 2-8°C, frozen at <-20 °C up to 13 weeks | 20 μL | No | Competitive immunoassay using direct chemiluminescent technology | 0.50–65.00 μmol/L | 0.50– 260.00 μmol/L | 3.7 – 13.9 μmol/L | NA | NA | NA | Atellica IM HCY CAL | 4 hours | 42 | 14 | Yes | Atellica IM HCY DIL |  | 0.50–130.00 μmol/L | NA | ≤ 10% change up to 1000 mg/dL of hemoglobin | ≤ 10% change up to 25.0 mg/dL of bilirubin | ≤ 10% change up to 1300 mg/dL of lipid | ≤ 10% change up to 3500 ng/mL of biotin | NA | Internal Standard
N/AMetabolic | Cor | 50 |  | 10 | 10 | Serum/Plasma | Serum: 8hrs at RT; 48hrs at 2-8C; frozen at -20C | 20 μL | No | Competitive immunoassay using direct chemiluminescent | 0.50-75.00 μg/dL | 0.50–150.00 μg/dL | 7 - 9; (a.m.): 5.27–22.45 μg/dL 3 - 5 (p.m.) 3.44 - 16.76 μg/dL | NA | NA | NA | Atellica IM CAL E | 14 days recon at 2–8°C 4 hours recons at RT | 19 | 10 | Yes | Atellica IM Multi‑Diluent 3 |  | 0.50–150.00 μg/dL | Use commercially available controls | insignificant effect up to 500 mg/dL of hemoglobin | insignificant effect up to 20 mg/dL of conjugated bilirubin; insignificant effect up to 20 mg/dL of unconjugated bilirubin | insignificant effect up to 1170 mg/dL of triglycerides | No Biotin in assay formulation | NA | Internal standards manufactured analytically which are traceable to gas chromatography-mass spectrometry (GCMS)
N/AAnemia | VB12 | 100 |  | 14 | 18 | Serum/Plasma | 8 hours at RT 2–8°C after 8 hours; ≤ -20°C after 48 hours | 100 μL | No | competitive immunoassay using direct chemiluminescent technology | 45–2000 pg/mL | 45 – 20,000 pg/mL | 211 to 911 pg/mL | NA | NA | NA | Atellica IM CAL C | 28 days recon at 2–8°C 4 hours recon at RT | 30 | 18 | Yes | Atellica IM VB12 DIL |  | 45–20000 pg/mL | NA | NSI up to 150 mg/dL of hemoglobin | NSI up to 20 mg/dL of bilirubin | NSI up to 3000 mg/dL of triglycerides | No Biotin in assay formulation | NA | Traceable to an internal standard manufactured using U.S.P. (United States Pharmacopeia) material
N/AAnemia | Fer | 90 |  | 14 | 28 | Serum/Plasma | up to 8 hrs at RT; 48 hrs at 2–8°C. Freeze samples at ≤ -20°C if the sample is not assayed within 48 hours. | 10 μL | No | 2-site sandwich immunoassay using direct chemiluminometric | 0.5–1650.0 ng/mL | 0.5 – 16,500 ng/mL | Normal Males 22–322 ng/mL; Normal Females 10–291 ng/mL |  |  |  | Atellica IM Calibrator C (CAL C) | 4 hours at RT; 28 days at 2–8°C | 50 | 28 | Yes | Atellica IM Multi‑Diluent 1 |  | 0.5 – 16,500 ng/mL | NA | ≤ 10% change up to 900 mg/dL (0.56 mmol/L) of hemoglobin | ≤ 10% change up to 60 mg/dL conjugated bilirubin (1021 μmol/L) 60 mg/dL unconjugated bilirubin (1021 μmol/L) | ≤ 10% change up to 2000 mg/dL (22.6 mmol/L) | No Biotin in assay formulation | 80,000 ng/mL | Traceable to World Health Organization 2nd International Standard (WHO 80/578)
Liver Fibrosis | TIMP-1 | 50 |  | 14 | 60 | Serum | 48 hr at RT, 7 days at 2-8C. Freeze samples at ≤ -20°C if the sample is not assayed within 7 days | 25 μL | No | 2-site sandwich immunoassay using direct chemiluminescent technology | 3.5–1300.0 ng/mL | 3.5–6500.0 ng/mL | Refer to the Atellica IM ELF Test instructions for use for ELF score expected values |  |  |  | Atellica IM ELF CAL | 8 hours at RT; 5 days at 2–8°C; 90 days at -20°C | 64 | 60 | Yes | Atellica IM Multi‑Diluent 10 |  | 3.5–6500.0 ng/mL | Atellica IM ELF QC | < 10% change up to 500 mg/dL hemoglobin | < 10% change up to 40 mg/dL conjugated bilirubin 40 mg/dL unconjugated bilirubin | < 10% change up to 1000 mg/dL triglycerides | No Biotin in assay formulation | 30,000 ng/mL | Traceable to an internal standard manufactured using highly purified material
Liver Fibrosis | PIIINP | 50 | PW3 | 14 | 30 | Serum | 48 hr at RT, 7 days at 2-8C. Freeze samples at ≤ -20°C if the sample is not assayed within 7 days | 20 | No | 2-site sandwich immunoassay using direct chemiluminescent technology | 0.50–150.00 ng/mL | 0.50–750.00 ng/mL | Refer to the Atellica IM ELF Test instructions for use for ELF score expected values |  |  |  | Atellica IM ELF CAL | 8 hours at RT; 5 days at 2–8°C; 90 days at -20°C | 63 | 30 | Yes | Atellica IM Multi‑Diluent 13 |  | 0.50–750.00 ng/mL | Atellica IM ELF QC | < 10% change up to 500 mg/dL hemoglobin | < 10% change up to 40 mg/dL conjugated bilirubin 40 mg/dL unconjugated bilirubin | < 10% change up to 1000 mg/dL triglycerides | <10% change up to 50 ng/mL of biotin | 2500 ng/mL | Traceable to an internal standard manufactured using bovine PIIINP
Liver Fibrosis | HA | 50 |  | 39 | 28 | Serum | 48 hr at RT, 7 days at 2-8C. Freeze samples at ≤ -20°C if the sample is not assayed within 7 days | 20 μL | No | 2-site sandwich assay using direct chemiluminescent technology | 1.60–1000.00 ng/mL | 1.60 – 5000.00 ng/mL | Refer to the Atellica IM ELF Test instructions for use for ELF score expected values |  |  |  | Atellica IM ELF CAL | 8 hours at RT; 5 days at 2–8°C; 90 days at -20°C | 24 | 14 | Yes | Atellica IM Multi‑Diluent 13 |  | 1.60 – 5000.00 ng/mL | Atellica IM ELF QC | < 10% change up to 500 mg/dL hemoglobin | < 10% change up to 40 mg/dL conjugated bilirubin 40 mg/dL unconjugated bilirubin | < 10% change up to 1000 mg/dL triglycerides | No Biotin in assay formulation | 25,000 ng/mL | Traceable to an internal standard manufactured using highly purified material
HIV | CHIV | 100 | PW3 | 39 | 42 | Serum/Plasma | RT not recommended; at 2–8°C up to 14 days; at ≤ -20°C up to 8 months | 100 μL | No | antigen-bridging microparticle chemiluminometric immunoassay | 0.05–12.00 Index | N/A | see IFU | see IFU | see IFU | see IFU | IM CHIV CAL | at 2–8°C until expiration 8 hours at RT | 24 | 21 | No | NA |  | N/A | Atellica IM CHIV QC | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 40 mg/dL of conjugated bilirubin, and up to 40 mg/dL of unconjugated bilirubin | ≤10% change upto 1000 mg/dL of triglycerides | Insignificant effect on the assay&nbsp; up to 3500 ng/mL | N/A | Standardization is based upon relative clinical agreement with commercially available anti-HIV combo assays
Hepatitis | aHBe | 100 |  | 52 | 26 | Serum/Plasma | RT for up to 24 hrs; 2–8°C for up to 7 days; ≤ -20°C for longer storage | 100 μL | No | competitive/neutralizing immunoassay using a 2‑step, singlewash format. | 0.05–4.50 Index | NA | NA | Nonreactive < 0.80 Index | Reactive > 1.20 Index | Equivocal ≥0.80 Index and <1.20 Index | Atellica IM aHBe Cal | at 2–8°C until expirationon board at RT 8hrs | 51 | 21 | No | No |  | No | Atellica IM aHBe QC | ≤ 20% change up to 500 mg/dL of hemoglobin | ≤ 20% change up to 40 mg/dL of conjugated bilirubin, up to 20 mg/dL of unconjugated bilirubin | ≤ 20% change up to 500 mg/dL of triglycerides, and up to 400 mg/dL of cholesterol | Insignificant effect on the assay up to 1500 ng/mL | NA | Standarization based upon relative clinical agreement withcommercially available anti‑HBe assays.
Hepatitis | HBcT | 200 | APW1, PW3 | 46 | 28 | Serum/Plasma | RT not recommended; 2–8°C up to 7 days; ≤ -20°C for longer storage | 50 μL | No | 2‑wash antigen sandwich immunoassay | NA | NA | NA | Nonreactive < 0.50 Index | Reactive ≥ 0.50 Index |  | Atellica IM HBcT CAL | at 2–8°C until expiration; RT at 8hrs on board | 33 | 14 | No | N/A |  | N/A | Atellica IM HBcT QC | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 60 mg/dL of conjugated bilirubin; 40 mg/dL of unconjugated bilirubin | ≤ 10% change up to 1000 mg/dL of intralipids | Insignificant effect on the assay up to 1500 ng/mL | N/A | based upon the relative clinical agreement with commercially available anti-HBc Total assays
Hepatitis | aHBcM | 100 |  | 46 | 41 | Serum/Plasma | RT not recommended 2–8°C up to 7 days 7 days at ≤ -20°C | 15 μL | No | IgM capture immunoassay using a 2-step format | 0.05–9.00 Index | NA | NA | Negative <0.80 Index | Equivocal 0.80-0.99 Index | Positive ≥ 1.0 Index | Atellica IM aHBcM CAL | at 2–8°C until expiration RT at 8hrs on board | 92 | 27 | No | N/A |  | N/A | Atellica IM aHBcM QC | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 60 mg/dL of conjugated bilirubin 40 mg/dL of unconjugated bilirubin | ≤ 10% change up to 1000 mg/dL of triglycerides | Insignificant effect on the assay up to 150 ng/mL | N/A | assay standardization is based upon relative clinical agreement with commercially available anti-HBc IgM assays
Hepatitis | HBsII | 200 | PW3 | 26 | 60 | Serum/Plasma | RT not recommended; 2–8°C up to 7 days ; ≤ -20°C for longer storage | 100 μL | No |  | 0.1–1000.0 Index | NA | NA | Non Reactive <1.0 Index | Reactive ≥ 1.00 Index |  | Atellica IM HBsII CAL | at 2–8°C until expiration RT at 8hrs on board | 48 | 21 | No | N/A |  | N/A | Atellica IM HBsII QC | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 40 mg/dL of conjugated bilirubin, and up to 40 mg/dL of unconjugated bilirubin | ≤ 10% change up to 1000 mg/dL of triglycerides | ≤ 10% change up to 10 ng/mL | 3.36 mg/mL | based upon the relative clinical agreement with commercially available HBsAg assays
Hepatitis | HBeAg | 50 | PW3 | 39 | 60 | Serum/Plasma | RT not recommended; at 2–8°C up to 3 days (primary tubes), up to 7 days (separated samples); ≤ -20°C for longer storage | 100 μL | No |  | 0.05–1000.00 Index | NA | NA | See IFU |  |  | Atellica IM HBeAg CAL | at 2–8°C until expiration 8 hours at RT | 74 | 30 | Yes | Atellica IM Multi‑Diluent 10 |  | na | Atellica IM HBeAg QC | NA | NA | NA | Insignificant effect on the assay up to 1500 ng/mL | N/A | standardization is based upon relative clinical agreement with commercially available HBeAg assays
Hepatitis | aHCV | 200 | APW1 | 46 | 41 | Serum/Plasma | 2–8°C for up to 7 days, ≤ -20°C for longer storage | 10 μL | No | indirect two wash sandwich immunoassay | 0.02–11.00 Index | NA | NA | nonreactive <0.80 Index | reactive ≥ 1.00 Index | equivocal ≥ 0.80 Index and <1.00 Index | Atellica IM aHCV CAL | at 2–8°C until expiration 8 hours at RT | 60 | 28 | No | NA |  | NA | Atellica IM aHCV QC | ≤ 10% change up to up to 500 mg/dL of hemoglobin | ≤ 10% change up to 60 mg/dL of conjugated bilirubin 40 mg/dL of unconjugated bilirubin | ≤ 10% change up to 1000 mg/dL of triglycerides | Insignificant effect on the assay up to 1500 ng/mL | NA | Based upon relative clinical agreement with commercially available anti-HCV assays
Hepatitis | aHBs2 | 200 |  | 14 | 90 | Serum/Plasma | 2–8°C for up to 7 days ≤ -20°C for longer storage | 100 μL | No | direct sandwich | 3.1–1000.0 mIU/mL | NA | NA | nonreactive <8.0 mIU/mL | reactive ≥ 12.0 mIU/mL | retest zone ≥8.0 mIU/mL and <12.0 mIU/mL | Atellica IM aHBs2 CAL | At 2–8°C Until expiration date, 8 hrs at RT | 55 | 42 | Yes | Atellica IM Multi‑Diluent 11 |  | 3.1 - 10,000 mIU/mL | Atellica IM aHBs2 QC | ≤ 15% change up to 500 mg/dL of hemoglobin | ≤ 15% change up to 40 mg/dL of conjugated and unconjugated bilirubin | ≤ 15% change up to 1000 mg/dL triglycerides | ≤15% change up to 1500 ng/mL of biotin | 200,000 mIU/mL | Standardized against the World Health Organization (WHO) 1st International Reference Preparation (1977)
DiabetesDiabetes | CpS | 100 |  | 14 | 28 | Serum | Serum : Up to 8 hrs at RT, Up to 24 hrs at 2-8°C, Frozen at <-20 °C | 50 μL | No | 2-site sandwich immunoassay using direct chemiluminescent technology | 0.05 - 30.00 ng/mL | 1:2 0.05 - 60.00ng/mL 1:5 0.05 - 150 .0 ng/mL | 0.81 - 3.85 ng/mL | NA | NA | NA | Atellica IM CpS CAL | 4 hours at RT | 50 | 14 | Yes | Multi-Diluent 10 |  | 1:2 0.05 - 60.00ng/mL 1:5 0.05 - 150 .0 ng/mL | Use comercially available control | ≤ 10% change up to 250 mg/dL of hemoglobin | ≤ 10% change up to 20 mg/dL of bilirubin | ≤ 10% change up to 1000 mg/dL of lipid | No Biotin in assay formulation | 125 ng/mL | WHO International Standard 84/510
DiabetesDiabetes | IRI | 100 |  | 14 | 21 | Serum | Up to 8 hrs at RT, Up to 24 hrs at 2-8°C, Freeze at -20 °C after 24 hrs | 25 μL | No | 2-site sandwich chemiluminescent with acridinium-ester technology | 0.5 - 300.0 mU/L | 1:2 05 - 600 mU/L 1:5 0.5 - 1500 mU/L | 3.0 - 25 mU/L | NA | NA | NA | Atellica IM IRI CAL | 8 hrs at RT | 91 | 14 | Yes | Atellica IM IRI DIL |  | 1:2 0.5 - 600 mU/L 1:5 0.5 - 1500 mU/L | Use Commercially Available Controls | ≤ 6% change up to 125 mg/dL of hemoglobin | ≤ 6% change up to 20 mg/dL of bilirubin | ≤ 6% change up to 1000 mg/dL of lipid | No Biotin in assay formulation | 3000 mU/L | WHO 1st IRP 66/3024
DiabetesDiabetes | CpS | 100 |  | 14 | 28 | Urine | Store at 2 - 8 C during the collection process. If not run within 2hours of collection completion aliquot and freeze at -20 C | 50 μL | No | 2-site sandwich chemiluminescent with acridinium-ester technology | 0.50 -300 ng/mL | All urines must be diluted 1:10 | <156.46 µg/24 hours | NA | NA | NA | Atellica IM CpS CAL | 4 hours at RT | 50 | 14 | No | Multi-Diluent 10 |  | All urines must be diluted 1:10 | Use comercially available control | NA | NA | NA | NA | NA | WHO International Standard 84/510
Bone Metabolism | VitD | 100 |  | 14 | 28 | Serum/Plasma | < 24 hours at RT >24 hours at 2–8°C for up to 7 days &gt;7 days at ≤ -20°C | 20 μL | No | Competitive immunoassay using direct chemiluminescent technology | 4.20–150.00 ng/mL |  |  | See IFU |  |  | Atellica IM VitD CAL | 28 days recon at 2–8°C 120 days recon at ≤ -20°C 10 hours recon at RT | 28 | 28 | Yes | Atellica IM VitD DIL |  | 4.20–300.00 ng/mL | Atellica IM VitD QC | ≤ 10% change up to 155 mg/dL of hemoglobin | ≤ 10% change up to 40 mg/dL of conjugated bilirubin 40 mg/dL of unconjugated bilirubin | ≤ 10% change up to 540 mg/dL of triglycerides | No Biotin in assay formulation | NA | Traceable to the ID LC/MS/MS 25(OH)vitamin D RMP
Autoimmune | aCCP | 100 |  | 33 | 35 | Serum/Plasma | 22 hrs at RT, 7 days at 2–8°C, > 7days store at -20°C or colder, avoid > 2 freeze/thaw cycles | 15 μL | No | 2-step sandwich immunoassay | 0.54–200.00 U/mL |  |  | Healthly individuals median 0.74 U/mL | Samples ≤ 5.00 U/mL indicate that anti-CCP IgG antibodies were not detected | Samples ≥5.00 U/mL indicate that anti-CCP IgG antibodies were detected | Atellica IM aCCP CAL | 2–8°C 90 days after opening, At room temperature 8 hours | 35 | 14 | Yes | Atellica IM Multi-Diluent 1 |  | 0.54–2000.00 U/mL | Atellica IM aCCP QC | ≤ 10% change up to 1000 mg/dL of hemoglobin | ≤ 10% change up to 40 mg/dL of conjugated bilirubin and unconjugated bilirubin | ≤ 10% change up to 1500 mg/dL of Intralipid | ≤ 10% change up to 1500 ng/dL | 2145 U/mL | Traceable to an internal standard
Anemia | Fer | 450 |  | 14 | 28 | Serum/Plasma | up to 8 hrs at RT; 48 hrs at 2–8°C. Freeze samples at ≤ -20°C if the sample is not assayed within 48 hours. | 10 μL | No | 2-site sandwich immunoassay using direct chemiluminometric | 0.5–1650.0 ng/mL | 0.5 – 16,500 ng/mL | Normal Males 22–322 ng/mL; Normal Females 10–291 ng/mL |  |  |  | Atellica IM Calibrator C (CAL C) | 4 hours at RT; 28 days at 2–8°C | 50 | 28 | Yes | Atellica IM Multi‑Diluent 1 |  | 0.5 – 16,500 ng/mL | NA | ≤ 10% change up to 900 mg/dL (0.56 mmol/L) of hemoglobin | ≤ 10% change up to 60 mg/dL conjugated bilirubin (1021 μmol/L) 60 mg/dL unconjugated bilirubin (1021 μmol/L) | ≤ 10% change up to 2000 mg/dL (22.6 mmol/L) | No Biotin in assay formulation | 80,000 ng/mL | Traceable to World Health Organization 2nd International Standard (WHO 80/578)
Anemia | Fol | 140 | Fol DTT/REL, APW1 | 14 | 14 | Serum/WB | Serum < 8 hrs at RT > 8 hrs at 2–8°C >48 hrs at ≤ -20°C | 100 μL | No | Competitive immunoassay using direct chemiluminescent technology | 0.35–24.00 ng/mL | 0.35 - 48.00 ng/mL | Serum Folate; Deficient 0.35–3.37 ng/mL Indeterminate 3.38–5.38 ng/mL Normal > 5.38 ng/mL |  |  |  | Atellica IM Fol CAL | 7 days at 2–8°C<br>8 hours at RT<br>28 days ≤ -20°C | 14 | 7 | Yes | Atellica IM Fol DIL |  | 0.35 - 48.00 ng/mL | NA | NA | ≤ 10% change up to 20 mg/dL conjugated bilirubin, 20 mg/dL unconjugated bilirubin | ≤ 10% change up to 2000 mg/dL Lipemia | ≤ 10% change up to 38 ng/mL of biotin | NA | Traceable to an internal standard manufactured using highly purified material (N‑5‑methyl tetrahydrofolate)

## IM (hidden)
TEST | AMR | Reference Range | Time | OBS | Sample Volume | Calibrator Name | Materials Required But Not Provided | Expected Values #1 | Expected Values #2 | Expected Values #3 | N
aCCP | 0.54–200.00 U/mL | ≤ 5.00 U/mL | 29 | 60 | 100 μL | Atellica IM Syph CAL | APW1 | Healthly individuals median 0.74 U/mL | Samples ≤ 5.00 U/mL indicate that anti-CCP IgG antibodies were not detected | Samples ≥5.00 U/mL indicate that anti-CCP IgG antibodies were detected | 200
aHBcM | 0.05–9.00 Index | Negative : <0.80 Equivocal : 0.8~0.99 Positive : >1.0 | 14 | 42 | 80 μL | Atellica IM CAL 30 |  | Negative <0.80 Index | Equivocal 0.80-0.99 Index | Positive ≥ 1.0 Index | 100
aHBe | 0.05–4.50 Index | Negative : <0.80 Equivocal : 0.8~1.2 Positive : >1.2 | 14 | 28 | 25 μL | Atellica IM CAL B |  | Nonreactive < 0.80 Index | Reactive > 1.20 Index | Equivocal ≥0.80 Index and <1.20 Index | 50
aHBs2 | 3.1–1000.0 mIU/mL | Negative : <8 Equivocal : 8~12 Positive : >12 | 14 | 28 | 10 μL | Atellica IM CAL E |  | nonreactive <8.0 mIU/mL | reactive ≥ 12.0 mIU/mL | retest zone ≥8.0 mIU/mL and <12.0 mIU/mL | 90
aHCV | 0.02–11.00 Index | Negative : <0.80 Equivocal : 0.8~0.99 Positive : >1.0 | 14 | 18 | 20 μL | Atellica IM TSTII CAL | APW1 | nonreactive <0.80 Index | reactive ≥ 1.00 Index | equivocal ≥ 0.80 Index and <1.00 Index | 100
aTgII | 1.3 - 1000 IU/mL | Cutoff 4.5 IU/mL | 14 | 28 | 50 μL | Atellica IM CAL B |  | NA | NA | NA | 190
aTPO | 28–1300 U/mL | Cutoff 60 U/mL | 14 | 28 | 25 μL | Atellica IM CAL B |  | NA | NA | NA | 110
CHIV | 0.05–12.00 Index | Negative : <1.0 Positive : >1.0 | 14 | 21 | 25 μL | Atellica IM CAL A |  | see IFU | see IFU | see IFU | 50
Cor | 0.50-75.00 μg/dL | 7 - 9; (a.m.): 5.27–22.45 μg/dL 3 - 5 (p.m.) 3.44 - 16.76 μg/dL | 14 | 28 | 30 | Atellica IM CAL A |  | NA | NA | NA | 120
CpS | 0.05 - 30.00 ng/mL | 0.81 - 3.85 ng/mL | 14 | 28 | 30 | Atellica IM CAL O |  | NA | NA | NA | 100
eE2 | 11.80–3000.00 pg/mL |  | 14 | 63 | 75 μL | Atellica IM TSH3-UL CAL |  | See IFU |  |  | 110
Fer | 0.5–1650.0 ng/mL | Normal Males 22–322 ng/mL; Normal Females 10–291 ng/mL | 14 | 28 | 20 μL | Atellica IM HCY CAL |  |  |  |  | 100
Fol | 0.35–24.00 ng/mL | Serum Folate; Deficient 0.35–3.37 ng/mL Indeterminate 3.38–5.38 ng/mL Normal > 5.38 ng/mL | 14 | 18 | 100 μL | Atellica IM CAL C |  |  |  |  | 100
FSH | 0.30–200.00 mIU/mL |  | 14 | 28 | 10 μL | Atellica IM Calibrator C (CAL C) |  | See IFU |  |  | 90
FT4 | 0.1 – 12.0 ng/dL | 0.89- 1.76 ng/dL | 14 | 60 | 25 μL | Atellica IM ELF CAL |  | NA | NA | NA | 50
HA | 1.60–1000.00 ng/mL | Refer to the Atellica IM ELF Test instructions for use for ELF score expected values | 14 | 30 | 20 | Atellica IM ELF CAL | PW3 |  |  |  | 50
HBcT | NA | cut off: 0.5 | 39 | 28 | 20 μL | Atellica IM ELF CAL |  | Nonreactive < 0.50 Index | Reactive ≥ 0.50 Index |  | 50
HBeAg | 0.05–1000.00 Index | NA | 39 | 42 | 100 μL | IM CHIV CAL | PW3 | See IFU |  |  | 100
HBsII | 0.1–1000.0 Index | Cut-off : 1.00 | 52 | 26 | 100 μL | Atellica IM aHBe Cal |  | Non Reactive <1.0 Index | Reactive ≥ 1.00 Index |  | 100
HCY | 0.50–65.00 μmol/L | 3.7 – 13.9 μmol/L | 46 | 28 | 50 μL | Atellica IM HBcT CAL | APW1, PW3 | NA | NA | NA | 200
IRI | 0.5 - 300.0 mU/L | 3.0 - 25 mU/L | 46 | 41 | 15 μL | Atellica IM aHBcM CAL |  | NA | NA | NA | 100
LH | 0.07–200.00 mIU/mL |  | 26 | 60 | 100 μL | Atellica IM HBsII CAL | PW3 | See IFU |  |  | 200
PIIINP | 0.50–150.00 ng/mL | Refer to the Atellica IM ELF Test instructions for use for ELF score expected values | 39 | 60 | 100 μL | Atellica IM HBeAg CAL | PW3 |  |  |  | 50
PRL | 0.30–200.00 ng/mL |  | 14 | 90 | 100 μL | Atellica IM aHBs2 CAL |  | See IFU |  |  | 200
Syph | 0.10-45.00 index |  | 14 | 28 | 50 μL | Atellica IM CpS CAL |  | Nonreactive, Reactive, Equivocal |  |  | 100
T3 | 0.10–8.00 ng/mL | 0.6 - 1.81 ng/mL | 14 | 21 | 25 μL | Atellica IM IRI CAL |  | NA | NA | NA | 100
TIMP-1 | 3.5–1300.0 ng/mL | Refer to the Atellica IM ELF Test instructions for use for ELF score expected values | 14 | 28 | 50 μL | Atellica IM CpS CAL |  |  |  |  | 100
TSH3UL | 0.008–150.000 μIU/mL | 0.55 - 4.78 μIU/mL | 14 | 28 | 20 μL | Atellica IM VitD CAL |  | NA | NA | NA | 100
TSTII | 7.00–1500.00 ng/dL |  | 33 | 35 | 15 μL | Atellica IM aCCP CAL |  | See IFUs |  |  | 100
VB12 | 45–2000 pg/mL | 211 to 911 pg/mL | 14 | 28 | 10 μL | Atellica IM Calibrator C (CAL C) |  | NA | NA | NA | 450
VitD | 4.20–150.00 ng/mL |  | 14 | 14 | 100 μL | Atellica IM Fol CAL | Fol DTT/REL, APW1 | See IFU |  |  | 140

## Sheet2 (hidden)
Atellica® CH;Atellica® IM; – OUS;Global;
Platform | SMN | Catalog Number | Index | GTIN | Availability | Discipline/Disease state | Alliance Application | Product Name | Abbreviation Product/Test Name | TDef Code | Number of Tests | Kit Configuration | Materials Required But Not Provided | Optional Materials | TTFR (min) Time To First Result | Reagent On Board Stability (days) For CH assay - per well | Reagent Material Origin | Sample Types | Sample Stability | Sample Volume | Sample Pretreatment | Dilution Factor | Assay Principle | Analytical Measuring Range (Common Units) | Analytical Measuring Range (SI Units) | Extended Range (Common Units) | Extended Range (SI Units) | Expected Values (Common Units) | Expected Values (SI Units) | Expected Values #1 | Expected Values #2 | Expected Values #3 | Calibrator Name | Calibrator SMN | Calibrator Open Vial Stability | Calibrator (Lyophilized/Liquid) | Calibrator Preparation | Calibration: Number of replicates | Levels used in Calibration curve | Calibrator Material Origin | Lot Calibration Interval (days) | Pack Calibration Interval (days) | Calibration Interval (days) | IMT | Diluent | Automatic Dilution | Sample Diluent | Diluent SMN | Dilution Factor2 | Reportable Range With Dilution (Common Units) | Reportable Range With Dilution (SI Units) | Siemens QC avaiable? | Quality Control Name | Quality Control SMN | Analytical Sensitivity (Common Units) | Analytical Sensitivity (SI Units) | Functional Sensitivity (Common Units) | Functional Sensitivity (SI Units) | LoD (Common Units) | LoB (Common Units) | LoQ (Common Units) | LoB (SI Units) | LoD (SI Units) | LoQ (SI Units) | Mean Low Value Repeatability (Common Units) | Mean Low Value Repeatability (SI Units) | Low Repeatability (%) | Low Within-Lab Precision (%) | Mean Mid Value Repeatability (Common Units) | Mean Mid Value Repeatability (SI Units) | Mid Repeatability (%) | Mid Within-Lab Precision (%) | Mean High Value Repeatability (Common Units) | Mean High Value Repeatability (SI Units) | High Repeatability (%) | High Within-Lab Precision (%) | Method Comparison (Common Units) | Method Comparison (SI Units) | Hemolysis | Icterus | Lipemia | Biotin | Hook Effect (Common Units) | Hook Effect (SI Units) | High-Dose Hook Effect (Common Units) | High-Dose Hook Effect (SI Units) | Standardization Reference Method/Material | IFU Reference | IFU Document Link | Weight (in LB) | Length (in IN) | Heigth (in IN) | Width (in IN) | Weight (in KG) | Length (in CM) | Heigth (in CM) | Width (in CM)
Atellica® CH | 11097619 |  | pl22 | 00630414595610 | Global | Specific Proteins | No | Atellica CH Immunoglobulin A_2 (IgA_2) | IgA_2IgA_2 | 0907 | 600 | 4 x 150 | N/A | N/A | 10 | 30 | antihuman IgA (goat), lot specific | ;#Serum;#Lithium Heparin Plasma;# | up to 3 days at room temperature up to 7 days Frozen up to 6 months at -20 °C | 4 μL | No |  | PEG-enhanced immunoturbidimetric | 33.0 - 540.0 mg/dL | 0.33 - 5.40 g/L | 15.0-5,400.0 mg/dL | 0.15-54.00 g/L | 40.0-350.0 mg/dL for adults | 0.40-3.50 g/L for adults | N/A | N/A | N/A | Atellica CH LSP CAL | 11099434 | 28 days | Liquid | Ready to Use | 3 | 6 | Proteins derived from human serum | 60 | 30 |  | No |  | Yes | NA | 11099300 |  | 15.0-5,400.0 mg/dL | 0.15-54.00 g/L | No | NA | NA | 1.8 mg/dL | 0.6 mg/dL | NA | NA | 1.8 mg/dL | 0.6 mg/dL | NA | 0.01 g/L | 0.02 g/L | NA | 141.7 mg/dL | 1.42 g/L | 1.3 | 1.9 | 266.8 mg/dL | 2.67 g/L | 0.5 | 0.7 | 443.7 mg/dL | 4.44 g/L | 0.6 | 1 | ADVIA Chemistry 1800 IgA_2 y = 0.99x - 5.9 mg/dL r=1.000 n=80 | ADVIA Chemistry 1800 IgA_2 y = 0.99x - 0.06 g/L r=1.000 n=80 | <10% up to 1000 mg/dL of hemoglobin | <10% up to 50 mg/dL of bilirubin | <10% up to 922 mg/dL of Trig; <10% up to 500 mg/dL of Intralipid. | No Biotin in assay formulation | 9000.0 mg/dL | 90.00 g/L | 9,000.0 mg/dL | 90.00 g/L | IRMM reference material CRM-470 | 11110182_EN Rev. 03, 2019-05 | IFU Document | 0.572 | 5.512 | 3.543 | 2.756 | 0.26 | 14 | 9 | 7
Atellica® CH | 11097616 |  | pl24 | 00630414596266 | Global | Specific Proteins | No | Atellica CH Immunoglobulin G_2 (IgG_2) | IgG_2IgG_2 | 0910 | 720 | 4 x 180 | N/A | N/A | 10 | 30 | antihuman IgG (goat), lot-specific | ;#Serum;#Lithium Heparin Plasma;# | up to 7 days at 2-8 °C Frozen up to 3 months at -20°C | 2 μL | No |  | PEG-enhanced immunoturbidimetric | 140 - 3400 mg/dL | 1.40 – 34.00 g/L | 70-34,000 mg/dL | 0.70-340.00 g/L | 650-1,600 mg/dL for adults | 6.50-16.00 g/L for adults | N/A | N/A | N/A | Atellica CH LSP CAL | 11099434 | 28 days | Liquid | Ready to Use | 3 | 6 | Proteins derived from human serum | 60 | 30 |  | No |  | Yes | NA | 11099300 |  | 70-34,000 mg/dL | 0.70-340.00 g/L | No | NA | NA | 21 mg/dL | 7 mg/dL | NA | NA | 21 mg/dL | 7 mg/dL | NA | 0.07 g/L | 0.21 g/L | NA | 806 mg/dL | 8.06 g/L | 1.6 | 1.8 | 1112 mg/dL | 11.12 g/L | 1.2 | 1.4 | 2309 mg/dL | 23.09 g/L | 0.8 | 1.6 | ADVIA Chemistry 1800 IgG_2 y = 0.99x + 4 mg/dL r = 0.999 n = 80 | ADVIA Chemistry 1800 IgG_2 y = 0.99x + 0.04 g/L)\ r = 0.999 n = 80 | <10% up to 1000 mg/dL of hemoglobin | <10% up to 50 mg/dL of bilirubin | <10% up to 1000 mg/dL of Intralipid. | No Biotin in assay formulation | 70,000 mg/dL | 700 g/L | 70,000 mg/dL | 700 g/L | IRMM reference material CRM-470 | 11110179_EN Rev. 03, 2019-05 | IFU Document | 0.614 | 5.625 | 3.5 | 2.75 | 0.279 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097620 |  | pl04 | 00630414595627 | Global | Specific Proteins | No | Atellica CH Immunoglobulin M_2 (IgM_2) | IgM_2IgM_2 | 0911 | 720 | 4 x 180 | N/A | N/A | 10 | 30 | antihuman IgM (goat), lot-specific | ;#Serum;#Lithium Heparin Plasma;#Potassium EDTA Plasma;# | up to 7 days at 2-8 °C Frozen up to 3 months at -20°C | 9 μL | No |  | PEG-enhanced immunoturbidimetric | 21.0–330.0 mg/dL | 0.21–3.30 g/L | 8.0-3,300.0 mg/dL | 0.08-33.00 g/L | 50.0-300.0 mg/dL for adults | 0.50-3.00 g/L for adults | N/A | N/A | N/A | Atellica CH LSP CAL | 11099434 | 28 days | Liquid | Ready to Use | 3 | 6 | Proteins derived from human serum | 60 | 30 |  | No |  | Yes | NA | 11099300 |  | 8.0-3,300.0 mg/dL | 0.08-33.00 g/L | No | NA | NA | 3.9 mg/dL | 1.9 mg/dL | NA | NA | 3.9 mg/dL | 1.9 mg/dL | NA | 0.02 g/L | 0.04 g/L | NA | 85.8 mg/dL | 0.86 g/L | 1.1 | 2.5 | 151.9 mg/dL | 1.52 g/L | 0.6 | 1.8 | 260.5 mg/dL | 1.52 g/L | 0.6 | 1.3 | ADVIA Chemistry 1800 IGM_2 y = 1.01x + 2.5 mg/dL r = 0.999 n = 80 | ADVIA Chemistry 1800 IGM_2 y = 1.01x + 0.02 g/L r = 0.999 n = 80 | <10% up to 1000 mg/dL of hemoglobin | <10% up to 60 mg/dL of bilirubin | <10% up to 1000 mg/dL of Intralipid. | No Biotin in assay formulation | 7,000.0 mg/dL | 70.00 g/L | 7,000.0 mg/dL | 70.00 g/L | IRMM reference material ERMDA-470 | 11110183_EN Rev. 03, 2019-05 | IFU Document | 0.605 | 5.625 | 3.5 | 2.75 | 0.275 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097601 |  | ge33 | 00630414596402 | Global | General Chemistry | No | Atellica CH Iron_2 (Iron_2) | Iron_2Iron_2 | 0912 | 1792 | 4 x 448 | N/A | N/A | 10 | 30 | Acetate buffer; Ferrozine; ascorbic acid; guanidinium chloride | ;#Serum;#Lithium Heparin Plasma;# | up to 4 days at 25 °C, up to 7 days at 2-8 °C or stored frozen up to 60 days at -20 °C | 25 μL | No |  | Ferrozine | 2 - 1000 µg/dL | 0.4 - 179.0 µmol/L | 2000 μg/dL | 358.0 μmol/L | Male serum&nbsp; 65-175 μg/dL Female serum: 50-170 μg/dL | Male serum:11.6 - 31.3 μmol/L Female serum; 9.0 - 30.4 μmol/L | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 days | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 180 | 30 |  | No |  | Yes | N/A | 11099300 |  | 2 - 2000 μg/dL | 0.4 - 358.0 μmol/L | No | N/A | N/A | 2 μg/dL | 0.4 μmol/L | N/A | N/A | 2 μg/dL | 1 μg/dL | N/A | 0.2 μmol/L | 0.4 μmol/L | N/A | 62 μg/dL | 11.1 μmol/L | 0.8 | 2.3 | 253 μg/dL | 45.3 μmol/L | 0.3 | 0.6 | 813 μg/dL | 145.5 μmol/L | 0.3 | 0.4 | ADVIA Chemistry 1800 Iron_2 y = 0.99x - 2 μg/dL r=0.994 n=103 | ADVIA Chemistry 1800 Iron_2 y = 0.99x - 0.4 μmol/L r=0.994 n=103 | N/A | <10% up to 50 mg/dL of bilirubin | <10% up to 500 mg/dL of Intralipid. | No Biotin in assay formulation | N/A | N/A | N/A | N/A | NIST SRM 937 | 11110164_EN Rev. 02, 2019-07 | IFU Document | 0.977 | 5.625 | 3.5 | 2.75 | 0.444 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097614 |  | ge35 | 00630414596280 | Global | General Chemistry | No | Atellica CH Lactate (Lac) | LacLac | 1201 | 380 | 4 x 95 | N/A | N/A | 4 | 30 per pack | PIPES buffer, TOOS, Peroxidase (plant), lactate oxidase and ascorbate oxidase (Microorganism);4-aminoantipyrine | ;#Lithium Heparin Plasma;#Potassium EDTA Plasma;#Sodium fluoride/potassium oxalate Plasma;# | up to 1 day at 2-8 °C or stored frozen up to 30 days at -20 °C | 4 μL | No |  | lactate oxidase to pyruvate and hydrogen peroxide | 0.6 - 110.0 mg/dL | 0.07 - 12.21 mmol/L | N/A | N/A | 4.5-19.8 mg/dL | 0.50-2.20 mmol/L | N/A | N/A | N/A | Atellica CH SPCL CHEM CAL | 11099438 | 7 days | Lyophilized | 5.0 mL of distilled water; Stand for 30 minutes at room temperature; Invert to mix Prior to use mix the contents by gently inverting the vials | 3 | 2 | Human serum | 62 | 30 |  | No |  | No | N/A | 11099300 |  | N/A | N/A | No | N/A | N/A | 0.6 mg/dL | 0.07 mmol/L | N/A | N/A | 0.6 mg/dL | 0.4 mg/dL | n/a | 0.04 mmol/L | 0.07 mmol/L | n/a | 12.5 mg/dL | 1.39 mmol/L | 1.9 | 4.9 | 48.7 mg/dL | 5.41 mmol/L | 0.9 | 4.8 | 99.0 mg/dL | 10.99 mmol/L | 0.8 | 4.6 | ADVIA Chemistry 1800 Lac y = 1.01x - 0.2 mg/dL r=0.995 n =101 | ADVIA Chemistry 1800 Lac y = 1.01x - 0.02 mmol/L r=0.995 n =101 | <10% up to 300 mg/dL of hemoglobin | <10% up to 3.75 mg/dL of bilirubin | <10% up to 1000 mg/dL of Intralipid. | No Biotin in assay formulation | N/A | N/A | N/A | N/A | Internal standard using highly purified material | 11110177_EN Rev. 02, 2019-07 | IFU Document | 0.442 | 5.625 | 3.5 | 2.75 | 0.201 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097606 |  | ge36 | 00630414596358 | Global | General Chemistry | No | Atellica CH Lipase (Lip) | LipLip | 1205 | 1280 | 4 x 320 | N/A | N/A | 8 | 9 | TAPS; -tartaric acid; sodium hydroxide; colipase (porcine); 2-Propanol; DGGMR; | ;#Serum;#Lithium Heparin Plasma;# | up to 24 hours at room temperature, up to 7 days at 2-8 °C or stored Frozen up to a year at -20 °C or colder | 3 μL | No |  | Colorimetric rate | 8 - 700 U/L | 8 - 700 U/L | 9- 3500 U/L | 9- 3500 U/L | 12 - 53 U/L | 12 - 53 U/L | N/A | N/A | N/A | Atellica CH ENZ 1 CAL | 11099317 | 30 days | Liquid | No preparation required | 3 | 2 | Bovine Serum, (lipase (porcine pancreas) | 61 | 9 |  | No |  | Yes | N/A | 11099300 |  | 9- 3500 U/L | 9- 3500 U/L | No | N/A | N/A | 6 U/L | 3 U/L | N/A | N/A | 6 U/L | 3 U/L | N/A | 3 U/L | 6 U/L | N/A | 36 U/L | 36 U/L | 3.1 | 4.1 | 204 U/L | 204 U/L | 0.9 | 2.8 | 655 U/L | 655 U/L | 0.6 | 2.5 | ADVIA Chemistry 1800 y = 0.97x - 1 U/L r=1.000 | ADVIA Chemistry 1800 y = 0.97x - 1 U/L r=1.000 | <10% up to 200 mg/dL of hemoglobin | <10% up to 25 mg/dL of bilirubin | <10% up to 1000 mg/dL of Intralipid. | No Biotin in assay formulation | N/A | N/A | N/A | N/A | Internal standard | 11110169_EN Rev. 03, 2019-06 | IFU Document | 0.862 | 5.625 | 3.5 | 2.75 | 0.392 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097612 |  | ge37 | 00630414595597 | Global | General Chemistry | No | Atellica CH Magnesium (Mg) (serum/plasma) | MgMg | 1301 | 1200 | 3 x 400 | N/A | N/A | 8 | 14 | Xylidyl blue | ;#Serum;#Lithium Heparin Plasma;#Urine;# | Up to 7 days at 2-8 °C; Frozen up to 12 months at -20 °C or colder | 12 μL | No |  | Xylidyl Blue | 0.50 - 5.00 mg/dL | 0.21 - 2.06 mmol/L | 10.00 mg/dL | 4.11 mmol/L | Adult 1.60 - 2.60 mg/dL | Adult 0.66 - 1.07 mmol/L | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 days | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 180 | 3 |  | No |  | Yes | Atellica CH Diluent | 11099300 |  | 0.50 - 10.00 mg/dL | 0.21 - 4.11 mmol/L | No | N/A | N/A | 0.02 mg/dL | 0.01 mmol/L | N/A | N/A | 0.02 mg/dL | 0.00 mg/dL | 0.46 mg/dL | 0.00 mmol/L | 0.01 mmol/L | 0.19 mmol/L | 0.78 mg/dL | 0.32 mmol/L | 3 | 3.9 | 1.51 mg/dL | 0.62 mmol/L | 2.3 | 3.4 | 4.22 mg/dL | 1.74 mmol/L | 0.6 | 1.1 | Dimension y = 0.94x + 0.09 mg/dL (y = 0.94x + 0.04 mmol/L) r=0.996 n=108 | Dimension y = 0.94x + 0.09 mg/dL (y = 0.94x + 0.04 mmol/L) r=0.996 n=108 | <10% up to 750 mg/dL of hemoglobin | <10% up to 30 mg/dL of bilirubin | <10% up to 500 mg/dL of Intralipid. | No Biotin in assay formulation | N/A | N/A | N/A | N/A | Atomic absorption reference method/ NIST materials | 11110175_EN Rev. 02, 2019-06 | IFU Document | 0.502 | 4.187 | 3.5 | 2.75 | 0.228 | 10.636 | 8.89 | 6.985
Atellica® CH | 11097612 |  | ge37 | 00630414595597 | Global | General Chemistry | No | Atellica CH Magnesium (Mg) (urine) | MgMg | 1301 | 1200 | 3 x 400 | N/A | N/A | 8 | 14 | Xylidyl blue | ;#Serum;#Lithium Heparin Plasma;#Urine;# | Up to 7 days at 2-8 °C: | 12 μL | No |  | Xylidyl Blue | Urine 1.00 - 14.00 mg/dL | 0.41 - 5.75 mmolL | 28.00 mg/dL | 11.51 mmol/L | Adult: 24 - 255 mg/24 hours | Adult: 0.99 - 10.45 mmol/24 hours | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 days | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 180 | 3 |  | No |  | Yes | Atellica CH Diluent | 11099300 |  | 1.00 - 28.00 mg/dL | 0.41 - 11.51 mmol/L | No | N/A | N/A | 0.04 mg/dL | 0.02 mmol/L | N/A | N/A | 0.04 mg/dL | 0.00 mg/dL | 0.57 mg/dL | 0.00 mmol/L | 0.02 mmol/L | 0.23 mmol/L | 4.61 mg/dL | 1.89 mmol/L | 0.8 | 2.1 | 11.19 mg/dL | 4.60 mmol/L | 1 | 1.3 | N/A | N/A |  |  | Dimension y = 0.96x - 0.06 mg/dL (y = 0.96x - 0.02 mmol/L) r=0.998 n=100 | Dimension y = 0.96x - 0.06 mg/dL (y = 0.96x - 0.02 mmol/L) r=0.998 n=100 | N/A | N/A | N/A | No Biotin in assay formulation | N/A | N/A | N/A | N/A | Atomic absorption reference method/ NIST materials | 11110175_EN Rev. 02, 2019-06 | IFU Document | 0.502 | 4.187 | 3.5 | 2.75 | 0.228 | 10.636 | 8.89 | 6.985
Atellica® CH | 11097531 |  | ge42 | 00630414595818 | Global | General Chemistry | No | Atellica CH Total Bilirubin_2 (TBil_2) | TBil_2TBil_2 | 0207 | 1792 | 4 x 448 | N/A | N/A | 10 | 30 | sodium metavanadate | ;#Serum;#Lithium Heparin Plasma;# | Protect daylight and fluorescent light 4 °C and analyzed within 5 days Stored frozen up to 3 months at -70 °C | 14.3 μL | No |  | Vanadate oxidation | 0.15 - 35.0 mg/dL | 3 - 599 µmol/L | 70.0 mg/dL | 1197 μmol/L | 0–1 day: <8.0 mg/dL 1–2 days: <12.0 mg/dL 3–5 days: <16.0 mg/dL > 5 days–60 years: 0.3–1.2 mg/dL 60–90 years: 0.2–1.1 mg/dL >90 years: 0.2–0.9 mg/dL | 0–1 day: <137 μmol/L 1–2 days: <205 μmol/L 3–5 days: <274 μmol/L > 5 days–60 years: 5 – 21 μmol/L 60–90 years: 3 –19 μmol/L >90 years: 3 – 15 μmol/L | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 8 hours | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 60 | 30 |  | No |  | Yes | N/A | 11099300 |  | 0.15 - 70.0 mg/dL | 3 - 1197 μmol/L | No | N/A | N/A | 0.0 mg/dL | 0 μmol/L | N/A | N/A | 0.0 mg/dL | 0.0 mg/dL | 0.1 mg/dL | 0 μmol/L | 0 μmol/L | 1.71 μmol/L | 0.8 mg/dL | 14 μmol/L | 2.1 | 6.3 | 1.5 mg/dL | 26 μmol/L | 1.1 | 3.3 | 25.2 mg/dL | 431 μmol/L | 0.2 | 1 | ADVIA Chemistry 1800 TBiL_2 y = 0.96x + 0.0 mg/dL r=1.000 n=109 see IFU for Neonatal | ADVIA Chemistry 1800 TBiL_2 y = 0.96x + 0 μmol/L r=1.000 n=109 see IFU for Neonatal | <10% up to 1000 mg/dL of hemoglobin | N/A | ≤10% up to 750 mg/dL of Lipemia (Triglycerides concentrate) | No Biotin in assay formulation | N/A | N/A | N/A | N/A | AACC reference method which uses reference materials from the National Institute of Standards and Technology (NIST SRM 916 | 11110149_EN Rev. 04, 2020-08 | IFU Document | 0.942 | 5.625 | 3.5 | 2.75 | 0.428 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097525 |  | ge07 | 00630414595535 | Global | General Chemistry | No | Atellica CH Total Iron Binding Capacity (TIBC) | TIBCTIBC | 2016 | 800 | 4 x 200 | N/A | N/A | 10 | 7 | Chromazurol B; cetrimide; ; ferric chloride; acetate buffer; Sodium bicarbonate buffer | ;#Serum;# | Separated specimens may be stored for up to 4 days at room temperature or for up to 7 daysat 2–8°C or stored frozen for up to 2 months at -20°C. | 24 μL | No |  | Sequential release and uptake of iron | 40 – 670 μg/dL | 7.16 –119.93 μmol/L | N/A | N/A | 250–425 μg/dL . | 44.75–76.08 μmol/L. | N/A | N/A | N/A | Atellica CH SPCL CHEM CAL | 11099438 | 7 | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Human serum | 180 | 7 |  | No |  | No | N/A | 11099300 |  | N/A | N/A | No | N/A | N/A | 9 μg/dL | 1.61 μmol/L | N/A | N/A | 9 μg/dL | 6 μg/dL | N/A | 1.07 μmol/L | 1.61 μmol/L | N/A | 222 μg/dL | 39.74 μmol/L | 0.8 | 1.3 | 370 μg/dL | 66.23 μmol/L | 0.4 | 0.7 | 536 μg/dL | 95.94 μmol/L | 0.4 | 0.6 | ADVIA Chemistry TIBC (serum) y = 1.05x - 20 μg/dL r=0.995 n=137 | ADVIA Chemistry TIBC (serum) y = 1.05x - 3.58 μmol/L r=0.995 n=137 | <10% up to 1000 mg/dL of hemoglobin | <10% up to 25 mg/dL of bilirubin | <10% up to 1000 mg/dL of Trig; <10% up to 375 mg/dL of Intralipid. | No Biotin in assay formulation | N/A | N/A | N/A | N/A | internal standard. | 11110143_EN Rev. 02, 2019-11 | IFU Document | 0.678 | 5.625 | 3.5 | 2.75 | 0.308 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097604 |  | ge44 | 00630414596372 | Global | General Chemistry | No | Atellica CH Total Protein II (TP) | TPTP | 2017 | 7400 | 4 x 1850 | N/A | N/A | 10 | 90 | Sodium hydroxide; Na-K-tartrate;potassium iodide; cupric sulfate | ;#Serum;#Lithium Heparin Plasma;# | up to 8 hours at room temperature, up to 3 days at 2-8 °C or stored frozen up to 180 days at -20 °C | 17.5 μL | No |  | biuret reagent (cupric sulfate in an alkaline solution) | 2.0 - 12.0 g/dL | 20 - 120 g/L | 24.0 g/dL | 240 g/L | 5.7-8.2 g/dL | 57-82 g/L | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 days | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 181 | 30 |  | No |  | Yes | N/A | 11099300 |  | N/A | N/A | No | N/A | N/A | 0.7 g/dL | 7 g/L | N/A | N/A | 0.7 g/dL | 0.6 g/dL | N/A | 6 g/L | 7 g/L | N/A | 4.0 g/dL | 40 g/L | 0.8 | 1 | 7.7 g/dL | 77 g/L | 1.7 | 1.7 | 9.6 g/dL | 96 g/L | 0.5 | 1.3 | ADVIA Chemistry 1800 Total Protein II y = 0.98x + 0.1 g/dL r=0.997 n=100 | ADVIA Chemistry 1800 Total Protein II y = 0.98x + 10 g/L r=0.997 n=100 | <10% up to 500 mg/dL of hemoglobin | <10% up to 25 mg/dL of bilirubin | <10% up to 500 mg/dL of Intralipid. | No Biotin in assay formulation | N/A | N/A | N/A | N/A | NIST SRM 927 | 11110167_EN Rev. 03, 2020-08 | IFU Document | 1.223 | 5.625 | 3.5 | 2.75 | 0.556 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097613 |  | pl34 | 00630414596297 | Global | Specific Proteins | No | Atellica CH Transferrin (Trf) | TrfTrf | 2018 | 880 | 4 x 220 | N/A | N/A | 10 | 30 | Antihuman transferrin antibody (goat) | ;#Serum;#Lithium Heparin Plasma;#Potassium EDTA Plasma;# | Separated specimens store for up to 8 hours at room temperature and up to 7 days at 2–8°C Frozen for up to 1 month at -20°C or 1 year at -70°C. | 2 μL | No |  | PEG-enhanced immunoturbidimetric | 1 - 440 mg/dL | 0.01 - 4.40 g/L | 4400 mg/dL | 44.00 g/L | Males Serum/Plasma 215-365 mg/dL Females Serum/Plasma 250-380 mg/dL | Males Serum/Plasma 2.15–3.65 g/L Females Serum/Plasma 2.50–3.80 g/L | N/A | N/A | N/A | Atellica CH LSP CAL | 11099434 | 28 days | Liquid | Ready to Use | 3 | 6 | Proteins derived from human serum | 60 | 14 |  | No |  | Yes | NA | 11099300 |  | 1 - 4400 mg/dL | 0.01 - 44.00 g/L | No | NA | NA | 1.0 mg/dL | 0.01 g/L | NA | NA | 1.0 mg/dL | 0 mg/dL | NA | 0.00 g/L | 0.01 g/L | NA | 130 mg/dL | 1.3 g/L | 0.7 | 1.9 | 257mg/dL | 2.57 g/L | 0.8 | 1.2 | 322 mg/dL | 3.22 g/L | 0.9 | 1.3 | ADVIA Chemistry 1800 TRF y = 0.98x - 3 mg/dL r = 0.997 n = 80 | ADVIA Chemistry 1800 TRF y = 0.98x - 0.03 g/L r = 0.997 n = 80 | <10% up to 525 mg/dL of hemoglobin | <10% up to 30 mg/dL of bilirubin | <10% up to 650 mg/dL of Intralipid. | No Biotin in assay formulation | 7200 mg/dL | 72.00 g/L | 7200 mg/dL | 72.00 g/L | Institute for Reference Materials and Measurements (IRMM) reference material CRM-470 from IFCC. | 11110176_EN Rev. 03, 2019-04 | IFU Document | 0.662 | 5.625 | 3.5 | 2.75 | 0.301 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097591 |  | ge43 | 00630414596495 | Global | General Chemistry | No | Atellica CH Triglycerides (concentrated) (Trig) | TrigTrig | 2020 | 2000 | 4 x 500 | N/A | N/A | 4 | 30 | 4‑Chlorophenol; peroxidase (plant) ; glycerol kinase (microorganism); glycerol‑3‑phosphate oxidase (microorganism); 4‑aminophenazone; lipase (microorganism) | ;#Serum;#Lithium Heparin Plasma;#Potassium EDTA Plasma;# | up to 7 days at 2-8 °C; Frozen up to 3 months at -15 to -25 °C | 4 μL | No |  | GPO | 10 - 550 mg/dL | 0.11 - 6.22 mmol/L | 1100 mg/dL | 12.43 mmol/L | Adult Serum Normal:150 mg/dL Borderline high: 150-199 mg/dL High: 200-499 mg/dL Very high: ≥ 500 mg/dL | Adult Serum: Normal: 1.70 mmol/L Borderline high: 1.70-2.25 mmol/L High: 2.26-5.64 mmol/L Very high: ≥ 5.65 mmol/L | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 days | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 60 | 14 |  | No |  | Yes | N/A | 11099300 |  | 10 - 1100 mg/dL | 0.11 - 12.43 mmol/L | No | N/A | N/A | 4 mg/dL | 0.05 mmol/L | N/A | N/A | 4 mg/dL | 3 mg/dL | 10 mg/dL | 0.03 mmol/L | 0.05 mmol/L | 0.11 mmol/L | 106 mg/dL | 1.2 mmol/L | 0.8 | 2.5 | 228 mg/dL | 2.58 mmol/L | 0.5 | 1.3 | 424 mg/dL | 4.79 mmol/L | 0.5 | 1 | ADVIA Chemistry 1800 TRIG_2 y = 0.98x + 0.5 mg/dL r=0.999 n=100 | ADVIA Chemistry 1800 TRIG_2 y = 0.98x + 0.01 mmol/L r=0.999 n=100 | <10% up to 1000 mg/dL of hemoglobin | <10% up to 30 mg/dL of bilirubin | N/A | No Biotin in assay formulation | N/A | N/A | N/A | N/A | NIST reference material SRM909c | 11110154_EN Rev. 03, 2019-11 | IFU Document | 0.365 | 2.875 | 3.5 | 2.75 | 0.166 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097608 |  | ge45 | 00630414596334 | Global | General Chemistry | No | Atellica CH Uric Acid (UA) | UAUA (serum) | 2102 | 4800 | 4 x 1200 | N/A | N/A | 7 | 30 | N-ethyl-N-(2-hydroxy-3-sulfopropyl)-3-methyl-aniline; 4-aminophenazone; peroxidase (plant); uricase (microorganism) | ;#Serum;#Lithium Heparin Plasma;# | Serum and plasma: up to 3–5 days at 4°C or Frozen for up to 6 months at -20°C. | 11 μL | No |  | Uricase/Peroxidase | 0.5 - 20 mg/dL | 30 - 1190 µmol/L | 100.0 mg/dL | 5950 μmol/L | Serum/Plasma Male: 3.7–9.2 mg/dL Female: 3.1–7.8 mg/dL | Serum/Plasma Male: 220–547 μmol/L Female: 184–464 μmol/L | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 days | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 183 | 7 |  | No |  | Yes | N/A | 11099300 |  | 0.5 - 100.0 mg/dL | 30 - 5950 μmol/L | No | N/A | N/A | 0.0 mg/dL | 0.0 μmol/L | N/A | N/A | 0.0 mg/dL | 0.0 mg/dL | N/A | 0.0 μmol/L | 0.0 μmol/L | N/A | 3.2 mg/dL | 190 μmol/L | 1 | 1.5 | 10.6 mg/dL | 631 μmol/L | 0.3 | 0.6 | 16.0 mg/dL | 952 μmol/L | 1.2 | 2 | ADVIA Chemistry 1800 UA y = 0.99x + 0.0 mg/dL r=1.000 n=100 | ADVIA Chemistry 1800 UA y = 0.99x + 0 μmol/L r=1.000 n=100 | <10% up to 250 mg/dL of hemoglobin | <10% up to 30 mg/dL of bilirubin | <10% up to 1000 mg/dL of Trig. | No Biotin in assay formulation | N/A | N/A | N/A | N/A | CDC candidate reference method which uses SRM 913 and SRM 909 reference materials from the National Institute of Standards and Technology | 11110171_EN Rev. 03, 2019-05 | IFU Document | 1.049 | 5.625 | 3.5 | 2.75 | 0.477 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097608 |  | ge45 | 00630414596334 | Global | General Chemistry | No | Atellica CH Uric Acid (UA) | UAUA (urine) | 2102 | 4800 | 4 x 1200 | N/A | N/A | 7 | 30 | N-ethyl-N-(2-hydroxy-3-sulfopropyl)-3-methyl-aniline; 4-aminophenazone; peroxidase; uricase | ;#Urine;# | 3–4 days at ambient temperature for alkaline urine. | 11 μL | No |  | Uricase/Peroxidase | 0.9 - 180 mg/dL | 54 - 10710 µmol/L | 360.0 mg/dL | 21420 μmol/L | Urine: 250–750 mg/day | Urine: 1.48–4.43 mmol/day | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 days | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 183 | 7 |  | No |  | Yes | N/A | 11099300 |  | 0.9 - 360.0 mg/dL | 54 - 21420 μmol/L | No | N/A | N/A | 0.5 mg/dL | 30 μmol/L | N/A | N/A | 0.5 mg/dL | 0.3 mg/dL | N/A | 18 μmol/L | 30 μmol/L | N/A | 12.9 mg/dL | 768 μmol/L | 2 | 2.9 | 21.1 mg/dL | 1255 μmol/L | 1.2 | 1.6 | N/A |  |  |  | ADVIA Chemistry 1800 UA y = 1.01x - 1.0 mg/dL r=0.997 n=100 | ADVIA Chemistry 1800 UA y = 1.01x - 60 μmol/L r=0.997 n=100 | N/A | N/A | N/A | No Biotin in assay formulation | N/A | N/A | N/A | N/A | CDC candidate reference method which uses SRM 913 and SRM 909 reference materials from the National Institute of Standards and Technology | 11110171_EN Rev. 03, 2019-05 | IFU Document | 1.049 | 5.625 | 3.5 | 2.75 | 0.477 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097593 |  | ge04 | 00630414596471 | Global | General Chemistry | No | Atellica CH Urea Nitrogen (UN_c) (serum/plasma) | UN_cUN_c (serum/plasma) | 2101 | 6240 | 4 x 1560 | N/A | N/A | 7 | 90 | NADH; Urease (plant); α‑ketoglutarate; GLDH (bovine) | ;#Serum;#Lithium Heparin Plasma;# | Separated blood urea nitrogen in separated serum plasma up to 3-5 days at room temperature or up to 7 days at 4 °C or stored frozen indefinitely at -20 °C; urine stored up to 4 days at 4-8 °C when preserved with thymol | 7 μL | No |  | Urease with GLDH | 5 - 150 mg/dL | 1.8 - 53.6 mmol/L | 300 mg/dL | 107.1 mmol/L | Adult 9-23 mg/dL | Adult: 3.2-8.2 mmol/L | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 days | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 75 | 6 |  | No |  | Yes | N/A | 11099300 |  | 5 - 300 mg/dL | 1.8 - 107.1 mmol/L | No | N/A | N/A | 2 mg/dL | 0.7 mmol/L | N/A | N/A | 2 mg/dL | 1 mg/dL | N/A | 0.4 mmol/L | 0.7 mmol/L | N/A | 16 mg/dL | 5.7 mmol/L | 3.1 | 3.4 | 67 mg/dL | 23.9 mmol/L | 0.7 | 1.6 | 118 mg/dL | 42.1 mmol/L | 0.8 | 2.5 | ADVIA Chemistry 1800 UN y = 0.96x + 2 mg/dL r=1.000 n=100 | ADVIA Chemistry 1800 UN y = 0.96x + 0.7 mmol/L r=1.000 n=100 | <10% up to 500 mg/dL of hemoglobin | <10% up to 30 mg/dL of bilirubin | <10% up to 650 mg/dL of Intralipid. | No Biotin in assay formulation | N/A | N/A | N/A | N/A | traceable to the CDC reference method, which uses SRM 912 and 909 reference materials from the National Institute of Standards and Technology (NIST). | 11110156_EN Rev. 03, 2020-08 | IFU Document | 1.21 | 5.625 | 3.5 | 2.75 | 0.55 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097593 |  | ge04 | 00630414596471 | Global | General Chemistry | No | Atellica CH Urea Nitrogen (UN_c) (urine) | UN_cUN_c (urine) | 2101 | 6240 | 4 x 1560 | N/A | N/A | 7 | 90 | NADH; Urease (plant); α‑ketoglutarate; GLDH (bovine) | ;#Urine;# | Separated blood urea nitrogen in separated serum plasma up to 3-5 days at room temperature or up to 7 days at 4 °C or stored frozen indefinitely at -20 °C; urine stored up to 4 days at 4-8 °C when preserved with thymol | 7 μL | No |  | Urease with GLDH | 35 - 1000 mg/dL | 12.5 - 357.0 mmol/L | 2000 mg/dL | 714 mmol/L | Adult 12-20 g/day | Adult: 0.43-0.71 mol/day | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 days | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 75 | 6 |  | No |  | Yes | N/A | 11099300 |  | 35 - 2000 mg/dL | 12.5 - 714 mmol/L | No | N/A | N/A | 7 mg/dL | 2.5 mmol/L | N/A | N/A | 7 mg/dL | 4 mg/dL | N/A | 1.4 mmol/L | 2.5 mmol/L | N/A | 158 mg/dL | 56.4 mmol/L | 1.4 | 3.8 | 538 mg/dL | 192.1 mmol/L | 1.5 | 2.9 | 815 mg/dL | 291.0 mmol/L | 0.9 | 2.5 | ADVIA Chemistry 1800 UN y = 0.95x + 1 mg/dL r=0.997 n=100 | ADVIA Chemistry 1800 UN y = 0.95x + 0.4 mmol/L r=0.997 n=100 | N/A | N/A | N/A | No Biotin in assay formulation | N/A | N/A | N/A | N/A | traceable to the CDC reference method, which uses SRM 912 and 909 reference materials from the National Institute of Standards and Technology (NIST). | 11110156_EN Rev. 03, 2020-08 | IFU Document | 1.21 | 5.625 | 3.5 | 2.75 | 0.55 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097610 |  | di01 | 00630414596310 | Global | General Chemistry | No | Atellica CH Microalbumin_2 (µALB_2) | μALB_2µALB_2 | 1308 | 840 | 4 x 210 | N/A | N/A | 10 | 30 | Antihuman albumin (goat - ploclonal | ;#Urine;# | 14 days at 2–8°C 5 months at -20°C | 13.7 μL | No |  | PEG-enhanced immunoturbidimetric | 0.3 - 38.0 mg/dL | 3 - 380 mg/L | 0.3 - 380.0 mg/dL | 3.0 - 3800 mg/L | < 30 mg/day | < 30 mg/day | N/A | N/A | N/A | Atellica CH μALB_2 CAL | 11099435 | 60 days | Liquid | Mix by gentle inversion at least five (5) times to ensure homogeneity prior to use. | 3 | 6 | Human serum albumin in buffered solution; sodium azide (0.09%) | 180 | 30 |  | No |  | Yes | N/A | 11099300 |  | 0.3 - 380.0 mg/dL | 3.0 - 3800 mg/L | No | N/A | N/A | 0.1 mg/dL | 1.0 mg/L | N/A | N/A | 0.1 mg/dL | 0.0 mg/dL | NA | 0.0 mg/L | 1.0 mg/L | N/A | 0.3 mg/dL | 30 mg/L | 1.4 | 3.6 | 29.2 mg/dL | 292 mg/L | 1 | 2 | 40.9 mg/dL | 409 mg/L | 1.2 | 1.4 | ADVIA Chemistry μALB_2 (urine) y = 1.05x + 0.1 mg/dL r = 0.999 n= 112 | ADVIA Chemistry μALB_2 (urine) y = 1.05x + 1.0 mg/L r = 0.999 n= 112 | ≤ 10% change up to 500 mg/dL of hemoglobin. | N/A | N/A | No Biotin in assay formulation | 20,000 mg/dL | 200,000 mg/L | N/A | N/A | Internal Standard | 11110173_EN Rev. 03, 2019-06 | IFU Document | 0.724 | 5.625 | 3.5 | 2.75 | 0.329 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097628 |  | sp02 | 00630414596228 | Global | Specific Proteins | No | Atellica CH Alpha-1-antitrypsin (AAT) | AATAAT | 1110 | 360 | 2 x 180 | N/A | N/A | 10 | 30 | Antihuman alpha-1-antitrypsin (goat) (polyclonal) | ;#Serum;#Lithium Heparin Plasma;# | 7 days at 4°C 3 months at -70°C. | 2.7 μL | No |  | PEG enhanced immunoturbidimetric | 5–500 mg/dL | 0.05–5.00 g/L | 1000 mg/dL | 10.0 g/L | 78–200 mg/dL for adults | 0.78–2.00 g/L for adults | N/A | N/A | N/A | Atellica CH Liquid Specific Protein Calibrators (LSP CAL) | 11099434 | 28 days | Liquid | No preparation required | 3 | 6 | Proteins derived from human serum; sodium azide | 60 | 7 |  | No |  | Yes | Atellica CH Diluent | 11099300 |  | 5-1000 mg/dL | 0.05-10.0 g/L | No | N/A | N/A | 3 mg/dL | 0.03 g/L | N/A | N/A | 3 mg/dL | 1.0 mg/dL | N/A | 0.01 g/L | 0.03 g/L | N/A | 76 mg/dL | 0.76 g/L | 0.9 | 2.1 | 177 mg/dL | 1.77 g/L | 0.7 | 2.7 | 352 mg/dL | 3.52 g/L | 1 | 1.8 | Serum: ADVIA Chemistry 1800 y = 0.99x - 4.0 mg/dL | Serum: ADVIA Chemistry 1800 y = 0.99x - 0.04 g/L | ≤ 10% change up to 1000 mg/dL of hemoglobin. | ≤ 10% change up to 25 mg/dL of bilirubin | ≤ 10% change up to 125 mg/dL of Intralipid. | No Biotin in assay formulation | 10,500 mg/dL | 105.0 g/L | 10,500 mg/dL | 105.0 g/L | Institute for Reference Materials and Measurements (IRMM) ERM‑DA 470 | 11110191_EN Rev. 03, 2019-04 | IFU Document | 0.319 | 2.875 | 3.5 | 2.75 | 0.145 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097625 |  | pl11 | 00630414596242 | Global | Specific Proteins | No | Atellica CH Apolipoprotein A-1 (APO A1) | APO A1APO A1 | 1120 | 300 | 2 x 150 | N/A | N/A | 10 | 14 | Antihuman Antihuman APO A1 (goat); Polyethylene glycol; Tris Buffer, sodium azide; | ;#Serum;#Lithium Heparin Plasma;# | 7 days at 2–8°C 3 months at -20°C. | 5 μL | No |  | PEG-enhanced immunoturbidimetric | 15 - 200 mg/dL | 0.15 - 2.00 g/L | 400 mg/dL | 4.00 g/L | Serum/Plasma: Adult Males: 79–169 mg/dL Adult Females: 76–214 mg/dL | Serum/Plasma: Adult Males: 0.79–1.69 g/L Adult Females:0.76–2.14 g/L | N/A | N/A | N/A | Atellica CH Apolipoprotein A‑1 & B Calibrator (APO A1 & B CAL) | 11099329 | 28 days | Lyophilized | Add 1.0 mL of distilled water into each vial. Replace the stopper. Let the vials stand for 30 min. at room temperature; Gently swirl and invert the vials. Users manually dilute the reconstituted calibrator 1:4 with isotonic saline | 3 | 5 | Lyophilized human serum | 60 | 14 |  | No |  | Yes | Atellica CH Diluent | 11099300 |  | 15-400 mg/dL | 0.15 - 4.00 g/L | No | N/A | N/A | 1.0 mg/dL | 0.01 g/L | N/A | N/A | 1.0 mg/dL | 0.0 mg/dL | N/A | 0.00 g/L | 0.01 g/L | N/A | 83 mg/dL | 0. 83 g/L | 1 | 2.4 | 121 mg/dL | 1.21 g/L | 0.8 | 2.2 | 166 mg/dL | 1.66 g/L | 0.7 | 1.7 | ADVIA Chemistry 1800 y =1.05x - 1.0 mg/dL r = 0.989 n = 50 | ADVIA Chemistry 1800 y =1.05x - 0.014 g/L r=0.989 | ≤ 10% change up to 525 mg/dL of hemoglobin. | ≤ 10% change up to 30 mg/dL of bilirubin | ≤ 10% change up to 650 mg/dL of Intralipid. | No Biotin in assay formulation | 2,205 mg/dL | 22.05 g/L | 2,205 mg/dL | 22.05 g/L | WHO-IFCC reference material SP1-01 | 11110188_EN Rev. 02, 2019-06 | IFU Document | 0.317 | 2.875 | 3.5 | 2.75 | 0.144 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097626 |  | pl01 | 00630414596235 | Global | Specific Proteins | No | Atellica CH Apolipoprotein B (APO B) | APO BAPO B | 1122 | 300 | 2 x 150 | N/A | N/A | 10 | 30 | Antihuman Antihuman APO B (goat); Polyethylene glycol; Tris Buffer, sodium azide; | ;#Serum;#Lithium Heparin Plasma;# | 7 days at 2–8°C 3 months at -20°C | 5 μL | No |  | PEG-enhanced immunoturbidimetric | 15 - 200 mg/dL | 0.15 - 2.00 g/L | 400 mg/dL | 4.00 g/L | Serum/ Plasma: Males: 46–174 mg/dL Females:46–142 mg/dL | Serum/ Plasma: Males: 0.46–1.74g/L Females: 0.46–1.42g/L | N/A | N/A | N/A | Atellica CH Apolipoprotein A‑1 & B Calibrator (APO A1 & B CAL) | 11099329 | 28 days | Lyophilized | Add 1.0 mL of distilled water; let the vials stand for 30 minutes at room temperature; gently swirl and invert the vials to ensure homogeneit. Reconstituted calibrator requires a 1:4 manual dilution (saline) prior to use. | 3 | 5 | Lyophilized human serum | 60 | 30 |  | No |  | Yes | Atellica CH Diluent | 11099300 |  | 15 - 400 mg/dL | 0.15 - 4.00 g/L | No | N/A | N/A | 2.0 mg/dL | 0.02 g/L | N/A | N/A | 2.0 mg/dL | 1.0 mg/dL | N/A | 0.01 g/L | 0.02 g/L | N/A | 61 mg/dL | 0. 61g/L | 1.6 | 6.1 | 120 mg/dL | 1.20 g/L | 0.9 | 3.5 | 158 mg/dL | 1.58 g/L | 0.6 | 2.8 | ADVIA Chemistry 1800 y = 0.99x - 4.0 mg/dL r = 0.999 | ADVIA Chemistry 1800 y = 0.99x - 0.04 g/L r = 0.999 | ≤ 10% change up to 525 mg/dL of hemoglobin. | ≤ 10% change up to 30 mg/dL of bilirubin | ≤ 10% change up to 650 mg/dL of Intralipid. | No Biotin in assay formulation | 5600 mg/dL | 56.0 g/L | 5600 mg/dL | 56.0 g/L | WHO-IFCC reference material SP3-07 | 11110189_EN Rev. 02, 2019-05 | IFU Document | 0.31 | 2.875 | 3.5 | 2.75 | 0.141 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097634 |  | sp03 | 00630414596167 | Global | Specific Proteins | No | Atellica CH Anti-Streptolysin-O_2 (ASO_2) | ASO_2ASO_2 | 0149 | 1200 | 3 x 400 | N/A | N/A | 10 | 30 | Latex particles coated with streptolysin-O antigen (microorganism), Glycerine buffer; | ;#Serum;#Lithium Heparin Plasma;#Potassium EDTA Plasma;# | 2 days at 2–8°C 6 months at -20°C | 7.5 μL | No |  | latex enhanced immunoturbidimetric | 25.0 - 1000 IU/mL | 25.0 - 1000 IU/mL | 2000 IU/mL | 2000 IU/mL | Children (preschool) ≤ 100.0 IU/mL SErum/ Plasma: Children (school age)≤ 250.0 IU/mL Adults≤ 200.0 IU/mL Siemens-determined adult reference range ≤ 194.0 IU/mL | Children (preschool) ≤ 100.0 IU/mL SErum/ Plasma: Children (school age)≤ 250.0 IU/mL Adults≤ 200.0 IU/mL Siemens-determined adult reference range ≤ 194.0 IU/mL | N/A | N/A | N/A | Atellica CH Liquid Specific Protein Calibrators (LSP CAL) | 11099434 | 28 days | Liquid | No preparation required | 3 | 2 | Proteins derived from human serum; sodium azide | 60 | 30 |  | No |  | Yes | Atellica CH Diluent | 11099300 |  | 25.0 - 2000 IU/mL | 25.0 - 2000 IU/mL | No | N/A | N/A | 12.7 IU/mL | 12.7 IU/mL | N/A | N/A | 12.7 IU/mL | 5.9 IU/mL | N/A | 5.9 IU/mL | 12.7 IU/mL | N/A | 177.5 IU/mL | 177.5 IU/mL | 2.7 | 3 | 402.1 IU/mL | 402.1 IU/mL | 2.1 | 2.3 | 607.1 IU/mL | 607.1 IU/mL | 1.1 | 2.4 | ADVIA Chemistry 1800 y = 1.01x - 1.0 IU/mL r= 0.999 n=100 | ADVIA Chemistry 1800 y = 1.01x - 1.0 IU/mL r= 0.999 | ≤ 10% change up to 1000 mg/dL of hemoglobin. | ≤ 10% change up to 60 mg/dL of bilirubin | ≤ 10% change up to 1000 mg/dL of Intralipid. | No Biotin in assay formulation | 8500 IU/mL. | 8500 IU/mL. | 8500 IU/mL. | 8500 IU/mL. | WHO 1st international reference material for ASO | 11110197_EN Rev. 03, 2019-11 | IFU Document | 0.777 | 4.187 | 3.5 | 2.75 | 0.353 | 10.636 | 8.89 | 6.985
Atellica® CH | 11097592 |  | ge29 | 00630414596488 | Global | General Chemistry | No | Atellica CH Glucose Hexokinase_3 (GluH_3) | GluH_3GluH_3 (serum/ plasma) | 0704 | 6240 | 4 x 1560 | N/A | N/A | 7 | 62 | hexokinase (microbial); G6PDH (microbial) | ;#Serum;#Lithium Heparin Plasma;#Potassium EDTA Plasma;#Sodium fluoride/potassium oxalate Plasma;# | 72 hours at 4°C variable stability is observed with longer storage conditions | 3.4 μL | No |  | Hexokinase | 4 -700 mg/dL | 0.2 - 38.9 mmol/L | 2100 mg/dL | 116.6 mmol/L | Adults; 74–106 mg/dL Newborn; 40–60 mg/dL Newborn; 1 day 50–80 mg/dL Child 60–100 mg/dL | Adults; 4.1–5.9 mmol/L Newborn 2.2–3.3 mmol/L Newborn; 1 day; 2.8–4.4 mmol/L Child; 3.3–5.6 mmol/L | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 days | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 182 | 30 |  | No |  | Yes | N/A | 11099300 |  | 4 - 2100 mg/dL | 0.2 - 116.6 mmol/L | No | N/A | N/A | 1.0 mg/dL | 0.1 mmol/L | N/A | N/A | 1.0 mg/dL | 0.0 mg/dL | N/A | 0.0 mmol/L | 0.1 mmol/L | N/A | 59 mg/dL | 3.3 mmol/L | 0.8 | 2 | 90 mg/dL | 5.0 mmol/L | 0.7 | 1.4 | 292 mg/dL | 16.2 mmol/L | 0.5 | 1.1 | ADVIA Chemistry 1800 y = 1.03x - 1.0 mg/dL y = 1.03x - 0.1 mmol/L r=1.000 n=105 | ADVIA Chemistry 1800 y = 1.03x - 1.0 mg/dL y = 1.03x - 0.1 mmol/L r=1.000 n=105 | ≤ 10% change up to 1000 mg/dL of hemoglobin. | ≤ 10% change up to 30 mg/dL of bilirubin | ≤ 10% change up to 1000 mg/dL of Intralipid. | No Biotin in assay formulation | N/A | N/A | N/A | N/A | NIST SRM 965a | 11110155_EN Rev. 04, 2020-08 | IFU Document | 1.188 | 5.512 | 3.543 | 2.756 | 0.54 | 14 | 9 | 7
Atellica® CH | 11097592 |  | ge29 | 00630414596488 | Global | General Chemistry | No | Atellica CH Glucose Hexokinase_3 (GluH_3) | GluH_3GluH_3 (urine) | 0704 | 6240 | 4 x 1560 | N/A | N/A | 7 | 62 | hexokinase (microbial); G6PDH (microbial) | ;#Urine;# | 4°C at collection | 3.4 μL | No |  | Hexokinase | 4 -700 mg/dL | 0.2 - 38.9 mmol/L | 2100 mg/dL | 116.6 mmol/L | < 0.5 g/day | 2.8 mmol/day | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 days | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 182 | 30 |  | No |  | Yes | N/A | 11099300 |  | 4 - 2100 mg/dL | 0.2 - 116.6 mmol/L | No | N/A | N/A | 0.0 mg/dL | 0.0 mmol/L | N/A | N/A | 0.0 mg/dL | 0 mg/dL | N/A | 0.0 mmol/L | 0.0 mmol/L | N/A | 38 mg/dL | 2.1 mmol/L | 1.9 | 3.8 | 285 mg/dL | 15.8 mmol/L | 0.4 | 1.3 | N/A | N/A |  |  | ADVIA Chemistry 1800 y = 1.04x - 1.0 mg/dL y = 1.04x - 0.1 mmol/L r=1.000 n=100 | ADVIA Chemistry 1800 y = 1.04x - 1.0 mg/dL y = 1.04x - 0.1 mmol/L r=1.000 n=100 | N/A | N/A | N/A | No Biotin in assay formulation | N/A | N/A | N/A | N/A | NIST SRM 965a | 11110155_EN Rev. 04, 2020-08 | IFU Document | 1.188 | 5.512 | 3.543 | 2.756 | 0.54 | 14 | 9 | 7
Atellica® CH | 11097592 |  | ge29 | 00630414596488 | Global | General Chemistry | No | Atellica CH Glucose Hexokinase_3 (GluH_3) | GluH_3GluH_3 (CSF) | 0704 | 6240 | 4 x 1560 | N/A | N/A | 7 | 62 | hexokinase (microbial); G6PDH (microbial) | ;#Cerebrospinal fluid (CSF);# | analyze immediately centrifuge and store at 4°C or -20°C if delay is unavoidable | 3.4 μL | No |  | Hexokinase | 4 -700 mg/dL | 0.2 - 38.9 mmol/L | 2100 mg/dL | 116.6 mmol/L | Adult: 40-70 mg/dL Infant/child 60-80 mg/dL | Adult 2.2-3.9 mmol/L Infant/child3.3-4.4 mmol/L | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 days | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 182 | 30 |  | No |  | Yes | N/A | 11099300 |  | 4 - 2100 mg/dL | 0.2 - 116.6 mmol/L | No | N/A | N/A | 2.0 mg/dL | 0.1 mmol/L | N/A | N/A | 2.0 mg/dL | 1 mg/dL | N/A | 0.1 mmol/L | 0.1 mmol/L | N/A | 42 mg/dL | 2.3 mmol/L | 2 | 2.3 | 96 mg/dL | 5.3 mmol/L | 0.7 | 1.2 | N/A | N/A |  |  | ADVIA Chemistry 1800 y = 1.03x + 2.0 mg/dL y = 1.03x + 0.1 mmol/L r= 1.000 n=110 | ADVIA Chemistry 1800 y = 1.03x + 2.0 mg/dL y = 1.03x + 0.1 mmol/L r= 1.000 n=110 | N/A | N/A | N/A | No Biotin in assay formulation | N/A | N/A | N/A | N/A | NIST SRM 965a | 11110155_EN Rev. 04, 2020-08 | IFU Document | 1.188 | 5.512 | 3.543 | 2.756 | 0.54 | 14 | 9 | 7
Atellica® CH | 11097621 |  | ge31 | 00630414595634 | Global | General Chemistry | No | Atellica CH Glucose Oxidase (GluO) | GluOGluO (serum/plasma) | 0703 | 5600 | 4 x 1400 | N/A | N/A | 10 | 51 | glucose oxidase (microbial); ; peroxidase (horseradish) | ;#Serum;#Lithium Heparin Plasma;# | 8 hours at 25°C 72 hours at 4°C3 days at room temperature by addition of sodium iodoacetate or sodium fluoride | 4 μL | No |  | Glucose Oxidase Trinder | 6 -750 mg/dL | 0.3 - 41.6 mmol/L | 1,500 mg/dL | 83.3 mmol/L | 74-106 mg/dL | 4.1-5.9 mmol/L | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 days | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 35 | 7 |  | No |  | Yes | N/A | 11099300 |  | 6 - 1,500 mg/dL | 0.3 - 83.3 mmol/L | No | N/A | N/A | 1 mg/dL | 0.1 mmol/L | N/A | N/A | 1 mg/dL | 0 mg/dL | N/A | 0.0 mmol/L | 0.1 mmol/L | N/A | 61 mg/dL | 3.4 mmol/L | 1 | 2.3 | 261 mg/dL | 14.5 mmol/L | 0.3 | 2 | 567 mg/dL | 31.5 mmol/L | 0.3 | 2.9 | ADVIA Chemistry 1800 GLUO y = 0.99x - 3.0 mg/dL r = 0.999 n= 101 | ADVIA Chemistry 1800 GLUO y = 0.99x - 0.2 mmol/L r = 0.999 n= 101 | ≤ 10% change up to 200 mg/dL of hemoglobin. | ≤ 10% change up to 7.5 mg/dL of bilirubin | ≤ 10% change up to 250 mg/dL of Intralipid. | No Biotin in assay formulation | N/A | N/A | N/A | N/A | NIST SRM 965a | 11110184_EN Rev. 03, 2020-08 | IFU Document | 0.535 | 2.875 | 3.5 | 2.75 | 0.243 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097621 |  | ge31 | 00630414595634 | Global | General Chemistry | No | Atellica CH Glucose Oxidase (GluO) | GluOGluO (urine) | 0703 | 5600 | 4 x 1400 | N/A | N/A | 10 | 51 | glucose oxidase (microbial); ; peroxidase (horseradish) | ;#Urine;# | 4°C at collection | 4 μL | No |  | Glucose Oxidase Trinder | 6 -750 mg/dL | 0.3 - 41.6 mmol/L | 1,500 mg/dL | 83.3 mmol/L | Adults: < 0.5 g/day | Adults: < 0.5 g/day | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 days | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 35 | 7 |  | No |  | Yes | N/A | 11099300 |  | 6 - 1,500 mg/dL | 0.3 - 83.3 mmol/L | No | N/A | N/A | 1 mg/dL | 0.1 mmol/L | N/A | N/A | 1 mg/dL | 0 mg/dL | N/A | 0.0 mmol/L | 0.1 mmol/L | N/A | 28 mg/dL | 1.6 mmol/L | 1.6 | 4.2 | 299 mg/dL | 16.6 mmol/L | 0.4 | 2 | N/A | N/A |  |  | ADVIA Chemistry 1800 GLUO y = 1.00x + 4.0 mg/dL r = 0.998 n=105 | ADVIA Chemistry 1800 GLUO y = 1.00x + 0.2 mmol/L r = 0.998 n=105 | N/A | N/A | N/A | No Biotin in assay formulation | N/A | N/A | N/A | N/A | NIST SRM 965a | 11110184_EN Rev. 03, 2020-08 | IFU Document | 0.535 | 2.875 | 3.5 | 2.75 | 0.243 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097621 |  | ge31 | 00630414595634 | Global | General Chemistry | No | Atellica CH Glucose Oxidase (GluO) | GluOGluO (CSF) | 0703 | 5600 | 4 x 1400 | N/A | N/A | 10 | 51 | glucose oxidase (microbial); ; peroxidase (horseradish) | ;#Cerebrospinal fluid (CSF);# | CSF should be analyzed immediately, centrifuged and stored at 4°C or -20°C if delay is unavoidable | 4 μL | No |  | Glucose Oxidase Trinder | 6 -750 mg/dL | 0.3 - 41.6 mmol/L | 1,500 mg/dL | 83.3 mmol/L | Adult: 40-70 mg/dL Infants/children: 60-80 mg/dL | Adult .2-3.9 mmol/L Infants/children: 3.3-4.4 mmol/L | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 days | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 35 | 7 |  | No |  | Yes | N/A | 11099300 |  | 6 - 1,500 mg/dL | 0.3 - 83.3 mmol/L | No | N/A | N/A | 1 mg/dL | 0.1 mmol/L | N/A | N/A | 1 mg/dL | 0 mg/dL | N/A | 0.0 mmol/L | 0.1 mmol/L | N/A | 53 mg/dL | 2.9 mmol/L | 0.8 | 2.6 | 104 mg/dL | 5.8 mmol/L | 0.5 | 3.5 | N/A | N/A |  |  | ADVIA Chemistry 1800 GLUO y = 0.98x + 3.0 mg/dL r = 1.000 n=122 | ADVIA Chemistry 1800 GLUO y = 0.98x + 0.2 mmol/L r = 1.000 n=122 | N/A | N/A | N/A | No Biotin in assay formulation | N/A | N/A | N/A | N/A | NIST SRM 965a | 11110184_EN Rev. 03, 2020-08 | IFU Document | 0.535 | 2.875 | 3.5 | 2.75 | 0.243 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097633 |  | pl06 | 00630414596174 | Global | Specific Proteins | No | Atellica CH High Sensitivity C-Reactive Protein (hsCRP) | hsCRPhsCRP | 0334 | 740 | 2 x 370 | N/A | N/A | 8 | 30 | CRP antibody synthetic latex (rabbit) | ;#Serum;#Lithium Heparin Plasma;#Potassium EDTA Plasma;# | 3 days at 4-8 °C 6 months at -20 °C | 10 μL | No |  | latex enhanced immunoturbidimetric | 0.16–10.0 mg/L | 0.16–10.0 mg/L | 200 mg/L | 200 mg/L | Healthy individuals ≤ 3 mg/L Low risk for cardiovascular disease prediction <1 mg/L Average risk for cardiovascular disease prediction >1 and <3 mg/L High risk for cardiovascular disease prediction >3 mg/L | Healthy individuals ≤ 3 mg/L Low risk for cardiovascular disease prediction <1 mg/L Average risk for cardiovascular disease prediction >1 and <3 mg/L High risk for cardiovascular disease prediction >3 mg/L | N/A | N/A | N/A | Atellica CH hsCRP CAL | 11099412 | 60 days | Liquid | Liquid | 3 | 6 | Recombinant human c‑reactive protein (CRP) in a stabilized protein matrix | 60 | 30 |  | No |  | Yes | NA | 11099300 |  | 0.16–200 mg/L | 0.16–200 mg/L | No | NA | NA | 0.07 mg/L | 0.05 mg/L | NA | NA | 0.07 mg/L | 0.05 mg/L | 0.16 mg/L | 0.05 mg/L | 0.07 mg/L | 0.16 mg/L | 0.24 mg/L | NA | 6.8 | 7.4 | 0.73 mg/L | NA | 2.2 | 2.7 | 1.52 mg/L | NA | 1.3 | 1.7 | ADVIA Chemistry 1800 y = 0.96x + 0.02 mg/L r = 0.990 | ADVIA Chemistry 1800 y = 0.96x + 0.02 mg/L r = 0.990 | ≤ 10% change up to 500 mg/dL of hemoglobin. | ≤ 10% change up to 30 mg/dL of bilirubin | ≤ 10% change up to 1000 mg/dL of Intralipid. | No Biotin in assay formulation | 2000 mg/L | 2000 mg/L | 2,000 mg/L | 2,000 mg/L | IRMM -CRM-470 | 11110196_EN Rev. 02, 2019-07 | IFU Document | 0.431 | 2.875 | 3.5 | 2.75 | 0.196 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097627 |  | sp04 | 00630414595559 | Global | Specific Proteins | No | Atellica CH Lipoprotein(a) (Lp(a)) | Lp(a)Lp(a) | 1210 | 400 | 2 x 200 | N/A | N/A | 10 | 30 | Anti-LPA antibody coupled to latex particles (rabbit), lot-specific | ;#Serum;#Lithium Heparin Plasma;# | 14 days at 2-8 °C 6 months at ≤ -20 °C | 9 μL | No |  | latex enhanced immunoturbidimetric | 10.00–85.00 mg/dL | 0.100–0.850 g/L | N/A | N/A | < 30.00 mg/dL | < 0.300 g/L | N/A | N/A | N/A | Atellica CH Lp(a) CAL | 11099433 | 14 days | Lyophilized | Add 1.0 mL of distilled water; Replace stopper; Let the vials stand for 10 minutes at room temperature; to mix gently inverting the vials at least 5 times | 3 | 6 | Lyophilized human serum | 60 | 30 |  | No |  | No | NA | 11099300 |  | N/A | N/A | No | NA | NA | 2.45 mg/dL | 1.86 mg/dL | NA | NA | 2.45 mg/dL | 1.86 mg/dL | 9.30 mg/dL | 0.019 g/L | 0.025 g/L | 0.093 mg/dL | 19.27 mg/dL | 0.193 g/L | 2 | 2 | 33.60 mg/dL | 0.336 g/L | 0.7 | 1 | 47.45 mg/dL | 0.475 g/L | 0.4 | 1.5 | ADVIA Chemistry 1800 y = 0.97x + 0.90 mg/dL r = 0.997 n = 80 | ADVIA Chemistry 1800 y = 0.97x + 0.009 g/L r = 0.997 n = 80 | ≤ 10% change up to 1000 mg/dL of hemoglobin. | ≤ 10% change up to 60 mg/dL of bilirubin | ≤ 10% change up to 1000 mg/dL of Intralipid. | No Biotin in assay formulation | 490.00 mg/dL | 4.900 g/L | 490.00 mg/dL | 4.900 g/L | Internal standard | 11110190_EN Rev. 02, 2019-05 | IFU Document | 0.334 | 2.875 | 3.5 | 2.75 | 0.152 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097617 |  | pl30 | 00630414596259 | Global | Specific Proteins | No | Atellica CH Prealbumin (PreAlb) | PreAlbPreAlb | 1608 | 400 | 2 x 200 | N/A | N/A | 10 | 30 | Antihuman prealbumin (goat) | ;#Serum;# | 8 hours at room temperature 2 days at 2-8 °C or frozen at -20 °C or colder | 7.5 μL | No |  | PEG-enhanced immunoturbidimetric | 5.0 - 70.0 mg/dL | 50 - 700 mg/L | 140.0 mg/dL | 1,400 mg/L | 10.0-40.0 mg/dL | 100-400 mg/L | N/A | N/A | N/A | Atellica CH LSP CAL | 11099434 | 28 days | Liquid | Ready to use | 3 | 6 | Proteins derived from human serum | 60 | 7 |  | No |  | Yes | N/A | 11099300 |  | 5.0-140.0 mg/dL | 50-1,400 mg/L | No | N/A | N/A | 1.8 mg/dL | 18 mg/L | NA | NA | 1.8 mg/dL | 0.0 mg/dL | N/A | 0 mg/L | 18 mg/L | NA | 20.8 mg/dL | 208 mg/L | 1.4 | 2.6 | 40.3 mg/dL | 403 mg/L | 1.2 | 1.7 | 60.3 mg/dL | 603 mg/L | 0.9 | 1.7 | ADVIA Chemistry 1800 y = 1.05x + 0.0 mg/dL r = 0.993 n = 80 | ADVIA Chemistry 1800 y = 1.05x + 0.0 mg/L r = 0.993 n = 80 | ≤ 10% change up to 1000 mg/dL of hemoglobin. | ≤ 10% change up to 25 mg/dL of bilirubin | ≤ 10% change up to 250 mg/dL of Intralipid. | No Biotin in assay formulation | 420.0 mg/dL | 4200 mg/L | 420.0 mg/dL | 4200 mg/L | IRMM CRM-470 from IFCC | 11110180_EN Rev. 04, 2020-06 | IFU Document | 0.317 | 2.875 | 3.5 | 2.75 | 0.144 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097618 |  | pl32 | 00630414595603 | Global | Specific Proteins | No | Atellica CH Rheumatoid Factor (RF) | RFRF | 1802 | 360 | 2 x 180 | N/A | N/A | 10 | 21 | BSA (bovine); IgG synthetic latex (human) | ;#Serum;# | 7 days at 2-8°C 3 months at -20°C | 11.5 μL | No |  | Latex-enhanced immunoturbidimetric assay | 3.5 - 90.0 IU/mL | 3.5 - 90.0 IU/mL | 900 IU/mL | 900 IU/mL | < 14 IU/mL | < 14 IU/mL | N/A | N/A | N/A | Atellica CH LSP CAL | 11099434 | 28 days | Liquid | Ready to Use | 3 | 6 | Proteins derived from human serum | 60 | 14 |  | No |  | Yes | NA | 11099300 |  | 3.5 - 900 IU/mL | 3.5 - 900 IU/mL | No | NA | NA | 3.2 IU/mL | 3.2 IU/mL | NA | NA | 3.2 IU/mL | 1.6 IU/mL | NA | 1.6 IU/mL | 3.2 IU/mL | NA | 20.8 IU/mL | 20.8 IU/mL | 2.5 | 3.3 | 50.9 IU/mL | 50.9 IU/mL | 0.5 | 1 | 71.1 IU/mL | 71.1 IU/mL | 0.6 | 1 | ADVIA Chemistry 1800 RF y = 0.99x + 0.9 IU/mL r = 0.993 n = 80 | ADVIA Chemistry 1800 RF y = 0.99x + 0.9 IU/mL r = 0.993 n = 80 | ≤ 10% change up to 1000 mg/dL of hemoglobin. | ≤ 10% change up to 25 mg/dL of bilirubin | ≤ 10% change up to 1000 mg/dL of Intralipid. | No Biotin in assay formulation | 450 IU/mL | 450 IU/mL | 450 IU/mL | 450 IU/mL | Internal standard | 11110181_EN Rev. 02, 2019-11 | IFU Document | 0.174 | 2.875 | 3.5 | 2.75 | 0.079 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097532 |  | ge06 | 00630414595801 | Global | General Chemistry | No | Atellica CH Direct Bilirubin 2 (DBil_2) | DBil_2DBil_2 | 0205 | 1792 | 4 x 448 | N/A | N/A | 10 | 30 | Tartrate buffer; Phosphate buffer: sodium metavanadate | ;#Serum;#Lithium Heparin Plasma;# | 5 days at 4°C; Frozen at-70°C up to 3 months | 14.3 μL | No |  | Vanadate oxidation | 0.1–15.0 mg/dL | 2–256 μmol/L | 0.1 - 22.5 mg/dL | 2 - 385 μmol/L | ≤ 0.3 mg/dL | ≤ 5 μmol/L | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 8 hours | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 60 | 30 |  | No |  | Yes | N/A | 11099300 |  | 0.1 - 22.5 mg/dL | 2 - 385 μmol/L | No | N/A | N/A | 0.0 mg/dL | 0.0 μmol/L | N/A | N/A | 0.0 mg/dL | 0.0 mg/dL | N/A | 0.0 μmol/L | 0.0 μmol/L | N/A | 0.4 mg/dL | 7 μmol/L | 1 | 1.6 | 3.5 mg/dL | 60 μmol/L | 1 | 1.6 | 11.9 mg/dL | 203 μmol/L | 0.2 | 0.5 | ADVIA Chemistry 1800 D_BIL_2 y = 1.03x + 0.0 mg/dL r=1.00 n=103 | ADVIA Chemistry 1800 D_BIL_2 y = 1.03x + 0 μmol/L r=1.00 n=103 | ≤ 10% change up to 1000 mg/dL of hemoglobin. | N/A | ≤ 10% change up to 750 mg/dL of Lipemia (Triglycerides Concentrate) | No Biotin in assay formulation | N/A | N/A | N/A | N/A | AACC reference method NIST SRM 916 reference materials | 11110150_EN Rev. 03, 2020-08 | IFU Document | 0.939 | 5.625 | 3.5 | 2.75 | 0.427 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097629 |  | sp01 | 00630414596211 | Global | Specific Proteins | No | Atellica CH Alpha-1-acid Glycoprotein (AAG) | AAGAAG | 1109 | 400 | 2 x 200 | N/A | N/A | 10 | 30 | Antihuman alpha-1-acid glycoprotein antibody (goat) (polyclonal) | ;#Serum;#Lithium Heparin Plasma;#Potassium EDTA Plasma;# | Up to 3 days at 4°C ; Frozen at -20°C for up to 6 months | 3.2 μL | No |  | PEG-enhaced immunoturbidimetric assay | 5.0–612.0 mg/dL | 0.05–6.12 g/L | 5.0 - 1836 mg/dL | 0.05 - 18.36 g/L | 58.2–154.8 mg/dL | 0.58–1.55 g/L | N/A | N/A | N/A | Atellica CH Liquid Specific Protein Calibrators (LSP CAL) | 11099434 | 28 days | Liquid | No preparation required | 3 | 6 | Proteins derived from human serum; sodium azide | 60 | 14 |  | No |  | Yes | Atellica CH Diluent | 11099300 |  | 5.0 - 1836 mg/dL | 0.05 - 18.36 g/L | No | N/A | N/A | 2.0 mg/dL | 0.02 g/L | N/A | N/A | 2.0 mg/dL | 0.9 mg/dL | N/A | 0.01 g/L | 0.02 g/L | N/A | 43.2 mg/dL | 0.43 g/dL | 1.4 | 2.3 | 94.4 mg/dL | 0.94 g/dL | 0.8 | 2.4 | 354.6 mg/dL | 3.55 g/dL | 0.8 | 1 | ADVIA Chemistry 1800 y = 1.01x + 0.6 mg/dL r=0.996 n = 101 | ADVIA Chemistry 1800 y = 1.01x + 0.01 g/dL r=0.996 n=101 | ≤ 10% at 1000 mg/dL of hemoglobin | ≤ 10% at 60 mg/dL of bilirubin | ≤ 10% at 1000 mg/dL of Intralipid | No Biotin in assay formulation | 2165.0 mg/dL | 21.65 g/L | 2165.0 mg/dL | 21.65 g/L | Institute for Reference Materials and Measurements (IRMM) ERM‑DA 470 | 11110192_EN Rev. 03, 2019-05 | IFU Document | 0.345 | 2.875 | 3.5 | 2.75 | 0.157 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097600 |  | ge01 | 00630414596419 | Global | General Chemistry | No | Atellica CH Alkaline Phosphatase, Concentrated (ALP_2c) | ALP_2cALP_2c | 1108 | 4800 | 4 x 1200 | N/A | N/A | 8 | 30 | R1: 2-amino-2-methyl-1-propanol (AMP) buffer; HEDTA; magnesium acetate;; zinc sulfate; R2: Paranitrophenyl-phosphate (p‑NPP) substrate; 6 mmol; sodium azide; ProClin 300 | ;#Serum;#Lithium Heparin Plasma;# | Up to 8 hours at 25°C; up to 7 days at 2–8°C Frozen at -20°C or colder for up to 6 months. | 10 μL | No |  | IFCC Standardization | 10 - 1000 U/L | 10 - 1000 U/L | 10 - 2300 U/L | 10 - 2300 U/L | 46–116 U/L | 46–116 U/L | N/A | N/A | N/A | Atellica CH ALP_2 CAL | 11099316 | 30 days | Liquid | No preparation required | 3 | 2 | Human serum albumin; alkaline phosphatase (porcine kidney) | 60 | 17 |  | No |  | Yes | N/A | 11099300 |  | 10 - 2300 U/L | 10 - 2300 U/L | No | N/A | N/A | 0.4 U/L | 0.4 U/L | N/A | N/A | 0.4 U/L | 0 U/L | 6.2 U/L | 0 U/L | 0.4 U/L | 6.2 U/L | 87 U/L | 87 U/L | 0.4 | 1.3 | 277 U/L | 277 U/L | 0.2 | 0.7 | 841 U/L | 841 U/L | 0.2 | 1.5 | Dimension ALPI y = 1.05x -4U/L r=1.000 n=104 | Dimension ALPI y = 1.05x -4U/L r=1.000 n=104 | ≤ 10% at 1000 mg/dL of hemoglobin | ≤ 10% at 80 mg/dL of bilirubin | ≤ 10% at 600 mg/dL of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | traceable to IFCC primary reference procedure for the measurement of catalytic activity of alkaline phosphatase at 37°C | 11110163_EN Rev. 03, 2019-07 | IFU Document | 1.063 | 5.625 | 3.5 | 2.75 | 0.483 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097529 |  | ge18 | 00630414595832 | Global | General Chemistry | No | Atellica CH Ammonia (Amm) | AmmAmm | 0116 | 480 | 4 x 120 | N/A | N/A | 4 | 30 | α-Ketoglutarate (10 mmol/L); glutamate dehydrogenase (24 kU/L); NADPH analog (0.2 mmol/L); | ;#Lithium Heparin Plasma;#Potassium EDTA Plasma;# | Up to 2 hours at 2–8°C; Samples should be analyzed within 30 minutes of centrifugation. | 25 μL | No |  | Enzymatic | 17 – 1277 μg/dL | 10-750 μmol/L | 17 - 2554 μg/dL | 10 - 1499 μmol/L | 19–54 μg/dL | 11–32 μmol/L | N/A | N/A | N/A | Atellica CH CHEM III CAL | 11099335 | 30 days | Liquid | No preparation required | 3 | 3 | Ammonium bicarbonate | 30 | 3 |  | No |  | Yes | N/A | 11099300 |  | 17 - 2554 μg/dL | 10 - 1499 μmol/L | No | N/A | N/A | 11 μg/dL | 6 μmol/L | N/A | N/A | 11 μg/dL | 3 μg/dL | 17 μg/dL | 2 μmol/L | 6 μmol/L | 10 μmol/L | 48 μg/dL | 28 μmol/L | 6.2 | 7.3 | 197 μg/dL | 115 μmol/L | 2.1 | 2.6 | 1050 μg/dL | 616 μmol/L | 0.4 | 1 | Dimension RxL y = 1.06x - 10 μg/dL r=0.999 n=102 | Dimension RxL; y = 1.06x - 6 μmol/L r=0.999 n=102 | 18% bias at 375 mg/dL of hemoglobin | 18% bias at 60 mg/dLof bilirubin | ≤ 10% at 200 mg/dL of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | ACS Grade Ammonium Sulfate | 11110147_EN Rev. 04, 2020-03 | IFU Document | 0.185 | 2.875 | 3.5 | 2.75 | 0.084 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097615 |  | ge20 | 00630414596273 | Global | General Chemistry | No | Atellica CH Cholinesterase (CHE) | CHECHE | 0339 | 270 | 2 x 135 | N/A | N/A | 9 | 30 per pack | Phosphate buffer; 5,5’-dithiobis (2-Nitrobenzoic acid); Butyrylthiocholine iodide | ;#Serum;#Lithium Heparin Plasma;# | Up to 6 hours at 25°C; up to 7 days at 2–8°C; Frozen at -70°C or colder for up to 6 months. | 2 μL | No |  | Butyrylthiocholine substrate | 1,500 - 30,000 U/L | 1,500 - 30,000 U/L | 1500 - 60,000 U/L | 1500 - 60,000 U/L | 7000–19,000 U/L | 7000–19,000 U/L | N/A | N/A | N/A | Atellica CH ENZ 1 CAL | 11099317 | 30 days | Liquid | No preparation required | 3 | 2 | Bovine Serum; CHE (equine serum) | 120 | 7 |  | No |  | Yes | N/A | 11099300 |  | 1500 - 60,000 U/L | 1500 - 60,000 U/L | No | N/A | N/A | 100 U/L | 100 U/L | N/A | N/A | 100 U/L | 42 U/L | N/A | 42 U/L | 100 U/L | N/A | 4270 U/L | 4270 U/L | 0.3 | 2.9 | 12,499 U/L | 12,499 U/L | 0.3 | 1.4 | 27,050 U/L | 27,050 U/L | 0.3 | 1 | ADVIA Chemistry 1800 y = 1.00x - 56 U/L r=0.997 n=100 | ADVIA Chemistry 1800 y = 1.00x - 56 U/L r=0.997 n=100 | ≤ 10% at 500 mg/dL of hemoglobin | ≤ 10% at 25 mg/dL of bilirubin | ≤ 10% at 500 mg/dL of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | Traceable to the molar extinction coefficient of the reactionproduct | 11110178_EN Rev. 01, 2017-05 | IFU Document | 0.442 | 5.625 | 4.125 | 2.75 | 0.201 | 14.287 | 10.477 | 6.985
Atellica® CH | 11097521 |  | ge24 | 00630414595900 | Global | General Chemistry | No | Atellica CH Carbon Dioxide, Concentrated (CO2_c) | CO2_cCO2_c | 2015 | 7600 | 4 x 1900 | N/A | N/A | 2 | 27 | malate dehydrogenase (MDH) (mammalian); PEP: microorganism (PEPC); | ;#Serum;#Lithium Heparin Plasma;# | Up to 3 days at 2–8°C; Frozen at-20°C for up to 60 days | 5.7 μL | No |  | Enzymatic | 10.0 - 40.0 mEq/L | 10.0 - 40.0 mmol/L | N/A | N/A | 20.0–31.0 mEq/L | 20.0–31.0 mmol/L | N/A | N/A | N/A | Atellica CH CO2 CAL | 11099401 | 30 days | Liquid | No preparation required | 3 | 2 | Not provided | 30 | 6 |  | No |  | No | N/A | 11099300 |  | N/A | N/A | No | N/A | N/A | 0.9 mEq/L | 0.9 mmol/L | N/A | N/A | 0.9 mEq/L | ≤ 0.4 mEq/L | N/A | ≤ 0.4mmol/L | 0.9 mmol/L | N/A | 17.3 mEq/L | 17.3 mmol/L | 1.11 | 3.45 | 24.4 mEq/L | 24.4 mmol/L | 1.73 | 4.13 | 31.8 mEq/L | 31.8 mmol/L | 1.24 | 3.57 | ADVIA Chemistry 1800 CO2_L y = 0.97x - 0.2 mEq/L r=0.990 n=59 | ADVIA Chemistry 1800 CO2_L y = 0.97x- 0.2 mmol/L r=0.990 n=59 | ≤ 10% at 600 mg/dL of hemoglobin | ≤ 10% at 30 mg/dL of bilirubin | ≤ 10% at 750 mg/dL of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | Corning 965 method / NIST SRM 192 | 11110139_EN Rev. 02, 2019-04 | IFU Document | 0.554 | 2.875 | 3.5 | 2.75 | 0.252 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097632 |  | ge16 | 00630414596181 | Global | General Chemistry | No | Atellica CH LDL Cholesterol Direct (DLDL) | DLDLDLDL | 0335 | 1600 | 4 x 400 | N/A | N/A | 10 | 15 | ; N‑ethyl‑N‑(2‑hydroxy‑3‑sulfopropyl)‑3‑methylaniline (TOOS); cholesterol esterase (microbial); cholesterol oxidase (microbial): catalase (microbial); ; 4-aminoantipyrine; peroxidase (horseradish) | ;#Serum;#Lithium Heparin Plasma;# | Up to 3 days at 2–8°C; Frozen at-20°C for up to several weeks; Frozen at -70°C or lower for longer time | 5 μL | No |  | Elimination/catalase | 5.0 – 1000.0 mg/dL | 0.13 – 25.90 mmol/L | N/A | N/A | Adults Optimal; <100.0 mg/dL Near optimal: 100.0–129.0 mg/dL Borderline high: 130.0–159.0 mg/dL High: 160.0–189.0 mg/dL Very high: ≥ 190.0 mg/dL | Adults Optimal: <2.60 mmol/L Near optimall: 2.60–3.30 mmol/L Borderline high: 3.40–4.10 mmol/L High: 4.10–4.90 mmol/L Very high: ≥ 4.90 mmol/L | N/A | N/A | N/A | Atellica CH HDL/LDL CAL | 11099402 | 3 days | Lyophilized | 1.0 mL of distilled water; Let stand 30 minutes | 3 | 2 | Human serum | 121 | 15 |  | No |  | No | N/A | 11099300 |  | N/A | N/A | No | N/A | N/A | 3.0 mg/dL | 0.08 mmol/L | N/A | N/A | 3.0 mg/dL | 0.8 mg/dL | N/A | 0.02 mmol/L | 0.08 mmol/L | N/A | 104.5 mg/dL | 2.71 mmol/L | 1.3 | 2.3 | 183.0 mg/dL | 4.74 mmol/L | 0.7 | 1.6 | 953.8 mg/dL | 24.7 mmol/L | 0.4 | 1 | ADVIA Chemistry 1800 DLDL y = 0.99x - 0.8 mg/dL r=1.000 n=143 | ADVIA Chemistry 1800 DLDL y = 0.99x - 0.02 mmol/L r=1.000 n=143 | ≤ 10% at 810 mg/dL of hemoglobin | ≤ 10% at 20 mg/dL of conjugated bilirubin ≤ 10% at 30 mg/dL of unconjugated bilirubin | ≤ 10% at 505 mg/dL of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | CDC Reference Method via patient sample correlation | 11110195_EN Rev. 03, 2019-11 | IFU Document | 0.906 | 5.625 | 3.5 | 2.75 | 0.412 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097533 |  | ge27 | 00630414595795 | Global | General Chemistry | No | Atellica CH Enzymatic Creatinine_2 (ECre_2) | ECre_2 (serum/plasma)ECre_2 (serum/plasma) | 0340 | 1400 | 4 x 350 | N/A | N/A | 10 | 30 | Creatinase; sarcosine oxidase; N-(3 sulfopropyl)-3 methoxy-5 methylaniline; ; 4-aminoantipyrine | ;#Serum;#Lithium Heparin Plasma;#Potassium EDTA Plasma;#Urine;# | Up to 2 days at 2–8°C; Frozen at -20° or below | 13 μL | No |  | Enzymatic | 0.1 - 30.00 mg/dL | 9 - 2652 µmol/L | 0.10 - 150 mg/dL | 9 - 13,260 μmol/L | Adult males: Serum/Plasma 0.6–1.1 mg/dL Adult females: Serum/Plasma 0.5–0.8 mg/dL | Adult males Serum/Plasma: 53–97 μmol/L Adult females Serum/Plasma: 44–71 μmol/L | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 180 | 30 |  | No |  | Yes | N/A | 11099300 |  | 0.10 - 150 mg/dL | 9 - 13,260 μmol/L | No | N/A | N/A | 0.03 mg/dL | 3 μmol/L | N/A | N/A | 0.03 mg/dL | 0.02 mg/dL | N/A | 2 μmol/L | 3 μmol/L | N/A | 0.70 mg/dL | 62 µmol/L | 2.5 | 2.7 | 6.90 mg/dL | 610 µmol/L | 0.5 | 2.1 | 17.74 mg/dL | 1568 µmol/L | 0.3 | 2.2 | ADVIA Chemistry 1800 y = 1.00x + 0.04 mg/dL r=1.000 n=103 | ADVIA Chemistry 1800 y = 1.00x + 4 µmol/L r=1.000 n=103 | < 10% at 750 mg/dL of hemoglobin | < 10% at 30 mg/dL of bilirubin | < 10% at 1000 mg/dL of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | NIST SRM 967 | 11110151_EN Rev. 05, 2020-03 | IFU Document | 0.856 | 5.625 | 3.5 | 2.75 | 0.389 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097533 |  | ge27 | 00630414595795 | Global | General Chemistry | No | Atellica CH Enzymatic Creatinine_2 (ECre_2) | ECre_2 (urine)ECre_2 (urine) | 0340 | 1400 | 4 x 350 | N/A | N/A | 10 | 30 | Creatinase ; sarcosine oxidase; N-(3 sulfopropyl)-3 methoxy-5 methylaniline; 4-aminoantipyrinE; peroxidase | ;#Urine;# | Up to 4 days at 2–8°C; frozen at -20°C or below | 13 μL | No |  | Enzymatic | 1.0 - 245.00 mg/dL | Urine 88 - 21658 µmol/L | 1.00 - 1225 mg/dL | 88 - 108,290 μmol/L | Adult males Urine 800–2000 mg/day Adult females Urine 600–1800 mg/day | Adult males Urine 800–2000 mg/day Adult females Urine 600–1800 mg/day | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 180 | 30 |  | No |  | Yes | N/A | 11099300 |  | 1.00 - 1225 mg/dL | 88 - 108,290 μmol/L | No | N/A | N/A | 0.22 mg/dL | 19 μmol/L | NA | NA | 0.22 mg/dL | 0.02 mg/dL | N/A | 2 μmol/L | 19 μmol/L | N/A | 57.02 mg/dL | 5041 µmol/L | 0.6 | 2.5 | 123.18 mg/dL | 10,889 µmol/L | 0.3 | 2.3 | 199.78 mg/dL | 17,661 µmol/L | 0.7 | 3.2 | ADVIA Chemistry 1800 y = 0.97x + 0.00 mg/dL r=1.000 n=109 | ADVIA Chemistry 1800 y = 0.97x + 0 µmol/L r=1.000 n=109 | N/A | N/A | N/A | No Biotin in assay formulation | N/A | N/A | N/A | N/A | NIST SRM 967 | 11110151_EN Rev. 05, 2020-03 | IFU Document | 0.856 | 5.625 | 3.5 | 2.75 | 0.389 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097597 |  | ge28 | 00630414596440 | Global | General Chemistry | No | Atellica CH Gamma-Glutamyl Transferase (GGT) | GGTGGT | 0701 | 1792 | 4 x 448 | N/A | N/A | 10 | 22 | synthetic | ;#Serum;#Lithium Heparin Plasma;# | Up to 7 days at 25°C or at 2–8°C; Frozen at -20°C for up to 6 months | 10 μL | No |  | Modified IFCC | 7 - 1200 U/L | 7 - 1200 U/L | 7 - 3600 U/L | 7 - 3600 U/L | Adult males Serum/plasma: < 73 U/L Adult females Serum/plasma; < 38 U/L | Adult males Serum/plasma; 73 U/L Adult females Serum/plasma; 38 U/L | N/A | N/A | N/A | Atellica CH ENZ 1 CAL | 11099317 | 30 days | Liquid | No preparation required | 3 | 2 | Bovine serum; GGT (bovine kidney) | 60 | 22 |  | No |  | Yes | N/A | 11099300 |  | 7 - 3600 U/L | 7 - 3600 U/L | No | N/A | N/A | 2 U/L | 2 U/L | N/A | N/A | 2 U/L | 1 U/L. | N/A | 1 U/L. | 2 U/L | N/A | 42 U/L | 42 U/L | 1.9 | 2.6 | 81 U/L | 81 U/L | 1 | 1.5 | 1044 U/L | 1044 U/L | 0.2 | 0.9 | ADVIA Chemistry 1800 GGT y = 1.01x - 3 U/L r=0.999 n=100 | ADVIA Chemistry 1800 GGT y = 1.01x - 3 U/L r=0.999 n=100 | ≤ 10% at 1000 mg/dL of hemoglobin | ≤ 10% at 10 mg/dL of bilirubin | ≤ 10% at 500 mg/dL of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | IFCC-452 reference material | 11110160_EN Rev. 02, 2019-07 | IFU Document | 1.217 | 5.625 | 3.5 | 2.75 | 0.553 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097611 |  | ge32 | 00630414596303 | Global | General Chemistry | No | Atellica CH Inorganic Phosphorus (IP) | IPIP (serum/plasma) | 1605 | 5100 | 3 x 1700 | N/A | N/A | 10 | 30 | Ammonium molybdate; Sulfuric acid | ;#Serum;#Lithium Heparin Plasma;#Urine;# | Up to several days at 2-8 °C; Frozen at -70 °C for up to several months | 8 μL | No |  | Phosphomolybdate | 0.3 -20 mg/dL | 0.10 - 6.46 mmol/L | 0.3 - 40.0 mg/dL | 0.1 - 12.92 mmol/L | Adult 2.4-5.1 mg/dL | Adult 0.78-1.65 mmol/L | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 days | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 180 | 7 |  | No |  | Yes | N/A | 11099300 |  | 0.3 - 40.0 mg/dL | 0.1 - 12.92 mmol/L | No | N/A | N/A | 0.1 mg/dL | 0.03 mmol/L | N/A | N/A | 0.1 mg/dL | 0.00 mg/dL | N/A | 0.00 mmol/L | 0.03 mmol/L | N/A | 3.0 mg/dL | 0.97 mmol/L | 2 | 2.7 | 4.6 mg/dL | 1.49 mmol/L | 1.1 | 4.6 | 11.8 mg/dL | 3.81 mmol/L | 0.7 | 0.9 | ADVIA Chemistry 1800 IP y = 1.03x - 0.2 mg/dL r=1.000 n= 106 | ADVIA Chemistry 1800 IP y = 1.03x - 0.06 mmol/L r=1.000 n= 106 | ≤ 10% at 500 mg/dL of hemoglobin | ≤ 10% at 30 mg/dL of bilirubin | ≤ 10% at 650 mg/dL of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | NIST SRM 200 | 11110174_EN Rev. 04, 2020-07 | IFU Document | 0.785 | 4.187 | 3.5 | 2.75 | 0.357 | 10.636 | 8.89 | 6.985
Atellica® CH | 11097611 |  | ge32 | 00630414596303 | Global | General Chemistry | No | Atellica CH Inorganic Phosphorus (IP) | IPIP (urine) | 1605 | 5100 | 3 x 1700 | N/A | N/A | 10 | 30 | Ammonium molybdate; Sulfuric acid | ;#Serum;#Lithium Heparin Plasma;#Urine;# | Up to 6 months at 2-8 °C for refrigerated acidifed urine specimens | 8 μL | No |  | Phosphomolybdate | 4 -100 mg/dL | 1.29 - 32.30 mmol/L | 4.0 - 260 mg/dL | 1.29 - 83.98 mmol/L | 0.4-1.3 g/day for adults on an unrestricted diet | 12.9-42.0 mmol/day for adults on an unrestricted diet | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 days | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 180 | 7 |  | No |  | Yes | N/A | 11099300 |  | 4.0 - 260 mg/dL | 1.29 - 83.98 mmol/L | No | N/A | N/A | 1.4 mg/dL | 0.45 mmol/L | N/A | N/A | 1.4 mg/dL | 0.5 mg/dL | N/A | 0.16 mmol/L | 0.45 mmol/L | N/A | 17.3 mg/dL | 5.59 mmol/L | 1.5 | 4 | 51.8 mg/dL | 16.73 mmol/L | 1.1 | 2.1 | 84.1 mg/dL | 27.16 mmol/L | 0.5 | 2.3 | ADVIA Chemistry 1800 IP y = 1.01x + 0.7 mg/dL r=0.999 n=104 | ADVIA Chemistry 1800 IP y = 1.01x + 0.23 mmol/L r=0.999 n=104 | N/A | N/A | N/A | No Biotin in assay formulation | N/A | N/A | N/A | N/A | NIST SRM 200 | 11110174_EN Rev. 04, 2020-07 | IFU Document | 0.785 | 4.187 | 3.5 | 2.75 | 0.357 | 10.636 | 8.89 | 6.985
Atellica® CH | 11097594 |  | ge08 | 00630414596464 | Global | General Chemistry | No | Atellica CH Lactate Dehydrogenase L-P (LDLP) | LDLPLDLP | 1202 | 1792 | 4 x 448 | N/A | N/A | 9 | 28 | L-lactic acid; NAD; 2‑chloracetamide | ;#Serum;#Lithium Heparin Plasma;# | 7 days at 20–25°C; 4 days at 4–8°C; 42 days at -20°C | 14 μL | No |  | Lactate / NAD | 14-750 U/L | 14-750 U/L | 14 - 4500 U/L | 14 - 4500 U/L | 120 - 246 U/L | 120 - 246 U/L | N/A | N/A | N/A | Atellica CH ENZ 1 CAL | 11099317 | 30 days | Liquid | No preparation required | 3 | 2 | Bovine serum albumin; LDHL (chicken heart) | 60 | 28 |  | No |  | Yes | N/A | 11099300 |  | 14 - 4500 U/L | 14 - 4500 U/L | No | N/A | N/A | 3 U/L | 3 U/L | N/A | N/A | 3 U/L | 2 U/L | 11 U/L | 2 U/L | 3 U/L | 11 U/L | 189 U/L | 189 U/L | 0.9 | 1 | 403 U/L | 403 U/L | 0.7 | 1 | 557 U/L | 557 U/L | 0.3 | 0.7 | ADVIA Chemistry 1800 y = 0.96x + 6 U/L r=0.999 n=100 | ADVIA Chemistry 1800 y = 0.96x + 6 U/L r=0.999 n=100 | N/A | ≤ 10% at 30 mg/dL of bilirubin | ≤ 10% at 650 mg/dL of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | IFCC-453 reference material | 11110157_EN Rev. 04, 2019-11 | IFU Document | 0.988 | 5.625 | 3.5 | 2.75 | 0.449 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097524 |  | ge44 | 00630414595870 | Global | General Chemistry | No | Atellica CH Total Protein_2 (Urine) (UPro) | UProUpro | 2038 | 1600 | 4 x 400 | N/A | N/A | 10 | 30 | Pyrogallol red; sodium molybdate; | ;#Urine;#Cerebrospinal fluid (CSF);# | Up to 3 days at 4 °C; frozen at -20 °C for up to 6 months | 13.3 μL | No |  | Dye binding/ Pyrogallol red | 1.0 - 125.0 mg/dL | 10 - 1250 mg/L | 1.0 - 500.0 mg/dL | 10 - 5000 mg/L | Urine: 24-hour: 1-14 mg/dL; At rest: 50-80 mg/day; After intense exercise 250 mg/day CSF: 8-32 mg/dL | Urine: 24-hour: 10-140 mg/L; At rest: 50-80 mg/day; After intense exercise 250 mg/day CSF: 80-320 mg/L | N/A | N/A | N/A | Atellica CH UPro CAL | 11099441 | 21 days | Liquid | No preparation required | 3 | 2 | Human serum albumin | 60 | 30 |  | No |  | Yes | N/A | 11099300 |  | 1.0 - 500.0 mg/dL | 10 - 5000 mg/L | No | N/A | N/A | 0.9 mg/dL | 9 mg/L | N/A | N/A | 0.9 mg/dL | 0.3 mg/dL | N/A | 3 mg/L | 9 mg/L | N/A | 12.1 mg/dL | 121 mg/L | 4.9 | 5.3 | 54.9 mg/dL | 549 mg/L | 1.2 | 1.9 | 113.1 mg/dL | 1131 mg/L | 0.6 | 0.9 | ADVIA Chemistry 1800 Urine; y = 0.99x + 0.5 mg/dL r=0.994 n=100 CSF: y=1.00x + 0.2mg/dL r=0.999 n=101 | ADVIA Chemistry 1800 Urine; y = 0.99x + 5 mg/L r=0.994 n=100 CSF: y=1.00x + 2mg/L r=0.999 n=101 | 38% bias at 16 mg/dL of hemoglobin | N/A | N/A | No Biotin in assay formulation | N/A | N/A | N/A | N/A | Traceable to an internal standard | 11110142_EN Rev. 02, 2019-07 | IFU Document | 0.585 | 2.875 | 3.5 | 2.75 | 0.266 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097630 |  | ge12 | 00630414596204 | Global | General Chemistry | No | Atellica CH Direct HDL Cholesterol (D-HDL) | D-HDLD-HDL | 0313 | 1792 | 4 x 448 | N/A | N/A | 8 | 89 | N (2-hydroxy-3-sulfopropyl)-3,5-dimethoxyaniline;cholesterol esterase (microbial); cholesterol oxidase (microbial): catalase (microbial); ; 4-aminoantipyrine; peroxidase (plant) | ;#Serum;#Lithium Heparin Plasma;# | 7 days at 2–8°C; Frozen at -70°C or colder for up to 3 months | 5 μL | No |  | Elimination/catalase | 20.0 – 129.0 mg/dL | 0.52 - 3.34 mmol/L | 258.0 mg/dL | 6.68 mmol/L for | Adults Serum/plasma: Low (undesirable, high risk): 40.0 mg/dL High (desirable, low risk): ≥ 60.0 mg/dL | Adults Serum/plasma: Low (undesirable, high risk): 1.00 mmol/L High (desirable, low risk): ≥ 1.60 mmol/L | N/A | N/A | N/A | Atellica CH HDL/LDL CAL | 11099402 | 3 days | Lyophilized | 1.0 mL of distilled water; Let stand 30 minutes | 3 | 2 | Human serum | 60 | 30 |  | No |  | Yes | N/A | 11099300 |  | 20.0 - 258.0 mg/dL | 0.52 - 6.68 mmol/L for | No | N/A | N/A | 2.1 mg/dL | 0.05 mmol/L | N/A | N/A | 2.1 mg/dL | 0.5 mg/dL | 2.7 mg/dL | 0.01 mmol/L | 0.05 mmol/L | 0.07 mmol/L | 37.5 mg/dL | 0.97 mmol/L | 1.8 | 2 | 64.6 mg/dL | 1.67 mmol/L | 0.6 | 1.6 | 101.5 mg/dL | 2.63 mmol/L | 0.7 | 0.7 | ADVIA Chemistry 1800 D-HDL y = 0.97x + 1.4 mg/dL r=0.999 n= 135 | ADVIA Chemistry 1800 D-HDL y = 0.97x + 0.04 mmol/L r=0.999 n= 135 | ≤ 10% at 500 mg/dL of hemoglobin | ≤ 10% at 30 mg/dL of bilirubin | ≤ 10% at 1000 mg/dL of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | NCEP Designated Comparison Method via patient sample correlation. | 11110193_EN Rev. 04, 2020-08 | IFU Document | 0.964 | 5.625 | 3.5 | 2.75 | 0.438 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097522 |  | dr22 | 00630414595894 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Acetaminophen (Acet) | Acet | 0102 | 1200 | 4 x 300 |  |  | 10 | 14 |  | ;#Serum;#Lithium Heparin Plasma;# | 8 hours at 20–25°C; frozen at -20°C for up to 45 days | 6 μL |  |  | Colorimetric: acyl amidohydrolase | 0.2 - 20.0 mg/dL | 13.2 - 1322.0 μmol/L | 0.2 - 60.0 mg/dL | 13.2 - 3966.0 μmol/L | 1.0–2.0 mg/dL | 66.0 - 132.0 µmol/L |  |  |  | Atellica CH TOX CAL | 11099440 | 3 days | Liquid |  | 3 | 2 |  | 62 | 1 |  | No |  | Yes | [5 empty] | No | [6 empty] | ≤0.2 mg/dL | 0.1 mg/dL | n/a | 6.6 μmol/L | ≤13.2 μmol/L | n/a | 4.6 mg/dL | 304.1 µmol/L | 1.2 | 3.7 | 9.0 mg/dL | 594.9 µmol/L | 0.6 | 1.5 | 16.7 mg/dL | 1103.9 µmol/L | 1.3 | 1.9 | ADVIA Chemistry 1800 y = 0.97x - 0.1 mg/dL (y = 0.97x - 6.6 μmol/L) r = 0.998 |  | ≤ 6% change up to 500 mg/dL of hemoglobin | ≤ 12% change up to 10 mg/dL of bilirubin | ≤ -14% change up to 500 mg/dL of Intralipid | No Biotin in assay formulation |  |  | n/a | n/a | Traceable to an internal standard manufactured using highly purified material. | 11110140_EN Rev. 04, 2019-04 | IFU Document | 0.442 | 5.625 | 3.5 | 2.75 | 0.201 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097506 |  | dr01 | 00630414596068 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Amphetamines (Amp) | Amp300 | 0140 | 1520 | 4 x 380 |  |  | 7 | 30 |  | ;#Urine;# | 7 days at 25°C; 30 days at 2–8°C; 12 months at -20°C | 13.5 μL |  |  | EMIT | 300 ng/mL;Semiquantitative 100 - 1000 ng/mL | n/a | n/a | n/a | qualitative/semi-quantitative assay | n/a |  |  |  | Emit Calibrator 0, 1, 2, 3 | 10445406 10445407 10445408 10445409 | 5 weeks | Liquid |  | 3 | 5 |  | 60 | 20 |  | No |  | No | [12 empty] | 53 ng/mL | 41 ng/mL | n/a | n/a | n/a | n/a | 307.8 mAU/min | N/A | 0.2 | 0.4 | 326.2 mAU/min |  | 0.2 | 0.5 | 346.0 mAU/min |  | 0.2 | 0.5 | Viva-E Analyzer ± 50% of 300 ng/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | Traceable to Syva Emit Cal/Cont which are gravimetrically prepared standards- qualified by GC/MS within ± 10% of nominal. | 11110124_EN Rev. 03, 2019-07 | IFU Document | 0.915 | 5.625 | 3.5 | 2.75 | 0.416 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097506 |  | dr01 | 00630414596068 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Amphetamines (Amp) | Amp500 | 0137 | 1520 | 4 x 380 |  |  | 7 | 30 |  | ;#Urine;# | 7 days at 25°C; 30 days at 2–8°C; 12 months at -20°C | 8 μL |  |  | EMIT | 500 ng/mL;Semiquantitative 150 - 1800 ng/mL | n/a | n/a | n/a | qualitative/semi-quantitative assay | n/a |  |  |  | Emit Calibrator 0, 1, 2, 3, 5 | 10445406 10445407 10445408 10445409 10445411 | 5 weeks | Liquid |  | 3 | 5 |  | 60 | 20 |  | No |  | No | [12 empty] | 63 ng/mL | 44 ng/mL | n/a | n/a | n/a | n/a | 323.7 mAU/min | N/A | 0.2 | 0.4 | 345.3 mAU/min |  | 0.2 | 0.4 | 365,9 |  | 0.3 | 0.5 | Viva-E Analyzer ± 50% of 500 ng/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | Traceable to Syva Emit Cal/Cont which are gravimetrically prepared standards- qualified by GC/MS within ± 10% of nominal. | 11110124_EN Rev. 03, 2019-07 | IFU Document | 0.915 | 5.625 | 3.5 | 2.75 | 0.416 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097506 |  | dr01 | 00630414596068 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Amphetamines (Amp) | Amp1K | 0138 | 1520 | 4 x 380 |  |  | 7 | 30 |  | ;#Urine;# | 7 days at 25°C; 30 days at 2–8°C; 12 months at -20°C | 5 μL |  |  | EMIT | 1000 ng/mL;Semiquantitative 150 - 1800 ng/mL | n/a | n/a | n/a | qualitative/semi-quantitative assay | n/a |  |  |  | Emit Calibrator 0, 1, 2, 3, 5 | 10445406 10445407 10445408 10445409 10445411 | 5 weeks | Liquid |  | 3 | 5 |  | 60 | 20 |  | No |  | No | [12 empty] | 121 ng/mL | 83 ng/mL | n/a | n/a | n/a | n/a | 345.1 mAU/min | N/A | 0.2 | 0.6 | 373.8 mAU/min |  | 0.3 | 0.5 | 401.1 mAU/min |  | 0.2 | 0.5 | Viva-E Analyzer ± 50% of 1000 ng/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | Traceable to Syva Emit Cal/Cont which are gravimetrically prepared standards- qualified by GC/MS within ± 10% of nominal. | 11110124_EN Rev. 03, 2019-07 | IFU Document | 0.915 | 5.625 | 3.5 | 2.75 | 0.416 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097505 |  | dr03 | 00630414596075 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Benzodiazepines (Bnz) | Bnz200 | 0202 | 1520 | 4 x 380 |  |  | 7 | 30 |  | ;#Urine;# | 7 days at 25°C; 2–8°C or frozen at -20°C for up to 30 days | 17.5 μL |  |  | EMIT | 200 ng/mL;Semi-quant 50-900 ng/mL | n/a | n/a | n/a | qualitative/semi-quantitative assay | n/a |  |  |  | Emit Calibrator/Control Levels 0, 2, 3, 4, 5 | 10445406 10445408 10445409 10445410 10445411 | 5 weeks | Liquid |  | 3 | 5 |  | 60 | 20 |  | No |  | No | [12 empty] | 27 ng/mL | 15 ng/mL | n/a | n/a | n/a | n/a | 187.3 mAU/min | N/A | 0.3 | 0.7 | 236.7 mAU/min |  | 0.6 | 0.8 | 256.4 mAU/min |  | 0.3 | 0.6 | Viva-E Analyzer ± 50% of 200 ng/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | Traceable to Syva Emit Cal/Cont which are gravimetrically prepared standards- qualified by GC/MS within ± 10% of nominal. | 11110123_EN Rev. 04, 2019-11 | IFU Document | 0.913 | 5.625 | 3.5 | 2.75 | 0.415 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097505 |  | dr03 | 00630414596075 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Benzodiazepines (Bnz) | Bnz300 | 0202 | 1520 | 4 x 380 |  |  | 7 | 30 |  | ;#Urine;# | 7 days at 25°C; 2–8°C or frozen at -20°C for up to 30 days | 17.5 μL |  |  | EMIT | 300 ng/mLSemi-quant 50-900 ng/mL | n/a | n/a | n/a | qualitative/semi-quantitative assay | n/a |  |  |  | Emit Calibrator/Control Levels 0, 2, 3, 4, 5 | 10445406 10445408 10445409 10445410 10445411 | 5 weeks | Liquid |  | 3 | 5 |  | 60 | 20 |  | No |  | No | [12 empty] | 27 ng/mL | 15 ng/mL | n/a | n/a | n/a | n/a | 187.3 mAU/min | N/A | 0.3 | 0.7 | 249.8 mAU/min |  | 0.2 | 0.6 | 256.4 mAU/min |  | 0.3 | 0.6 | Viva-E Analyzer ± 50% of 300 ng/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | Traceable to Syva Emit Cal/Cont which are gravimetrically prepared standards- qualified by GC/MS within ± 10% of nominal. | 11110123_EN Rev. 04, 2019-11 | IFU Document | 0.913 | 5.625 | 3.5 | 2.75 | 0.415 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097507 |  | dr02 | 00630414596051 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Barbiturates (Brb) | Brb200 | 0203 | 1520 | 4 x 380 |  |  | 7 | 30 |  | ;#Urine;# | 7 days at 25°C; 2–8°C or frozen at -20°C for up to 30 days | 17.5 μL |  |  | EMIT | 200 ng/mL;Semi-quant 40 – 750 ng/mL | n/a | n/a | n/a | qualitative/semi-quantitative assay | n/a |  |  |  | Emit Calibrator/Control Levels 0, 2, 3, 4, 5 | 10445406 10445408 10445409 10445410 10445411 | 5 weeks | Liquid |  | 3 | 5 |  | 60 | 30 |  | No |  | No | [12 empty] | 9 ng/mL | 5 ng/mL | n/a | n/a | n/a | n/a | 123.2 mAU/min | N/A | 0.3 | 0.6 | 171.9 mAU/min |  | 0.3 | 0.5 | 214.6 mAU/min |  | 0.3 | 0.5 | Viva-E Analyzer ± 50% of 200 ng/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | Traceable to Syva Emit Cal/Cont which are gravimetrically prepared standards- qualified by GC/MS within ± 10% of nominal. | 11110125_EN Rev. 03, 2019-07 | IFU Document | 0.917 | 5.625 | 3.5 | 2.75 | 0.417 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097507 |  | dr02 | 00630414596051 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Barbiturates (Brb) | Brb300 | 0203 | 1520 | 4 x 380 |  |  | 7 | 30 |  | ;#Urine;# | 7 days at 25°C; 2–8°C or frozen at -20°C for up to 30 days | 17.5 μL |  |  | EMIT | 300 ng/mL;Semi-quant 40 – 750 ng/mL | n/a | n/a | n/a | qualitative/semi-quantitative assay | n/a |  |  |  | Emit Calibrator/Control Levels 0, 2, 3, 4, 5 | 10445406 10445408 10445409 10445410 10445411 | 5 weeks | Liquid |  | 3 | 5 |  | 60 | 30 |  | No |  | No | [12 empty] | 9 ng/mL | 5 ng/mL | n/a | n/a | n/a | n/a | 123.2 mAU/min | N/A | 0.3 | 0.6 | 197.9 mAU/min |  | 0.4 | 0.6 | 214.6 mAU/min |  | 0.3 | 0.5 | Viva-E Analyzer ± 50% of 300 ng/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | Traceable to Syva Emit Cal/Cont which are gravimetrically prepared standards- qualified by GC/MS within ± 10% of nominal. | 11110125_EN Rev. 03, 2019-07 | IFU Document | 0.917 | 5.625 | 3.5 | 2.75 | 0.417 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097515 |  | th02 | 00630414595986 | Global | Therapeutic Drug Monitoring (TDM) | No | Atellica CH Carbamazepine (Carb) | Carb | 0309 | 400 | 4 x 100 | Atellica CH DRUG CAL II |  | 8 | 30 |  | ;#Serum;#Lithium Heparin Plasma;# | 8 hours at 25°C; 2 days at 2-8°C; frozen at -20°C for up to 30 days | 3.8 μL | No |  | PETINIA | 0.4 - 20.0 µg/mL | 1.7 - 84.7 µmol/L | 0.4 - 40.0 μg/mL | 1.7 - 169.3 μmol/L | 4.0–12.0 μg/mL | 16.9–50.8 μmol/L |  |  |  | Atellica CH DRUG CAL II | 11099405 | 30 days | Liquid |  | 3 | 5 |  | 30 | 7 |  | No |  | Yes | [5 empty] | Not Specified | [6 empty] | 0.4 μg/mL | 0.3 μg/mL | n/a | 1.3 μmol/L | 1.7 μmol/L | n/a | 3.2 μg/mL | 13.4 μmol/L | 1.5 | 2.4 | 6.4 μg/mL | 27.2 μmol/L | 1 | 1.3 | 16.5 μg/mL | 69.8 μmol/L | 0.8 | 1.2 | Dimension y = 0.98x + 0.01 μg/mL (y = 0.98x + 0.04 μmol/L) r = 0.998 |  | ≤ 2% change up to 500 mg/dL of hemoglobin | 0% change up to 20 mg/dL of bilirubin | ≤ 8% change up to 800 mg/dL of Intralipid | No Biotin in assay formulation |  |  | N/A | N/A | USP Carbamazepine Standard | 11110133_EN Rev. 01, 2017-05 | IFU Document | 0.629 | 5.625 | 3.5 | 2.75 | 0.286 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097504 |  | dr04 | 00630414596082 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Cocaine Metabolite (Coc) | Coc150 | 0336 | 1520 | 4 x 380 |  |  | 7 | 30 |  | ;#Urine;# | 7 days at 25°C; 30 days at 2–8°C or frozen at -20°C. | 25 μL |  |  | EMIT | 150 ng/mL; Semi-quant 36-900 ng/mL | n/a | N/A | N/A | qualitative/semi-quantitative assay | N/A |  |  |  | Emit Calibrator 0, 2, 3, 4, 5 | 10445406 10445408 10445409 10445410 10445411 | 5 weeks | Liquid |  | 3 | 5 |  | 40 | 20 |  | No |  | No | [12 empty] | 7 ng/mL | 4 ng/mL | n/a | N/A | N/A | n/a | 129.6 mAU/min | N/A | 0.3 | 0.6 | 163.9 mAU/min |  | 0.3 | 0.6 | 183.0 mAU/min |  | 0.2 | 0.6 | Viva-E Analyzer ± 50% of 150 ng/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | Traceable to Syva Emit Cal/Cont which are gravimetrically prepared standards- qualified by GC/MS within ± 10% of nominal. | 11110122_EN Rev. 03, 2019-10 | IFU Document | 0.884 | 5.625 | 3.5 | 2.75 | 0.402 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097504 |  | dr04 | 00630414596082 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Cocaine Metabolite (Coc) | Coc300 | 0337 | 1520 | 4 x 380 |  |  | 7 | 30 |  | ;#Urine;# | 7 days at 25°C; 30 days at 2–8°C or frozen at -20°C. | 25 μL |  |  | EMIT | 300 ng/mL; Semi-quant 36-900 ng/mL | n/a | N/A | N/A | qualitative/semi-quantitative assay | N/A |  |  |  | Emit Calibrator 0, 2, 3, 4, 5 | 10445406 10445408 10445409 10445410 10445411 | 5 weeks | Liquid |  | 3 | 5 |  | 40 | 20 |  | No |  | No | [12 empty] | 7 ng/mL | 4 ng/mL | n/a | N/A | N/A | n/a | 129.6 mAU/min | N/A | 0.3 | 0.6 | 163.9 mAU/min |  | 0.3 | 0.6 | 183.0 mAU/min |  | 0.2 | 0.6 | Viva-E Analyzer ± 50% of 300 ng/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | Traceable to Syva Emit Cal/Cont which are gravimetrically prepared standards- qualified by GC/MS within ± 10% of nominal. | 11110122_EN Rev. 03, 2019-10 | IFU Document | 0.884 | 5.625 | 3.5 | 2.75 | 0.402 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097526 |  | th03 | 00630414595863 | Global | Therapeutic Drug Monitoring (TDM) | No | Atellica CH Digoxin (Dgn) | Dgn | 0405 | 1600 | 4 x 400 |  |  | 10 | 30 |  | ;#Serum;#Lithium Heparin Plasma;# | 8 hours at 20–25°C; up to 7 days at 2–8°C; frozen at -20°C for up to 6 months | 15.6 μL |  |  | latex-enhanced immunoturbidimetric | 0.14 - 5.00 ng/mL | 0.18 - 6.40 nmol/L | 0.14 - 10.00 ng/mL | 0.18 - 12.80 nmol/L | 0.80–2.00 ng/mL | 1.02–2.56 nmol/L |  |  |  | Atellica CH TDM CAL | 11099439 | 7 days when protected fromlight | Lyophilized |  | 3 | 6 |  | 60 | 7 |  | No |  | Yes | [12 empty] | 0.14 ng/mL | 0.08 ng/mL | n/a | 0.10 nmol/L | 0.18 nmol/L | n/a | 1.51 ng/mL | 1.93 nmol/L | 2.4 | 2.9 | 2.68 ng/mL | 3.43 nmol/L | 3.2 | 3.5 | 4.20 ng/mL | 5.38 nmol/L | 1.8 | 2 | ADVIA Chemistry 1800 y = 0.97x + 0.18 ng/mL (y = 0.97x + 0.23 nmol/L) r = 0.994 |  | ≤ 8% change up to 500 mg/dL of hemoglobin | ≤ 7% change up to 60 mg/dL of bilirubin | ≤ 9% change up to 399 mg/dL of Intralipid | No Biotin in assay formulation |  |  | N/A | N/A | Traceable to an internal standard manufactured using highly purified material traceable to USP. | 11110144_EN Rev. 01, 2017-05 | IFU Document | 0.845 | 5.625 | 3.5 | 2.75 | 0.384 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097501 |  | dr14 | 00630414596112 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Ethyl Alcohol (ETOH) | ETOH | 0141 | 1200 | 4 x 300 |  |  | 8 | 30 |  | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#Potassium EDTA Plasma;#Sodium fluoride/potassium oxalate Plasma;#Urine;# | 2 days at room temperature; 2 weeks at 2–8°C; or indefinitely frozen at ≤ -20°C | 8 μL |  |  | Enzymatic reaction | 3.0 - 300.0 mg/dL | 0.7 - 65.1 mmol/L | 3.0 - 900.0 mg/dL | 0.7 - 195.3 mmol/L | Varied. Reportedly fatal 400.0 mg/mL | Varied. Reportedly fatal 86.8 mmol/L |  |  |  | Atellica CH TOX CAL | 11099440 | 3 days | Liquid |  | 3 | 2 | Human serum albumin | 60 | 10 |  | No |  | Yes | [12 empty] | 2.8 mg/dL | 1.3 mg/dL | n/a | 0.3 mmol/L | 0.6 mmol/L | n/a | 50.7 mg/dL | 11.0 mmol/L | 1.5 | 5.4 | 106.6 mg/dL | 23.1 mmol/L | 0.9 | 2.9 | 266.7 mg/dL | 57.9 mmol/L | 0.5 | 1.7 | Dimension RxL y = 1.05x + 1.5 mg/dL (y = 1.05x + 0.3 mmol/L) r = 0.995 |  | ≤ 6% change up to 1000 mg/dL of hemoglobin | ≤ 4% change up to 80 mg/dL of bilirubin | ≤ 1% change up to 3000 mg/dL of Intralipid | No Biotin in assay formulation |  |  | N/A | N/A | USP grade ethyl alcohol | 11110091_EN Rev. 03, 2019-07 | IFU Document | 0.779 | 5.625 | 3.5 | 2.75 | 0.354 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097516 |  | th05 | 00630414595979 | Global | Therapeutic Drug Monitoring (TDM) | No | Atellica CH Gentamicin (Gent) | Gent | 0702 | 400 | 4 x 100 | Atellica CH DRUG CAL II |  | 8 | 30 per pack |  | ;#Serum;#Lithium Heparin Plasma;# | 8 hours at room temperature; 2 days at 2–8°C; 30 days at -20°C | 3.8 μL |  |  | PETINIA | 0.5 - 12.0 µg/mL | 1.1 - 25.9 µmol/L | 0.5 - 24.0 μg/mL | 1.1 - 51.8 μmol/L | 4.0 - 10.0 μg/mL | 8.64 - 21.6 μmol/L |  |  |  | Atellica CH DRUG II CAL | 11099405 | 30 days | Liquid |  | 3 | 5 |  | 30 | 7 |  | No |  | Yes | [12 empty] | 0.4 μg/mL | 0.3 μg/mL | n/a | 0.6 μmol/L | 0.9 μmol/L | n/a | 1.5 μg/mL | 3.2 µmol/L | 5.2 | 6.5 | 4.2 μg/mL | 9.0 µmol/L | 1.7 | 2.4 | 7.8 μg/mL | 16.8 µmol/L | 1.7 | 1.8 | Dimensiony = 0.98x + 0.02 µg/mL (y = 0.98x + 0.04 µmol/L) r = 0.984 |  | ≤ 2% change up to 500 mg/dL of hemoglobin | ≤ 2% change up to 20 mg/dL of bilirubin | ≤ 6% change up to 1000 mg/dL of Intralipid | No Biotin in assay formulation |  |  | N/A | N/A | USP Gentamicin Standard | 11110134_EN Rev. 01, 2017-05 | IFU Document | 0.629 | 5.625 | 3.5 | 2.75 | 0.286 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097535 |  | th06 | 00630414006789 | Global | Therapeutic Drug Monitoring (TDM) | No | Atellica CH Lithium (Li) | Li | 1208 | 400 | 2 x 200 |  |  | 10 | 13 |  | ;#Serum;#Sodium Heparin Plasma;#Potassium EDTA Plasma;# | 24 hours at room temperature; 7 days at 4°C; 180 days at -18°C | 13 μL |  |  | Colorimetric endpoint chemistry | 0.10 - 3.00 mmol/L | n/a | 0.10 - 6.00 mmol/L | n/a | 1.00 - 1.20 mmol/L | n/a |  |  |  | Atellica CH CHEM CAL | 11099411 | 2 days | Lyophilized | Reconstituted to 3.0 mL/vial | 3 | 2 | bovine serum | 63 | 4 |  | No |  | Yes | [12 empty] | 0.07 mmol/L | 0.05 mmol/L | n/a | n/a | n/a | n/a | 0.91 mmol/L | N/A | 1.2 | 2.1 | 1.99 mmol/L |  | 0.6 | 1.3 | 2.76 mmol/L |  | 0.5 | 1 | ADVIA Chemistry 1800 y = 0.98x + 0.02 mmol/L r = 0.997 |  | ≤ 7% change up to 750 mg/dL of hemoglobin | ≤ -8% change up to 40 mg/dL of bilirubin | ≤ -4% change up to 500 mg/dL of Intralipid | No Biotin in assay formulation |  |  | n/a | n/a | NIST SRM 908b1 and 909b2 | 11110126_EN Rev. 02, 2018-05 | IFU Document | 0.455 | 2.875 | 3.5 | 2.75 | 0.207 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097520 |  | dr06 | 00630414595917 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Methadone (Mdn) | Mdn150 | 1303 | 1520 | 4 x 380 |  |  | 7 | 30 |  | ;#Urine;# | 7 days at 25°C; 30 days at 2–8°C or storedfrozen at -20°C | 23 μL |  |  | EMIT | 150 ng/mL; Semi-quant 40 - 450 ng/mL | n/a | n/a | n/a | qualitative/semi-quantitative assay | n/a |  |  |  | Emit Calibrator 0, 2, 3, 4 | 10445406 10445408 10445409 10445410 | 5 weeks | Liquid |  | 3 | 4 |  | 60 | 20 |  | No |  | No | [12 empty] | 15 ng/mL | 12 ng/mL | n/a | n/a | n/a | n/a | 106.0 mAU/min | N/A | 0.3 | 0.4 | 198.6 mAU/min | N/A | 0.3 | 0.4 | 221.6 mAU/min | N/A | 0.4 | 0.4 | Viva-E Analyzer ± 50% of 150 ng/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | Traceable to Syva Emit Cal/Cont which are gravimetrically prepared standards- qualified by GC/MS within ± 10% of nominal. | 11110138_EN Rev. 02, 2019-11 | IFU Document | 0.618 | 5.625 | 3.5 | 2.75 | 0.281 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097520 |  | dr06 | 00630414595917 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Methadone (Mdn) | Mdn300 | 1314 | 1520 | 4 x 380 |  |  | 7 | 30 |  | ;#Urine;# | 7 days at 25°C; 30 days at 2–8°C or storedfrozen at -20°C | 11.5 μL |  |  | EMIT | 300 ng/mL; Semi-quant 90 - 900 ng/mL | n/a | n/a | n/a | qualitative/semi-quantitative assay | n/a |  |  |  | Emit Calibrator 0, 3, 4, 5 | 10445406 10445409 10445410 10445411 | 5 weeks | Liquid |  | 3 | 4 |  | 60 | 20 |  | No |  | No | [12 empty] | 22 ng/mL | 16 ng/mL | n/a | n/a | n/a | n/a | 119.0 mAU/min | N/A | 0.3 | 0.4 | 225.7 mAU/min | N/A | 0.3 | 0.5 | 248.8 mAU/min | N/A | 0.3 | 0.5 | Viva-E Analyzer ± 50% of 150 ng/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | Traceable to Syva Emit Cal/Cont which are gravimetrically prepared standards- qualified by GC/MS within ± 10% of nominal. | 11110138_EN Rev. 02, 2019-11 | IFU Document | 0.618 | 5.625 | 3.5 | 2.75 | 0.281 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097527 |  | dr07 | 00630414595856 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Methadone Metabolite (EDDP) (MetMtb) | MetMtb | 1315 | 2000 | 4 x 500 |  |  | 7 | 30 |  | ;#Urine;# | 5 days at 2–8°C or ≤ -20°C | 25 μL |  |  | Enzyme Immunoassay | 1000 ng/mL | n/a | n/a | n/a | qualitative assay | n/a |  |  |  | Atellica CH MetMtb CAL | 11099437 | 30 days | Liquid |  | 3 | 1 |  | 60 | 14 |  | No |  | No | [12 empty] | n/a | n/a | n/a | n/a | n/a | n/a | 193.5 mAU/min | N/A | 0.3 | 1.3 | 221.0 mAU/min |  | 0.5 | 1.3 | n/a |  |  |  | ADVIA 1800 @ 1000 ng/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | Traceable to Syva Emit Cal/Cont which are gravimetrically prepared standards- qualified by GC/MS within ± 10% of nominal. | 11110145_EN Rev. 02, 2019-07 | IFU Document | 0.609 | 5.625 | 3.5 | 2.75 | 0.277 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097502 |  | dr08 | 00630414596105 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Opiates (Op) | Op300 | 1504 | 1520 | 4 x 380 |  |  | 7 | 30 |  | ;#Urine;# | 7 days at 25°C; 30 days at 2–8°C or storedfrozen at -20°C | 24 μL |  |  | EMIT | 300 ng/mL; Semi-quant 75-1800 ng/mL | n/a | n/a | n/a | qualitative/semi-quantitative assay | n/a |  |  |  | Emit Calibrator 0, 1, 2, 3 | 10445406 10445407 10445408 10445409 | 5 weeks | Liquid |  | 3 | 4 |  | 60 | 30 |  | No |  | No | [12 empty] | 27 ng/mL | 17 ng/mL | n/a | n/a | n/a | n/a | 105 mAU/min | N/A | 0.5 | 0.6 | 142 mAU/min |  | 0.5 | 0.6 | 154 mAU/min |  | 0.4 | 0.6 | Viva-E Analyzer ± 50% of 300 ng/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | N/A | N/A | Traceable to Syva Emit Cal/Cont which are gravimetrically prepared standards- qualified by GC/MS within ± 10% of nominal. | 11110106_EN Rev. 03, 2019-10 | IFU Document | 0.904 | 5.625 | 3.5 | 2.75 | 0.411 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097502 |  | dr08 | 00630414596105 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Opiates (Op) | Op2000 | 1505 | 1520 | 4 x 380 |  |  | 7 | 30 |  | ;#Urine;# | 7 days at 25°C; 30 days at 2–8°C or storedfrozen at -20°C | 6 μL |  |  | EMIT | 2000 ng/mL; Semi-quant 320-3800 ng/mL | n/a | n/a | n/a | qualitative/semi-quantitative assay | n/a |  |  |  | Emit Calibrator 0, 2, 3, 5 | 10445406 10445408 10445409 10445411 | 5 weeks | Liquid |  | 3 | 4 |  | 60 | 30 |  | No |  | No | [12 empty] | 179 ng/mL | 122 ng/mL | n/a | n/a | n/a | n/a | 127 mAU/min | N/A | 0.5 | 0.6 | 211 mAU/min |  | 0.5 | 0.7 | 222 mAU/min |  | 0.5 | 0.6 | Viva-E Analyzer ± 50% of 2000 ng/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | N/A | N/A | Traceable to Syva Emit Cal/Cont which are gravimetrically prepared standards- qualified by GC/MS within ± 10% of nominal. | 11110106_EN Rev. 03, 2019-10 | IFU Document | 0.904 | 5.625 | 3.5 | 2.75 | 0.411 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097509 |  | dr09 | 00630414596044 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Phencyclidine (Pcp) | Pcp25 | 1602 | 1520 | 4 x 380 |  |  | 7 | 30 |  | ;#Urine;# | 7 days unrefrigerated; 30 days at 2 – 8°C; 12 months at ≤ -20°C. | 23.5 μL |  |  | EMIT | 25 ng/mL; Semi-quant 5 - 75 ng/mL | n/a | n/a | n/a | qualitative/semi-quantitative assay | n/a |  |  |  | Emit Calibrator 0, 2, 3, 4 | 10445406 10445408 10445409 10445410 | 5 weeks | Liquid |  | 3 | 4 |  | 60 | 19 |  | No |  | No | [12 empty] | n/a | n/a | n/a | n/a | n/a | n/a | 18 ng/mL | N/A | 1.8 | 3.3 | 25 ng/mL | N/A | 1.6 | 3.4 | 30 ng/mL | N/A | 2.2 | 2.7 | GC/MS ± 50% of the 25 ng/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | Traceable to Syva Emit Cal/Cont which are gravimetrically prepared standards- qualified by GC/MS within ± 10% of nominal. | 11110127_EN Rev. 02, 2019-07 | IFU Document | 0.913 | 5.625 | 3.5 | 2.75 | 0.415 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097514 |  | th09 | 00630414595993 | Global | Therapeutic Drug Monitoring (TDM) | No | Atellica CH Phenobarbital (Phnb) | Phnb | 1603 | 400 | 4 x 100 |  |  | 8 | 30 per pack |  | ;#Serum;#Lithium Heparin Plasma;# | 8 hours at 25°C; 2 days at 2-8°C; 30 days at -20°C | 5 μL |  |  | PETINIA | 3.0 - 80.0 µg/mL | 12.9 - 344.8 µmol/L | 3.0 - 160.0 μg/mL | 12.9 - 689.6 μmol/L | 15.0-40.0 μg/mL | 64.7-172.4 μmol/L |  |  |  | Atellica CH DRUG CAL | 11099336 | 90 days | Liquid |  | 3 | 5 |  | 30 | 7 |  | No |  | Yes | [12 empty] | 2.0 μg/mL | 1.5 μg/mL | n/a | 6.5 μmol/L | 8.6 μmol/L | n/a | 14.1 μg/mL | 60.8 µmol/L | 1.4 | 2.8 | 47.0 μg/mL | 202.6 µmol/L | 1 | 1.6 | 67.6 μg/mL | 291.4 µmol/L | 0.9 | 2.3 | Dimensiony = 0.92x + 0.7 µg/mL (y = 0.92x + 3.0 µmol/L) r = 0.994 |  | ≤ 5% change up to 500 mg/dL of hemoglobin | ≤ 6% change up to 80 mg/dL of bilirubin | ≤ 10% change up to 200 mg/dL of Intralipid | No Biotin in assay formulation |  |  | n/a | n/a | USP Phenobarbital Standard | 11110132_EN Rev. 01, 2017-05 | IFU Document | 0.625 | 5.625 | 3.5 | 2.75 | 0.284 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097510 |  | th11 | 00630414596037 | Global | Therapeutic Drug Monitoring (TDM) | No | Atellica CH Phenytoin (Phny) | Phny | 1604 | 400 | 4 x 100 |  |  | 8 | 30 per pack |  | ;#Serum;#Lithium Heparin Plasma;# | 24 hours at ambient temperature; 48 hours at 2-8°C; 5 months at -20°C | 5 μL |  |  | PETINIA | 2.0 - 40.0 µg/mL | 7.9 - 158.4 µmol/L | 2.0 - 80.0 μg/mL | 7.9 - 316.8 μmol/L | 10.0-20.0 μg/mL | 39.6-79.2 μmol/L |  |  |  | Atellica CH DRUG CAL | 11099336 | 90 days | Liquid |  | 3 | 5 |  | 28 | 7 |  | No |  | Yes | [12 empty] | 0.8 μg/mL | 0.4 μg/mL | n/a | 1.6 μmol/L | 3.2 μmol/L | n/a | 8.2 μg/mL | 32.5 µmol/L | 2.2 | 3.1 | 16.0 μg/mL | 63.4 µmol/L | 1.3 | 1.8 | 35.6 μg/mL | 141.0 µmol/L | 1.1 | 1.9 | Dimension RxL y = 0.95x - 0.1 µg/mL (y = 0.95x - 0.4 µmol/L) r = 0.995 |  | ≤ 9% change up to 300 mg/dL of hemoglobin | ≤ 4% change up to 20 mg/dL of bilirubin | ≤ 7% change up to 250 mg/dL of Intralipid | No Biotin in assay formulation |  |  | n/a | n/a | USP Phenytoin Standard | 11110128_EN Rev. 01, 2017-05 | IFU Document | 0.623 | 5.625 | 3.5 | 2.75 | 0.283 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097519 |  | dr10 | 00630414595948 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Propoxyphene (Ppx) | Ppx | 2702 | 1520 | 4 x 380 |  |  | 7 | 30 |  | ;#Urine;# | 7 days at 25°C; 30 days at 2–8°C or storedfrozen at -20°C | 20 μL |  |  | EMIT | 300 ng/mL; Semi-quant 103 - 1000 ng/mL | n/a | n/a | n/a | qualitative/semi-quantitative assay | n/a |  |  |  | Emit Calibrator 0, 2, 3, 4, 5 | 10445406 10445408 10445409 10445410 10445411 | 5 weeks | Liquid |  | 3 | 5 |  | 60 | 30 |  | No |  | No | [12 empty] | 18 ng/mL | 6 ng/mL | n/a | n/a | n/a | n/a | 166 mAU/min | N/A | 0.7 | 0.9 | 215 mAU/min |  | 0.6 | 0.7 | 226 mAU/min |  | 0.5 | 0.5 | Viva-E Analyzer ± 50% of 300 ng/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | Traceable to Syva Emit Cal/Cont which are gravimetrically prepared standards- qualified by GC/MS within ± 10% of nominal. | 11110137_EN Rev. 02, 2019-10 | IFU Document | 0.911 | 5.625 | 3.5 | 2.75 | 0.414 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097523 |  | dr15 | 00630414595887 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Salicylate (Sal) | Sal | 1901 | 1200 | 4 x 300 |  |  | 10 | 30 |  | ;#Serum;#Lithium Heparin Plasma;# | 7 days at 25°C; 14 days at 2-8 °C; 6 months at ≤ -20 °C | 8 μL |  |  | Colorimetric | 3.0 - 100.0 mg/dL | 0.22 - 7.20 mmol/L | 3.0 - 400.0 mg/dL | 0.22 - 28.8 mmol/L | Toxic >30 mg/dL | n/a |  |  |  | Atellica CH TOX CAL | 11099440 | 3 days | Liquid |  | 3 | 2 |  | 180 | 21 |  | No |  | Yes | [12 empty] | 0.9 mg/dL | 0.5 mg/dL | n/a | 0.04 mmol/L | 0.06 mmol/L | n/a | 13.2 mg/dL | 0.95 mmol/L | 2.4 | 2.8 | 30.9 mg/dL | 2.2 mmol/L | 1.2 | 1.4 | 85.8 mg/dL | 6.2 mmol/L | 0.4 | 0.5 | ADVIA Chemistry 1800 Sal y = 1.01x - 0.9 mg/dL (y = 1.01x - 0.1 mmol/L) r = 0.997 r = 0.997 |  | ≤ -9% change up to 1000 mg/dL of hemoglobin | ≤ -9% change up to 30 mg/dL of bilirubin | ≤ -5% change up to 1000 mg/dL of Intralipid | No Biotin in assay formulation |  |  | n/a | n/a | Internal HPLC Standard | 11110141_EN Rev. 03, 2018-05 | IFU Document | 0.702 | 5.625 | 3.5 | 2.75 | 0.319 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097503 |  | dr11 | 00630414596099 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Cannabinoids THC (Thc) | Thc 20 | 0346 | 1520 | 4 x 380 |  |  | 7 | 30 |  | ;#Urine;# | 7 days at 25°C; 30 days at 2–8°C; 12 months at -20°C. | 25 μL |  |  | EMIT | 20 ng/mL; Semi-quant 10 - 55 ng/mL | n/a | n/a | n/a | qualitative/semi-quantitative assay | n/a |  |  |  | Emit Calibrator 0, 2, 3, 4 | 10445406 10445408 10445409 10445410 | 5 weeks | Liquid |  | 3 | 5 |  | 60 | 5 |  | No |  | No | [12 empty] | 7 ng/mL | 6 ng/mL | n/a | n/a | n/a | n/a | 19 ng/mL | N/A | 2.8 | 4.8 | 19 ng/mL |  | 6.1 | 6.5 | 23 ng/mL |  | 1.6 | 2.3 | Viva-E Analyzer ± 50% of 20 ng/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | Traceable to Syva Emit Cal/Cont which are gravimetrically prepared standards- qualified by GC/MS within ± 10% of nominal. | 11110121_EN Rev. 03, 2019-05 | IFU Document | 0.911 | 5.625 | 3.5 | 2.75 | 0.414 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097503 |  | dr11 | 00630414596099 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Cannabinoids THC (Thc) | Thc 50 | 0308 | 1520 | 4 x 380 |  |  | 7 | 30 |  | ;#Urine;# | 7 days at 25°C; 30 days at 2–8°C; 12 months at -20°C. | 13 μL |  |  | EMIT | 50 ng/mL; Semi-quant 15 - 180 ng/mL | n/a | n/a | n/a | qualitative/semi-quantitative assay | n/a |  |  |  | Emit Calibrator 0, 2, 3, 4, 5 | 10445406 10445408 10445409 10445410 10445411 | 5 weeks | Liquid |  | 3 | 5 |  | 60 | 5 |  | No |  | No | [12 empty] | 7 ng/mL | 6 ng/mL | n/a | n/a | n/a | n/a | 23 ng/mL | N/A | 3.7 | 8 | 47 ng/mL | N/A | 2 | 3.2 | 62 ng/mL | N/A | 1.2 | 4 | Viva-E Analyzer ± 50% of 50 ng/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | Traceable to Syva Emit Cal/Cont which are gravimetrically prepared standards- qualified by GC/MS within ± 10% of nominal. | 11110121_EN Rev. 03, 2019-05 | IFU Document | 0.911 | 5.625 | 3.5 | 2.75 | 0.414 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097503 |  | dr11 | 00630414596099 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Cannabinoids THC (Thc) | Thc 100 | 0345 | 1520 | 4 x 380 |  |  | 7 | 30 |  | ;#Urine;# | 7 days at 25°C; 30 days at 2–8°C; 12 months at -20°C. | 9.5 μL |  |  | EMIT | 100 ng/mL; Semi-quant 15 - 180 ng/mL | n/a | n/a | n/a | qualitative/semi-quantitative assay | n/a |  |  |  | Emit Calibrator 0, 3, 4, 5 | 10445406 10445409 10445410 10445411 | 5 weeks | Liquid |  | 3 | 5 |  | 60 | 5 |  | No |  | No | [12 empty] | 9 ng/mL | 7 ng/mL | n/a | n/a | n/a | n/a | 73 ng/mL | N/A | 2.2 | 4.5 | 98 ng/mL | N/A | 1.2 | 2.6 | 121 ng/mL | N/A | 0.7 | 2 | Viva-E Analyzer ± 50% of 100 ng/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | Traceable to Syva Emit Cal/Cont which are gravimetrically prepared standards- qualified by GC/MS within ± 10% of nominal. | 11110121_EN Rev. 03, 2019-05 | IFU Document | 0.911 | 5.625 | 3.5 | 2.75 | 0.414 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097513 |  | th12 | 00630414596006 | Global | Therapeutic Drug Monitoring (TDM) | No | Atellica CH Theophylline (Theo) | Theo | 2008 | 400 | 4 x 100 |  |  | 6 | 30 per pack |  | ;#Serum;#Lithium Heparin Plasma;# | 8 hours at 25 °C; 7 days at 2-8 °C; 90 days at -20 °C | 5.0 μL |  |  | PETINIA | 2.0 - 40.0 µg/mL | 11.1 - 222.0 µmol/L | 2.0 - 80.0 μg/mL | 11.1 - 444.0 μmol/L | 10-20 μg/mL | 56-111 μmol/L |  |  |  | Atellica CH DRUG CAL | 11099336 | 90 days | Liquid |  | 3 | 5 |  | 30 | 7 |  | No |  | Yes | [12 empty] | 0.4 μg/mL | 0.3 μg/mL | n/a | 1.7 μmol/L | 2.2 μmol/L | n/a | 7.0 μg/mL | 38.9 µmol/L | 1.4 | 2.2 | 14.8 μg/mL | 82.1 µmol/L | 0.9 | 2.2 | 27.4 μg/mL | 152.1 µmol/L | 1.5 | 2 | Dimension RxL y = 0.97x - 0.1 µg/mL (y = 0.97x - 0.6 µmol/L) r = 0.993 |  | ≤ -3% change up to 1000 mg/dL of hemoglobin | ≤ 4% change up to 80 mg/dL of bilirubin | ≤ 10% change up to 1000 mg/dL of Intralipid | No Biotin in assay formulation |  |  | n/a | n/a | USP Theophylline Standard | 11110131_EN Rev. 02, 2017-10 | IFU Document | 0.634 | 5.625 | 3.5 | 2.75 | 0.288 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097517 |  | th13 | 00630414595962 | Global | Therapeutic Drug Monitoring (TDM) | No | Atellica CH Tobramycin (Tob) | Tob | 2014 | 400 | 4 x 100 | Atellica CH DRUG CAL II |  | 9 | 30 per pack |  | ;#Serum;#Lithium Heparin Plasma;# | 8 hours at 25°C; 72 hours at 2–8°C; 30 days at -20°C. | 3.8 μL |  |  | PETINIA | 0.3 - 12.0 µg/mL | 0.6 - 25.7 µmol/L | 0.3 - 24.0 μg/mL | 0.6 - 51.4 μmol/L | 4–10 μg/mL | 8.6–21.4 µmol/L |  |  |  | Atellica CH DRUG CAL II | 11099405 | 30 days | Liquid |  | 3 | 5 |  | 30 | 7 |  | No |  | Yes | [12 empty] | 0.2 μg/mL | 0.1 μg/mL | n/a | 0.2 μmol/L | 0.4 μmol/L | n/a | 1.3 μg/mL | 2.8 µmol/L | 5 | 7.8 | 2.8 μg/mL | 6.0 µmol/L | 3.2 | 4.6 | 9.8 μg/mL | 21.0 µmol/L | 1.2 | 1.5 | Dimension RxL y = 1.01x + 0.07 µg/mL (y = 1.01x + 0.15 µmol/L) r = 0.996 |  | ≤ 2% change up to 500 mg/dL of hemoglobin | ≤ 2% change up to 20 mg/dL of bilirubin | ≤ 4% change up to 800 mg/dL of Intralipid | No Biotin in assay formulation |  |  | n/a | n/a | USP TobramycinStandard | 11110135_EN Rev. 01, 2017-05 | IFU Document | 0.638 | 5.625 | 3.5 | 2.75 | 0.29 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097511 |  | th01 | 00630414596020 | Global | Therapeutic Drug Monitoring (TDM) | No | Atellica CH Vancomycin (Vanc) | Vanc | 2202 | 400 | 4 x 100 | Atellica CH DRUG CAL II |  | 9 | 30 per pack |  | ;#Serum;#Lithium Heparin Plasma;# | 8 hours at room temperature; 2 days at 2-8°C; 30 days at -20°C | 2.5 μL |  |  | PETINIA | 3.0 - 50.0 µg/mL | 2.1 - 34.5 µmol/L | 3.0 - 100.0 μg/mL | 2.1 - 69.0 μmol/L | Peak 18 - 26 µg/mLTrough 5 - 10 µg/mL | Peak 12.4 - 17.9 µmol/LTrough 3.5 - 6.9 µmol/L |  |  |  | Atellica CH DRUG II CAL | 11099405 | 30 days | Liquid |  | 3 | 5 |  | 30 | 7 |  | No |  | Yes | [12 empty] | 0.2 μg/mL | 0.1 μg/mL | 3.0 μg/mL | 0.1 μmol/L | 0.1 μmol/L | 2.1 μmol/L | 11.2 μg/mL | 7.7 μmol/L | 1.8 | 2 | 32.7 μg/mL | 22.6 μmol/L | 1.4 | 2 | 45.8 μg/mL | 31.6 μmol/L | 1.7 | 1.8 | Dimension y = 1.04x - 1.04 µg/mL (y = 1.04x - 0.7 µmol/L) r = 0.997 |  | ≤ 2.1% change up to 600 mg/dL of hemoglobin | ≤ 0.9% change up to 20 mg/dL of bilirubin | ≤ 2.7% change up to 1000 mg/dL of Intralipid | No Biotin in assay formulation |  |  | n/a | n/a | USP material | 11110129_EN Rev. 01, 2017-05 | IFU Document | 0.638 | 5.625 | 3.5 | 2.75 | 0.29 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097512 |  | th15 | 00630414596013 | Global | Therapeutic Drug Monitoring (TDM) | No | Atellica CH Valproic Acid (VPA) | VPA | 2201 | 400 | 4 x 100 | Atellica CH DRUG CAL II |  | 8 | 30 per pack |  | ;#Serum;#Lithium Heparin Plasma;# | 8 hours at 25°C; 2 days at 2-8°C; 30 days at -20°C | 2.5 μL |  |  | PETINIA | 3.0 - 150.0 µg/mL | 20.8 - 1040.1 µmol/L | 3.0 - 300.0 μg/mL | 20.8 - 2080.2 μmol/L | 50-100 μg/mL | 346-693 μmol/L |  |  |  | Atellica CH DRUG II CAL | 11099405 | 30 days | Liquid |  | 3 | 5 |  | 30 | 7 |  | No |  | Yes | [12 empty] | 1.5 μg/mL | 0.9 μg/mL | n/a | 6.2 μmol/L | 10.4 μmol/L | n/a | 24.0 μg/mL | 166.4 μmol/L | 1.7 | 2.9 | 53.3 μg/mL | 369.6 µmol/L | 1.2 | 1.7 | 97.5 μg/mL | 676.1 μmol/L | 1.3 | 1.7 | Dimension RxL y = 1.01x - 1.4 µg/mL (y = 1.01x - 9.7 µmol/L) r = 0.997 |  | 6% bias up to 1000 mg/dL (0.62 mmol/L) of hemoglobin | ≤ 1% bias up to 80 mg/dL of bilirubin | 4% bias up to 1000 mg/dL (11.3 mmol/L) of Intralipid | No Biotin in assay formulation |  |  | n/a | n/a | USP-grade material | 11110130_EN Rev. 02, 2017-05 | IFU Document | 0.636 | 5.625 | 3.5 | 2.75 | 0.289 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097518 |  | dr05 | 00630414595955 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Ecstasy (Xtc) | Xtc300 | 0510 | 400 | 4 x 100 |  |  | 7 | 30 |  | ;#Urine;# | 7 days at 25°C; 30 days at 2–8°C or stored at -20°C. | 23 μL |  |  | EMIT | 300 ng/mL; Semi-quant 75 – 500 ng/mL | n/a | n/a | n/a | qualitative/semi-quantitative assay | n/a |  |  |  | Emit Cal/Control 0 Emit II Plus 6-AM/Ecstasy Cal/Control Levels 1, 2, 3 | 10445406 10470441 10470442 10470443 | expiration date printed on the vial label | Liquid |  | 3 | 4 |  | 60 | 15 |  | No |  | No | [12 empty] | 40 ng/mL | 32 ng/mL | n/a | n/a | n/a | n/a | 231 ng/mL | N/A | 1.5 | 2.2 | 302 ng/mL |  | 1.5 | 7.6 | 391 ng/mL |  | 3.4 | 4.7 | Viva-E Analyzer ± 50% of 300 ng/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | Traceable to Syva Emit Cal/Cont and Emit II Plus 6-AM Ecstasy which are gravimetrically prepared standards- qualified by GC/MS within ± 10% of nominal. | 11110136_EN Rev. 02, 2019-09 | IFU Document | 0.636 | 5.625 | 3.5 | 2.75 | 0.289 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097518 |  | dr05 | 00630414595955 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Ecstasy (Xtc) | Xtc500 | 0511 | 400 | 4 x 100 |  |  | 7 | 30 |  | ;#Urine;# | 7 days at 25°C; 30 days at 2–8°C or stored at -20°C. | 16 μL |  |  | EMIT | 500 ng/mL; Semi-quant 75 – 1000 ng/mL | n/a | n/a | n/a | qualitative/semi-quantitative assay | n/a |  |  |  | Emit Cal/Control 0 Emit II Plus 6-AM/Ecstasy Cal/Control Levels 1, 2, 3, 4 | 10445406 10470441 10470442 10470443 10470444 | expiration date printed on the vial label | Liquid |  | 3 | 4 |  | 60 | 15 |  | No |  | No | [12 empty] | 46 ng/mL | 35 ng/mL | n/a | n/a | n/a | n/a | 391 ng/mL | N/A | 1.9 | 2.3 | 512 ng/mL |  | 4.4 | 4.9 | 681 ng/mL |  | 7.1 | 8 | Viva-E Analyzer ± 50% of 500 ng/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | Traceable to Syva Emit Cal/Cont and Emit II Plus 6-AM Ecstasy which are gravimetrically prepared standards- qualified by GC/MS within ± 10% of nominal. | 11110136_EN Rev. 02, 2019-09 | IFU Document | 0.636 | 5.625 | 3.5 | 2.75 | 0.289 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097590 |  | ge13 | 00630414596501 | Global | General Chemistry | No | Atellica CH Albumin BCG (Alb) | AlbAlb | 1106 | 6800 | 4 x 1700 | N/A | N/A | 10 | 60 | Bromocresol green | ;#Serum;#Lithium Heparin Plasma;# | Up to 3 days at 2–8°C Frozen at -20°C for up to 60 days | 2 μL | No |  | Dye Binding: bromocresol green | 1.0 - 6 g/dL | 10 - 60 g/L | 12.0 g/dL | 120 g/L | 3.2–4.8 g/dL | 32–48 g/L | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 days | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine serum | 97 | 60 |  | No |  | Yes | N/A | 11099300 |  | 1.0 - 12.0 g/dL | 10 - 120 g/L | No | N/A | N/A | 0.0 g/dL | 0 g/L | N/A | N/A | 0.0 g/dL | 0.0 g/dL | N/A | 0 g/L | 0 g/L | N/A | 2.1 g/dL | 21 g/L | 1.7 | 2.4 | 3.4 g/dL | 34 g/L | 1.3 | 2 | 5.1 g/dL | 51 g/L | 1 | 1.6 | ADVIA Chemistry 1800 y = 0.99x + 0.1 g/dL r=0.995 n=121 | ADVIA Chemistry 1800 y = 0.99x + 1 g/L r=0.995 n=121 | ≤ 10% at 400 mg/dL (0.25 mmol/L) of hemoglobin | ≤ 10% at 30 mg/dL (513 µmol/L) of bilirubin, conjugated ≤ 10% at 30 mg/dL (513 µmol/L) of bilirubin, unconjugated | ≤ 10% at 1000 mg/dL (11.3 mmol/L) of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | BCG reference method NIST SRM 927reference materials | 11110153_EN Rev. 03, 2019-07 | IFU Document | 0.535 | 2.875 | 3.5 | 2.75 | 0.243 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097530 |  | ge15 | 00630414595825 | Global | General Chemistry | No | Atellica CH Albumin BCP (AlbP) | AlbPAlbP | 0148 | 6400 | 4 x 1600 | N/A | NA | 3 | 20 | Bromocresol Purple | ;#Serum;#Lithium Heparin Plasma;#Potassium EDTA Plasma;# | Up to 3 days at 2–8°C Frozen at -20°C for up to 30 days Thawed frozen specimens which are turbid must be clarified by centrifugation prior to testing | 4 μL | No |  | Dye Binding: bromocresol purple | 0.5 - 8 g/dL | 5 - 80 g/L | N/A | N/A | 3.4–5.0 g/dL | 34–50 g/L | N/A | N/A | N/A | Atellica CH AlbP CAL | 11099310 | 8 hours | Lyophilized | Preparation required refer to IFU | 3 | 2 | Human serum-based product containing albumin | 30 | 8 |  | No |  | No | N/A | 11099300 |  | N/A | N/A | No | N/A | N/A | 0.2 g/dL | 2 g/L | N/A | N/A | 0.2 g/dL | 0.1 g/dL | 0.4 g/dL | 1 g/L | 2 g/L | 4 g/L | 2.7 g/dL | 27 g/L | 1.17 | 1.17 | 3.6 g/dL | 36 g/L | 1.07 | 1.42 | 7.1 g/dL | 71 g/L | 0.73 | 1.02 | ADVIA Chemistry 1800 ALBP y = 0.99x + 0.01 g/dL r=0.996 n=59 | ADVIA Chemistry 1800 ALBP y = 0.99x + .1 g/L r=0.996 n=59 | ≤ 10% at 600 mg/dL (0.37 mmol/L) of hemoglobin | ≤ 10% at 30 mg/dL (513 µmol/L) of bilirubin, conjugated ≤ 10% at 30 mg/dL (513 µmol/L) of bilirubin, unconjugated | ≤ 10% at 500 mg/dL (5.65 mmol/L) of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | ERM DA470k Reference Material | 11110148_EN Rev. 01, 2017-05 | IFU Document | 0.519 | 2.875 | 3.5 | 2.75 | 0.236 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097605 |  | ge17 | 00630414596365 | Global | General Chemistry | No | Atellica CH Alanine Aminotransferase (ALT) | ALTALT | 1105 | 2550 | 3 x 850 | N/A | N/A | 10 | 90 | L-Alanine; LD (pig heart) | ;#Serum;#Lithium Heparin Plasma;# | Separated specimens may be stored for up to 7 days at 2–8°C or stored frozen for up to 30 days at -20°C or colder. Avoid repeating freezing and thawing | 25 μL | No |  | Modified IFCC | 7 - 1100 U/L | 7 - 1100 U/L | 7 - 3300 U/L | 7 - 3300 U/L | 10–49 U/L | 10–49 U/L | N/A | N/A | N/A | Atellica CH ENZ 2 CAL | 11099318 | 30 days at 2-8 °C | Liquid | Thaw the calibrator at room temperature for 30–45 minutes; Invert 10 times to mix | 3 | 2 | Bovine serum albumin: (alanine aminotransferase (porcine heart) | 131 | 30 |  | No |  | Yes | N/A | 11099300 |  | 7 - 3300 U/L | 7 - 3300 U/L | No | N/A | N/A | 7.0 U/L | 7.0 U/L | N/A | N/A | 2 U/L | 1 U/L | N/A | 1 U/L | 2 U/L | N/A | 27 U/L | 27 U/L | 1.2 | 2.8 | 79 U/L | 79 U/L | 1.7 | 2.2 | 843 U/L | 843 U/L | 0.3 | 1 | ADVIA Chemistry 1800 ALT y = 1.00x -1 U/L r=1.00 n=105 | ADVIA Chemistry 1800 ALT y = 1.00x -1 U/L r=1.00 n=105 | ≤ 10% at 450 mg/dL (0.28 mmol/L) of hemoglobin | ≤ 10% at 30 mg/dL (513 µmol/L) of bilirubin, conjugated ≤ 10% at 30 mg/dL (513 µmol/L) of bilirubin, unconjugated | ≤ 10% at 450 mg/dL (5.09 mmol/L) of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | IFCC reference method, IFCC-454 | 11110168_EN Rev. 04, 2019-06 | IFU Document | 0.719 | 4.187 | 3.5 | 2.75 | 0.327 | 10.636 | 8.89 | 6.985
Atellica® CH | 11097638 |  | ge17 | 00630414009056 | Global | General Chemistry | No | Atellica CH Alanine Aminotransferase P5P (ALTPLc) | ALTPLcALTPLc | 4035 | 2080 | 2 x 1040 | N/A | N/A | 10 | 19 per pack | L-Alanine; LD (pig heart) | ;#Serum;#Lithium Heparin Plasma;# | Separated specimens stored up to 7 days at 2–8°C; Frozen at -20°C for up to 30 days; avoid repeating freezing and thawing | 25 μL | No |  | IFCC Standardization | 9 - 1000 U/L | 9 - 1000 U/L | 6000 U/L | 6000 U/L | 7–40 U/L | 7–40 U/L | N/A | N/A | N/A | Atellica CH ENZ 2 CAL | 11099318 | 30 days | Liquid | Thaw the calibrator at room temperature for 30–45 minutes; Invert 10 times to mix | 3 | 2 | Bovine serum albumin: (alanine aminotransferase (porcine heart) | 130 | 4 |  | No |  | Yes | N/A | 11099300 |  | 9 - 6000 U/L | 9 - 6000 U/L | No | N/A | N/A | 3 U/L | 3 U/L | N/A | N/A | 3 U/L | 2 U/L | 5 U/L | 2 U/L | 3 U/L | 5 U/L | 36 U/L | 36 U/L | 2 | 2.8 | 102 U/L | 102 U/L | 1.8 | 2.7 | 218 U/L | 218 U/L | 0.6 | 1.3 | ADVIA Chemistry 1800 y = 1.02x - 0 U/L r=0.999 n=103 | ADVIA Chemistry 1800 y = 1.02x - 0 U/L r=0.999 n=103 | ≤ 10% at 200 mg/dL (0.12 mmol/L) of hemoglobin | ≤ 10% at 20 mg/dL (342 µmol/L) of bilirubin, conjugated ≤ 10% at 20 mg/dL (342 µmol/L) of bilirubin, unconjugated | ≤ 10% at 400 mg/dL (4.5 mmol/L) of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | IFCC reference method | 11111573_EN Rev. 02, 2019-11 | IFU Document | 0.636 | 5.625 | 3.5 | 2.75 | 0.289 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097603 |  | ge19 | 00630414596389 | Global | General Chemistry | No | Atellica CH Amylase (Amylas) Plasma/ Serum | AmylasAmylas | 0118 | 1050 | 3 x 350 | N/A | N/A | 9 | 30 | α-Glucosidase (microorganism); Ethylidene-4-NP-G7 | ;#Serum;#Lithium Heparin Plasma;# | Separated specimens stored up to 7 days at room temperature; Frozen at -20°C or colder for longer storage | 16 μL | No |  | Ethylidene Blocked-pNPG7 | 5 - 1500 U/L | 5 - 1500 U/L | 4500 U/L | 4500 U/L | 30–118 U/L | 30–118 U/L | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 3 | 1 | N/A | N/A | 1 |  | No |  | Yes | N/A | 11099300 |  | 5 - 4500 U/L | 5 - 4500 U/L | No | N/A | N/A | 4 U/L | 4 U/L | N/A | N/A | 4 U/L | 2 U/L | N/A | 2 U/L | 4 U/L | N/A | 43 U/L | 43 U/L | 1.9 | 2.2 | 134 U/L | 134 U/L | 0.5 | 0.9 | 1334 U/L | 1334 U/L | 0.3 | 0.4 | ADVIA Chemistry 1800 y = 1.05x - 2 U/L r=1.000 n=100 | ADVIA Chemistry 1800 y = 1.05x - 2 U/L r=1.000 n=100 | ≤ 10% at 525 mg/dL (0.336 mmol/L) of hemoglobin | ≤ 10% at 30 mg/dL (513 µmol/L) of bilirubin, conjugated ≤ 10% at 30 mg/dL (513 µmol/L) of bilirubin, unconjugated | ≤ 10% at 650 mg/dL (7.35 mmol/L) of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | IFCC reference method, IFCC-456 | 11110166_EN Rev. 02, 2019-09 | IFU Document | 0.607 | 4.187 | 3.5 | 2.75 | 0.276 | 10.636 | 8.89 | 6.985
Atellica® CH | 11097603 |  | ge19 | 00630414596389 | Global | General Chemistry | No | Atellica CH Amylase (Amylas) (urine) | AmylasAmylas | 0118 | 1050 | 3 x 350 | N/A | N/A | 9 | 30 | α-Glucosidase (microorganism); Ethylidene-4-NP-G7 | ;#Urine;# | Unstable in acid, adjust to pH of 7.0 then store refrigerated | 16 μL | No |  | Ethylidene Blocked-pNPG7 | 5 - 1500 U/L | 5 - 1500 U/L | 4500 U/L | 4500 U/L | ≤ 650 U/L | ≤ 650 U/L | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 3 | 1 | N/A | N/A | 1 |  | No |  | Yes | N/A | 11099300 |  | 5 - 4500 U/L | 5 - 4500 U/L | No | N/A | N/A | 3 U/L | 3 U/L | N/A | N/A | 3 U/L | 2 U/L | N/A | 2 U/L | 3 U/L | N/A | 51 U/L | 51 U/L | 2.2 | 2.9 | 491 U/L | 491 U/L | 0.3 | 1.1 | 1396 U/L | 1396 U/L | 0.5 | 1.5 | ADVIA Chemistry 1800 y = 1.05x + 2 U/L r=1.000 n=100 | ADVIA Chemistry 1800 y = 1.05x + 2 U/L r=1.000 n=100 | N/A | N/A | N/A | No Biotin in assay formulation | N/A | N/A | N/A | N/A | IFCC reference method, IFCC-456 | 11110166_EN Rev. 02, 2019-09 | IFU Document | 0.607 | 4.187 | 3.5 | 2.75 | 0.276 | 10.636 | 8.89 | 6.985
Atellica® CH | 11097607 |  | ge03 | 00630414596341 | Global | General Chemistry | No | Atellica CH Aspartate Aminotransferase (AST) | ASTAST | 1127 | 2550 | 3 x 850 | N/A | N/A | 10 | 90 | L-aspartic acid (580 mmol/L); MDH; LD (pig heart); α-Ketoglutarate; NADH | ;#Serum;#Lithium Heparin Plasma;# | Separated specimens stored up to 3 days at 20–25°C; up to 7 days at 2–8°C; frozen at -20°C or colder for up to 30 days | 25 μL | No |  | Modified IFCC | 8 - 1000 U/L | 8 - 1000 U/L | 6000 U/L | 6000 U/L | <34 U/L | <34 U/L | N/A | N/A | N/A | Atellica CH ENZ 2 CAL | 11099318 | 30 days | Liquid | Thaw the calibrator at room temperature for 30–45 minutes; Invert 10 times to mix | 3 | 2 | Bovine serum albumin: (alanine aminotransferase (porcine heart) | 131 | 30 |  | No |  | Yes | N/A | 11099300 |  | 8 - 6000 U/L | 8 - 6000 U/L | No | N/A | N/A | 2 U/L | 2 U/L | N/A | N/A | 2 U/L | 1 U/L | n/a | 1 U/L | 2 U/L | N/A | 30 U/L | 30 U/L | 2.9 | 3.4 | 116 U/L | 116 U/L | 1.3 | 1.7 | 861 U/L | 861 U/L | 0.4 | 0.9 | ADVIA Chemistry 1800 AST y = 1.04x - 5 U/L r=1.000 n=104 | ADVIA Chemistry 1800 AST y = 1.04x - 5 U/L r=1.000 n=104 | Avoid hemolyzed samples | ≤ 10% at 30 mg/dL (513 µmol/L) of bilirubin, conjugated ≤ 10% at 30 mg/dL (513 µmol/L) of bilirubin, unconjugated | ≤ 10% at 400 mg/dL (4.52 mmol/L) of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | IFCC reference method | 11110170_EN Rev. 03, 2020-08 | IFU Document | 0.724 | 4.187 | 3.5 | 2.75 | 0.329 | 10.636 | 8.89 | 6.985
Atellica® CH | 11097639 |  | ge03 | 00630414009063 | Global | General Chemistry | No | Atellica CH Aspartate Aminotransferase P5P (ASTPLc) | ASTPLcASTPLc | 4036 | 2080 | 2 x 1040 | N/A | N/A | 10 | 30 per pack | L-aspartic acid (580 mmol/L); MDH; LD (pig heart); α-Ketoglutarate; NADH | ;#Serum;#Lithium Heparin Plasma;# | Separated specimens stored up to 3 days at 20–25°C; up to 7 days at 2–8°C; frozen at -20°C or colder for up to 30 days | 25 μL | No |  | IFCC | 8 - 1000 U/L | 8 - 1000 U/L | 6000 U/L | 6000 U/L | 13–40 U/L | 13–40 U/L | N/A | N/A | N/A | Atellica CH ENZ 2 CAL | 11099318 | 30 days | Liquid | Thaw the calibrator at room temperature for 30–45 minutes; Invert 10 times to mix | 3 | 2 | Bovine serum albumin: aspartate aminotransferase (porcine heart)) | 130 | 11 |  | No |  | Yes | N/A | 11099300 |  | 8 - 6000 U/L | 8 - 6000 U/L | No | N/A | N/A | 1 U/L | 1 U/L | N/A | N/A | 1 U/L | 0 U/L | 7 U/L | 0 U/L | 1 U/L | 7 U/L | 34 U/L | 34 U/L | 1.6 | 1.8 | 98 U/L | 98 U/L | 1.4 | 1.5 | 238 U/L | 238 U/L | 0.9 | 1 | ADVIA Chemistry 1800; y = 1.00x - 4 U/L r=0.999 n=110 | ADVIA Chemistry 1800; y = 1.00x - 4 U/L r=0.999 n=110 | Avoid hemolyzed samples | ≤ 10% at 20 mg/dL (342 µmol/L) of bilirubin, conjugated ≤ 10% at 20 mg/dL (342 µmol/L) of bilirubin, unconjugated | ≤ 10% at 200 mg/dL (2.26 mmol/L) of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | IFCC reference method | 11111574_EN Rev. 01, 2017-05 | IFU Document | 0.64 | 5.625 | 3.5 | 2.75 | 0.291 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097595 |  | ge05 | 00630414595672 | Global | General Chemistry | No | Atellica CH Calcium (CPC) | Ca (Serum/plasma)Ca (Serum/plasma) | 0307 | 3560 | 4 x 890 | N/A | N/A | 8 | 30 | Ethanolamine buffer; o‑cresolphthalein complexone; 8‑hydroxyquinoline | ;#Serum;#Lithium Heparin Plasma;# | Up to 8 hours at room temperature; up to 2 days at 2–8°C; Frozen at -20°C or colder for up to 6 months | 16 μL | No |  | CPC | 1 - 15 mg/dL | 0.25 - 3.75 mmol/L | 30.0 mg/dL | 7.50 mmol/L | 8.3–10.6 mg/dL | 2.08–2.65 mmol/L | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 days | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 33 | 3 |  | No |  | Yes | N/A | 11099300 |  | 1 - 30.0 mg/dL | 0.25 - 7.50 mmol/L | No | N/A | N/A | 0.3 mg/dL | 0.08 mmol/L | N/A | N/A | 0.3 mg/dL | 0.2 mg/dL | N/A | 0.05 mmol/L | 0.08 mmol/L | N/A | 6.1 mg/dL | 1.53 mmol/L | 2.4 | 2.9 | 9.1 mg/dL | 2.28 mmol/L | 1.7 | 1.8 | 12.6 mg/dL | 3.15 mmol/L | 1.2 | 1.6 | ADVIA Chemistry 1800 Ca (serum) y = 0.98x + 0.00 mg/dL r=0.999 n=100 | ADVIA Chemistry 1800 Ca (serum) y = 0.98x + 0.00 mmol/L r=0.999 n=100 | ≤ 10% at 1000 mg/dL (0.62 mmol/L) of hemoglobin | ≤ 10% at 30 mg/dL (513 µmol/L) of bilirubin, conjugated ≤ 10% at 30 mg/dL (513 µmol/L) of bilirubin, unconjugated | ≤ 10% at 650 mg/dL (7.35 mmol/L) of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | NIST atomic absorption reference method, NIST SRM 915 and SRM 909b reference materials | 11110158_EN Rev. 02, 2019-10 | IFU Document | 1.133 | 5.625 | 3.5 | 2.75 | 0.515 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097595 |  | ge05 | 00630414595672 | Global | General Chemistry | No | Atellica CH Calcium (CPC) | Ca (Urine)Ca (Urine) | 0307 | 3560 | 4 x 890 | N/A | N/A | 8 | 30 | Ethanolamine buffer; o‑cresolphthalein complexone; 8‑hydroxyquinoline | ;#Urine;# | 2 days at 20–25°C, 4 days at 4–8°C, Frozen for up to 3 weeks at -20°C | 16 μL | No |  | CPC | 1 - 30 mg/dL | 0.25 - 7.50 mmol/L | 60.0 mg/dL | 15.00 mmol/L | 100–300 mg/day | 2.50–7.50 mmol/day | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 days | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 33 | 3 |  | No |  | Yes | N/A | 11099300 |  | 1 - 60.0 mg/dL | 0.25 - 15.00 mmol/L | No | N/A | N/A | 0.5 mg/dL | 0.13 mmol/L | N/A | N/A | 0.5 mg/dL | 0.3 mg/dL | N/A | 0.08 mmol/L | 0.13 mmol/L | N/A | 7.5 mg/dL | 1.88 mmol/L | 3.9 | 4.4 | 11.2 mg/dL | 2.80 mmol/L | 2.5 | 4.5 | 22.1 mg/dL | 5.53 mmol/L | 1.4 | 2.6 | ADVIA Chemistry 1800 Ca; (urine) y = 1.03x - 0.2 mg/dL r=0.999 n= 100 | ADVIA Chemistry 1800 Ca; (urine) y = 1.03x - 0.05 mmol/L r=0.999 n= 100 | N/A | N/A | N/A | No Biotin in assay formulation | N/A | N/A | N/A | N/A | NIST atomic absorption reference method, NIST SRM 915 and SRM 909b reference materials | 11110158_EN Rev. 02, 2019-10 | IFU Document | 1.133 | 5.625 | 3.5 | 2.75 | 0.515 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097609 |  | ge21 | 00630414596327 | Global | General Chemistry | No | Atellica CH Cholesterol_2 (Chol_2) | Chol_2Chol_2 | 0314 | 8400 | 4 x 2100 | N/A | N/A | 10 | 26 | 4-aminoantipyrine ; peroxidase (horseradish); cholesterol oxidase (microbial); phenolcholesterol esterase (Pseudomonas) | ;#Serum;#Lithium Heparin Plasma;# | Up to 8 hours at 25°C; Up to 2 days at 2–8°C; Frozen at -20°C or colder for longer storage | 4 μL | No |  | Colorimetric: cholesterol esterase and cholesterol oxidase conversion followed by a Trinder endpoint | 25 – 618 mg/dL | 0.65 – 16.01 mmol/L | 1236 mg/dL | 32.01 mmol/L | Low-risk levels (desirable); 200 mg/dL Moderate-risk levels (borderline) 200–239 mg/dL High-risk levels ≥ 240 mg/dL | Low-risk levels (desirable); 5.18 mmol/L Moderate-risk levels (borderline) 5.18–6.19 mmol/L High-risk levels ≥ 6.22 mmol/L | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 days | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 50 | 7 |  | No |  | Yes | N/A | 11099300 |  | 25 - 1236 mg/dL | 0.65 - 32.01 mmol/L | No | N/A | N/A | 2 mg/dL | 0.04 mmol/L | N/A | N/A | 2 mg/dL | 1 mg/dL | 25 mg/dL | 0.03 mmol/L | 0.04 mmol/L | 0.65 mmol/L | 115 mg/dL | 2.97 mmol/L | 0.8 | 1.3 | 276 mg/dL | 7.14 mmol/L | 0.5 | 0.8 | 584 mg/dL | 15.13 mmol/L | 0.5 | 0.9 | ADVIA Chemistry 1800 Chol_2 y = 0.97x + 1 mg/dL r=0.997 n=100 | ADVIA Chemistry 1800 Chol_2 y = 0.97x + 0.03 mmol/L r=0.997 n=100 | ≤ 10% at 500 mg/dL (0.311 mmol/L) of hemoglobin | ≤ 10% at 60 mg/dL (1026 µmol/L) of bilirubin, conjugated ≤ 10% at 20 mg/dL (342 µmol/L) of bilirubin, unconjugated | ≤ 10% at 1000 mg/dL (11.3 mmol/L) of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | NCEP/CDC reference method; which uses NIST SRM 909 reference materials | 11110172_EN Rev. 04, 2020-11 | IFU Document | 0.541 | 2.875 | 3.5 | 2.75 | 0.246 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097640 |  | ge22 | 00630414006642 | Global | General Chemistry | No | Atellica CH Creatine Kinase (CK_L) | CK_LCK_L | 0321 | 996 | 3 x 332 | N/A | N/A | 10 | 29 | Imidazole buffer; EDTA; ADP; AMP; Diadenosine Pentaphosphate; NADP;G-6-PDH; HK; N-Acetylcysteine; Mg; Glucose; Creatine Phosphate; | ;#Serum;#Lithium Heparin Plasma;# | Up to 4 hours at 25°C,up to 5 days at 2–8°C Frozen for up to 2 months at -20°C | 4.5 μL | No |  | Enzymatic | 15–1300 U/L | 15–1300 U/L | 7800 U/L | 7800 U/L | Males is 46–171 U/L Females is 34–145 U/L | Males is 46–171 U/L Females is 34–145 U/L | N/A | N/A | N/A | Atellica CH ENZ 3 CAL | 11099319 | 30 days | Liquid | Thaw at room temperature. 30 -45 minutes; Invert 10 times | 3 | 2 | Human serum albumin; creatine kinase MM (human heart) | 202 | 21 |  | No |  | Yes | N/A | 11099300 |  | 15 - 7800 U/L | 15 - 7800 U/L | No | N/A | N/A | 6 U/L | 6 U/L | N/A | N/A | 6 U/L | 1 U/L | 6 U/L | 1 U/L | 6 U/L | 6 U/L | 198 U/L | 198 U/L | 0.8 | 1.1 | 962 U/L | 962 U/L | 0.4 | 0.7 | 1152 U/L | 1152 U/L | 0.3 | 0.7 | ADVIA Chemistry 1800 CK_L y = 0.96x - 3.1 U/L r=1.00 n=177 | ADVIA Chemistry 1800 CK_L y = 0.96x - 3.1 U/L r=1.00 n=177 | ≤ 10% at 125 mg/dL (1.25 g/L) of hemoglobin | ≤ 10% at 60 mg/dL (1026 µmol/L) of bilirubin, conjugated ≤ 10% at 60 mg/dL (1026 µmol/L) of bilirubin, unconjugated | ≤ 10% at 1000 mg/dL (11.3 mmol/L) of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | IFCC reference method | 11110165_EN Rev. 03, 2019-09 | IFU Document | 0.656 | 4.187 | 3.5 | 2.75 | 0.298 | 10.636 | 8.89 | 6.985
Atellica® CH | 11097596 |  | ge25 | 00630414596457 | Global | General Chemistry | No | Atellica CH Creatinine_2 (Crea_2) | Crea_2Crea_2 (serum/plasma) | 0322 | 5888 | 4 x 1472 | N/A | N/A | 7 | 17 | Sodium hydroxide; Picric acid | ;#Serum;#Lithium Heparin Plasma;# | Up to 24 hours at 18–26°C; Up to 7 days at 2–8°C; Frozen at-20°C or colder for up to 3 months | 24 μL | No |  | Jaffe, alkaline picrate, kinetic with blank rate correction | 0.15 - 30.00 mg/dL | 13 - 2652 µmol/L | 60 mg/dL | 5304 μmol/L | Males 0.70–1.30 mg/dL Females 0.55–1.02 mg/dL | Males 62–115 μmol/L Females 49–90 μmol/L | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 days | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 180 | 6 |  | No |  | Yes | N/A | 11099300 |  | 0.15 - 60 mg/dL | 13 - 5304 μmol/L | No | N/A | N/A | 0.08 mg/dL | 7 μmol/L | N/A | N/A | 0.08 mg/dL | 0.03 mg/dL | 0.15 mg/dL | 3 μmol/L | 7 μmol/L | 13 μmol/L | 0.66 mg/dL | 58 µmol/L | 1.2 | 2.8 | 6.35 mg/dL | 561 µmol/L | 0.7 | 1 | 26.00 mg/dL | 2298 µmol/L | 0.2 | 0.6 | ADVIA Chemistry 1800 ECRE_2 y = 0.98x + 0.00 mg/dL r=0.999 n=140 | ADVIA Chemistry 1800 ECRE_2 y = 0.99x + 0.00 µmol/L r=0.999 n=140 | Do not use hemolyzed samples ≤ 10% at 500 mg/dL (0.311 mmol/L) of hemoglobin | ≤ 10% at 20 mg/dL (342 µmol/L) of bilirubin, conjugated ≤ 10% at 10 mg/dL (171 µmol/L) of bilirubin, unconjugated | ≤ 10% at 500 mg/dL (5.65 mmol/L) of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | IDMS reference method via correlation ofpatient samples and NIST SRM967 reference material | 11110159_EN Rev. 04, 2019-07 | IFU Document | 1.096 | 5.625 | 3.5 | 2.75 | 0.498 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097596 |  | ge25 | 00630414596457 | Global | General Chemistry | No | Atellica CH Creatinine_2 (Crea_2) | Crea_2Crea_2 (urine) | 0322 | 5888 | 4 x 1472 | N/A | N/A | 7 | 17 | Sodium hydroxide; Picric acid | ;#Urine;# | Up to 4 days at 2–8°C; Frozen for storage longer than 4 days | 24 μL | No |  | Jaffe, alkaline picrate, kinetic with blank rate correction | 3.00 -245.00 mg/dL | 265 - 21658 µmol/L | 735 mg/dL | 64,975 μmol/L | Males 950–2490 mg/day Female 600–1800 mg/day | Males 8.4–22.0 mmol/dayFemales 5.3–15.9 mmol/day | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 2 days | Lyophilized | 3.0 mL of deionized water | 3 | 2 | Bovine Serum | 180 | 6 |  | No |  | Yes | N/A | 11099300 |  | 3.00 - 735 mg/dL | 265 - 64,975 μmol/L | No | N/A | N/A | 0.51 mg/dL | 45 μmol/L | N/A | N/A | 0.51 mg/dL | 0.35 mg/dL | 3.00 mg/dL | 31 μmol/L | 45 μmol/L | 265 μmol/L | 59.62 mg/dL | 5270 µmol/L | 0.3 | 0.6 | 133.31 mg/dL | 11,785 µmol/L | 0.2 | 0.7 | 188.61 mg/dL | 16,673 µmol/L | 0.3 | 0.9 | ADVIA Chemistry 1800 ECRE_2 y = 0.95x + 0.07 mg/dL r=0.999 n=109 | ADVIA Chemistry 1800 ECRE_2 y = 0.95x + 6 µmol/L r=0.999 n=109 | N/A | N/A | N/A | No Biotin in assay formulation | N/A | N/A | N/A | N/A | IDMS reference method via correlation ofpatient samples and NIST SRM967 reference material | 11110159_EN Rev. 04, 2019-07 | IFU Document | 1.096 | 5.625 | 3.5 | 2.75 | 0.498 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097635 |  | pl13 | 00630414595566 | Global | Specific Proteins | No | Atellica CH β2-Microglobulin (B2M) | B2MB2M | 0204 | 400 | 2 x 200 | N/A | N/A | 8 | 30 | Latex particles coated with antihuman β2-microglobulin antibodies (goat); BSA (bovine) | ;#Serum;#Lithium Heparin Plasma;#Potassium EDTA Plasma;# | Up to 7 days at 2–8°C; Frozen at-20°C or colder for up to 2 months | 5 μL | No |  | Latex-enhanced immunoturbidimetric assay | 0.25 – 18.00 mg/L | 0.25 – 18.00 mg/L | 90.00 mg/L | 90.00 mg/L | 1.00–2.40 mg/L for adults | 1.00–2.40 mg/L for adults | N/A | N/A | N/A | Atellica CH B2M CAL | 11099442 | 30 | Lyophilized | 1.0 mL of reagent grade water; stand for 10 minutes at room temperature; Invert to mix | 3 | 2 | Human β2‑microglobulin; bovine serum albumin (50 mg) | 60 | 21 |  | No |  | Yes | NA | 11099300 |  | 0.25 - 90.00 mg/L | 0.25 - 90.00 mg/L | No | NA | NA | 0.20 mg/L | 0.10 mg/L | NA | NA | 0.20 mg/L | 0.10 mg/L | N/A | 0.20 mg/L | 0.10 mg/L | N/A | 2.05 mg/L | 2.05 mg/L | 2.8 | 2.9 | 4.88 mg/L | 4.88 mg/L | 0.6 | 0.9 | 11.15 mg/L | 11.15 mg/L | 0.5 | 0.6 | ADVIA Chemistry 1800 B2M y = 1.02x - 0.18 mg/L r = 0.999 n = 101 | ADVIA Chemistry 1800 B2M y = 1.02x - 0.18 mg/L r = 0.999 n = 101 | ≤ 10% at 1000 mg/dL of hemoglobin | ≤ 10% at 60 mg/dL of bilirubin, conjugated ≤ 10% at 60 mg/dL of bilirubin, unconjugated | ≤ 10% at 1000 mg/dL of Intralipid | No Biotin in assay formulation | 90 mg/dL | 90 mg/dL | 90 mg/dL | 90 mg/dL | World Health Organization (WHO) 1stInternational Standard for β2-microglobulin | 11110198_EN Rev. 03, 2019-11 | IFU Document | 0.323 | 2.875 | 3.5 | 2.75 | 0.147 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097624 |  | pl02 | 00630414595665 | Global | Specific Proteins | No | Atellica CH Complement C3 (C3) | C3C3 | 0301 | 400 | 2 x 200 | N/A | N/A | 10 | 30 | Anti-C3 antibody (goat); Tris/HCI buffer; sodium chloride; sodium azide; Polyethylene glycol | ;#Serum;# | Up to 3 days at 2–8°C; Frozen at-20°C up to 3 weeks | 4.5 μL | No |  | PEG-enhanced immunoturbidimetric | 0.5 - 500.0 mg/dL | 0.01 - 5.00 g/L | 0.5 - 1125.0 mg/dL | 0.01 - 11.25 g/L | Newborns 58 –108 mg/dL 3 months 67 – 124 mg/dL 6 months 74 – 138 mg/dL 9 months 78 – 144 mg/dL 12 months 80 – 150 mg/dL 2–10 years 80 – 150 mg/dL 12–18 years 85 – 160 mg/dL 20 years 82–160 mg/dL 30 years; 84–160 mg/dL 40–70 years 90–170 mg/dL | Newborns 0.60–1.10 g/L 3 months 0.70–1.20 g/L 6 months 0.70–1.40 g/L 9 months; 0.80–1.40 g/L 12 months; 0.80–1.50 g/L 2–10 years 0.80–1.50 g/L 12–18 years 0.90–1.60 g/L 20 years; 0.80–1.60 g/L 30 years; 0.80–1.60 g/L 40–70 years 0.90–1.70 g/L | N/A | N/A | N/A | Atellica CH Liquid Specific Protein Calibrators (LSP CAL) | 11099434 | 28 days | Liquid | No preparation required | 3 | 6 | Proteins derived from human serum; sodium azide | 60 | 14 |  | No |  | Yes | Atellica CH Diluent | 11099300 |  | 0.5 - 1125.0 mg/dL | 0.01 - 11.25 g/L | No | N/A | N/A | 0.2 mg/dL | 0.00 g/L | NA | NA | 0.2 mg/dL | 0.0 mg/dL | N/A | 0.00 g/L | 0.00 g/L | NA | 80.6 mg/dL | 0.81 g/L | 0.6 | 0.2 | 158.1 mg/dL | 1.58 g/L | 0.9 | 1.4 | 346.0 mg/dL | 3.46 g/L | 1.2 | 1.4 | ADVIA Chemistry 1800: y = 0.99x + 0.6 mg/dL r=0.999n=103 | ADVIA Chemistry 1800: y = 0.99x + 0.01 g/L r=0.999 n=103 | ≤ 10% at 1000 mg/dL of hemoglobin | ≤ 10% at 25 mg/dL of bilirubin, conjugated ≤ 10% at 18.8 mg/dL of bilirubin, unconjugated | ≤ 10% at 1000 mg/dL of Lipemia (avian triglycerides) | No Biotin in assay formulation | 2100 mg/dL | 21.00 g/L | 2100 mg/dL | 21.00 g/L | IRMM reference material CRM 470 from IFCC | 11110187_EN Rev. 02, 2019-11 | IFU Document | 0.328 | 2.875 | 3.5 | 2.75 | 0.149 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097623 |  | pl15 | 00630414595658 | Global | Specific Proteins | No | Atellica CH Complement C4 (C4) | C4C4 | 0302 | 400 | 2 x 200 | N/A | N/A | 10 | 30 | Anti-C4 antibody (goat); Tris/HCI buffer; sodium chloride; sodium azide; Polyethylene glycol | ;#Serum;# | Up to 7 days at 2–8°C; Frozen at 20°C for up to up to 3 months | 6 μL | No |  | PEG-enhanced immunoturbidimetric | 0.4 - 140.0 mg/dL | 0.01 - 1.40 g/L | 0.4 - 420.0 mg/dL | 0.01 - 4.20 g/L | 12.0–36.0 mg/dL for adults | 0.12–0.36 g/L for adults | N/A | N/A | N/A | Atellica CH Liquid Specific Protein Calibrators (LSP CAL) | 11099434 | 28 days | Liquid | No preparation required | 3 | 6 | Proteins derived from human serum; sodium azide | 60 | 14 |  | No |  | Yes | Atellica CH Diluent | 11099300 |  | 0.4 - 420.0 mg/dL | 0.01 - 4.20 g/L | No | N/A | N/A | 0.2 mg/dL | 0.00 g/L | NA | NA | 0.2 mg/dL | 0.1 mg/dL | N/A | 0.00 g/L | 0.00 g/L | NA | 15.7 mg/dL | 0.16 g/L | 0.7 | 1.8 | 42.3 mg/dL | 0.42 g/L | 0.9 | 1.3 | 100.8 mg/dL | 1.01 g/L | 1.1 | 1.4 | ADVIA Chemistry 1800: y = 0.96x + 0.4 mg/dL r=0.999 n=103 | ADVIA Chemistry 1800: y = 0.96x + 0.00 g/L r=0.999 n=103 | ≤ 10% at 750 mg/dL of hemoglobin | ≤ 10% at 18.8 mg/dL of bilirubin, conjugated ≤ 10% at 18.8 mg/dL of bilirubin, unconjugated | ≤ 10% at 1000 mg/dL of Lipemia (avian triglycerides) | No Biotin in assay formulation | 300 mg/dL | 3.00 g/L | 300 mg/dL | 3.00 g/L | e IRMM reference material CRM 470 from IFCC | 11110186_EN Rev. 02, 2019-11 | IFU Document | 0.326 | 2.875 | 3.5 | 2.75 | 0.148 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097631 |  | pl17 | 00630414596198 | Global | Specific Proteins | No | Atellica CH C-Reactive Protein_2 (CRP_2) | CRP_2CRP_2 | 0320 | 1000 | 2 x 500 | N/A | N/A | 8 | 30 | CRP antibody synthetic latex (rabbit); sodium azide; sodium chloride; EDTA disodium salt hydrate; sodium azide | ;#Serum;#Lithium Heparin Plasma;# | Up to 3 days at 4–8°C; Frozen at-20°C or colder for up to 6 months | 4 μL | No |  | Latex-enhanced immunoturbidimetric assay | 0.4 - 30.4 mg/dL | 4 - 304 mg/L | 91.2 mg/dL | 912 mg/L | < 1.0 mg/dL for adults | < 10 mg/L for adults | N/A | N/A | N/A | Atellica CH C‑Reactive Protein_2 Calibrators (CRP_2 CAL) | 11099430 | 60 days | Liquid | No preparation required | 3 | 6 | Recombinant human c-reactive protein (CRP) in a stabilized protein matrix; sodium azide | 60 | 30 |  | No |  | Yes | Atellica CH Diluent | 11099300 |  | 0.4 - 91.2 mg/dL | 4 - 912 mg/L | No | N/A | N/A | 0.0 mg/dL | 0.0 mg/L | N/A | N/A | 0.0 mg/dL | 0.0 mg/dL | 0.3 mg/dL | 0.0 mg/L | 0.0 mg/L | 3 mg/L | 0.7 mg/dL | 7 mg/L | 2.3 | 2.3 | 3.3 mg/dL | 33 mg/L | 1.4 | 1.4 | 6.5 mg/dL | 65 mg/L | 0.7 | 0.9 | ADVIA Chemistry 1800 y = 1.02x - 0.1 mg/dL r=0.998 | ADVIA Chemistry 1800 y = 1.02x - 1 mg/L r=0.998 | ≤ 10% at 1000 mg/dL of hemoglobin | ≤ 10% at 60 mg/dL of bilirubin, conjugated ≤ 10% at 60 mg/dL of bilirubin, unconjugated | ≤ 10% at 1000 mg/dL of Intralipid | No Biotin in assay formulation | 100.0 mg/dL | 1000 mg/L | 100.0 mg/dL | 1000 mg/L | IRMM reference material ERMDA-470. | 11110194_EN Rev. 02, 2019-10 | IFU Document | 0.458 | 2.875 | 3.5 | 2.75 | 0.208 | 7.302 | 8.89 | 6.985
Atellica® CH | 11099315 |  | ge23 | 00630414596518 | Global | General Chemistry | No | Atellica CH A‑LYTE Integrated Multisensor (IMT Na K Cl) | Cl (serum/plasma)Cl (serum/plasma) | 0312 | 20000 | 4 x 5000 | A-LYTE IMT Standard A SMN 11099304 A‑LYTE IMT Standard B + Salt Bridge SMN 11099306 A‑LYTE IMT Diluent SMN11099305 A‑LYTE IMT Dilution Check SMN 11099325 | N/A | 18 sec | 14 days or 5000 samples per multisensor | N/A | ;#Serum;#Lithium Heparin Plasma;# | up to 7 days at 2–8°C; Frozen for up to 30 days at -20°C. | 25 μL | No |  | Indirect Integrated Multisensor Technology (IMT) | 50 - 200 mmol/L | 50 - 200 mEq/L | N/A | N/A | Serum-98–107 mmol/L Plasma-98–107 mmol/L | Serum-98–107 mEq/L Plasma-98–107 mEq/L | N/A | N/A | N/A | A‑LYTE IMT Standard A; A‑LYTE IMT Standard B + Salt Bridge | 11099304 11099306 | 30 days onboard | Liquid | No preparation required | 3 | 2 | N/A | automatic calibration every 4 hours | automatic calibration every 4 hours |  | Yes |  | No | N/A | 11099305 |  | N/A | N/A | No | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 38.3 mmol/L | N/A | N/A | 38.3 mEq/L | 78.1 mmol/L | 78.1 mEq/L | 0.9 | 1.5 | 114 mmol/L | 114 mEq/L | 0.7 | 0.9 | 189 mmol/L | 189 mEq/L | 0.5 | 0.7 | ADVIA Chemistry 1800 Cl y = 1.00x + 0.00 mmol/L r=0.998 n=108 | ADVIA Chemistry 1800 Cl y = 1.00&nbsp; + 0.00 mEq/L r=0.998 n=108 | ≤ 10% at 750 mg/dL (0.47 mmol/L)) of hemoglobin | ≤ 10% at 60 mg/dL of bilirubin, conjugated ≤ 10% at 60 mg/dL of bilirubin, unconjugated | ≤ 10% at 3000 mg/dL of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | coulometric reference method, NIST Reference Serum and materials | 11109447_EN Rev. 03, 2019-09 | IFU Document | 0.167 | 6.125 | 2.19 | 3.75 | 0.076 | 15.557 | 5.562 | 9.525
Atellica® CH | 11099315 |  | ge23 | 00630414596518 | Global | General Chemistry | No | Atellica CH A‑LYTE Integrated Multisensor (IMT Na K Cl) | Cl (urine)Cl (urine) | 0312 | 20000 | 4 x 5000 | A-LYTE IMT Standard A SMN 11099304 A‑LYTE IMT Standard B + Salt Bridge SMN 11099306 A‑LYTE IMT Diluent SMN11099305 A‑LYTE IMT Dilution Check SMN 11099325 | N/A | 27 sec | 14 days or 5000 samples per multisensor | N/A | ;#Urine;# | up to 7 days at 2–8°C or stored frozen for up to 30 days at -20°C. 24 hour urine collection for sodium, potassium, and chloride should be made without addition of preservatives. Store refrigerated at 2–8°C or frozen for delayed analysis. | 25 μL | No |  | Indirect Integrated Multisensor Technology (IMT) | 20 - 330 mmol/L | 20 - 330 mEq/L | 2-fold manual urine dilution | 2-fold manual urine dilution | 110–250 mmol/L/24 hr | 110–250 mEq/L/24 hr | N/A | N/A | N/A | A‑LYTE IMT Standard A; A‑LYTE IMT Standard B + Salt Bridge | 11099304 11099306 | 30 days onboard | Liquid | No preparation required | 3 | 2 | N/A | automatic calibration every 4 hours | automatic calibration every 4 hours |  | Yes |  | No | N/A | 11099305 |  | 2-fold manual urine dilution | 2-fold manual urine dilution | No | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 15.8 mmol/L | N/A | N/A | 15.8 mEq/L | 42.7 mmol/L | 42.7 mEq/L | 0.6 | 1.7 | 194 mmol/L | 194 mEq/L | 0.5 | 0.9 | 280 mmol/L | 280 mEq/L | 0.4 | 0.8 | ADVIA Chemistry 1800 Cl y = 0.99x - 0.28 mmol/L r=0.999 n=102 | ADVIA Chemistry 1800 Cl y = 0.99x - 0.28 mEq/L r=0.999 n=102 | ≤ 10% at 500 mg/dL of hemoglobin | ≤ 10% at 60 mg/dL of bilirubin, conjugated ≤ 10% at 60 mg/dL of bilirubin, unconjugated | ≤ 10% at 2000 mg/dL of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | coulometric reference method, NIST Reference Serum and materials | 1109447_EN Rev. 03, 2019-09 | IFU Document | 0.167 | 6.125 | 2.19 | 3.75 | 0.076 | 15.557 | 5.562 | 9.525
Atellica® CH | 11099315 |  | ge34 | 00630414596518 | Global | General Chemistry | No | Atellica CH A‑LYTE Integrated Multisensor (IMT Na K Cl) | K (serum/plasma)K (serum/plasma) | 1607 | 20000 | 4 x 5000 | A-LYTE IMT Standard A SMN 11099304 A‑LYTE IMT Standard B + Salt Bridge SMN 11099306 A‑LYTE IMT Diluent SMN11099305 A‑LYTE IMT Dilution Check SMN 11099325 | N/A | 18 sec | 14 days or 5000 samples per multisensor | N/A | ;#Serum;#Lithium Heparin Plasma;#Urine;# | up to 7 days at 2–8°C or stored frozen for up to 30 days at -20°C. | 25 μL | No |  | Indirect Integrated Multisensor Technology (IMT) | 1 - 10 mmol/L | 1 - 10 mEq/L | N/A | N/A | Serum-3.5–5.1 mmol/L Plasma-3.4–4.5 mmol/L | Serum-3.5–5.1&nbsp; mEq/L ; Plasma-3.4–4.5 mEq/L | N/A | N/A | N/A | A‑LYTE IMT Standard A; A‑LYTE IMT Standard B + Salt Bridge | 11099304 11099306 | 30 days onboard | Liquid | No preparation required | 3 | 2 | N/A | automatic calibration every 4 hours | automatic calibration every 4 hours |  | Yes |  | No | N/A | 11099305 |  | N/A | N/A | No | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 0.429 mmol/L | N/A | N/A | 0.429 mEq/L | 3.95 mmol/L | 3.95 mEq/L | 0.5 | 0.9 | 5.92 mmol/L | 5.92 mEq/L | 0.5 | 0.8 | 7.39 mmol/L | 7.39 mEq/L | 0.5 | 0.8 | ADVIA Chemistry 1800 y = 0.99x - 0.04 mmol/L r=1.000 n=103 | ADVIA Chemistry 1800 y = 0.99x - 0.04 mEq/L r=1.000 n=103 | N/A | <10% up to 60 mg/dL of bilirubin, conjugated and unconjugated | <10% up to 3000 mg/dL of Intralipid® | No Biotin in assay formulation | N/A | N/A | N/A | N/A | flame emission spectrophotometry reference method | 1109447_EN Rev. 03, 2019-09 | IFU Document | 0.167 | 6.125 | 2.19 | 3.75 | 0.076 | 15.557 | 5.562 | 9.525
Atellica® CH | 11099315 |  | ge34 | 00630414596518 | Global | General Chemistry | No | Atellica CH A‑LYTE Integrated Multisensor (IMT Na K Cl) | K (urine)K (urine) | 1607 | 20000 | 4 x 5000 | A-LYTE IMT Standard A SMN 11099304 A‑LYTE IMT Standard B + Salt Bridge SMN 11099306 A‑LYTE IMT Diluent SMN11099305 A‑LYTE IMT Dilution Check SMN 11099325 | N/A | 27 sec | 14 days or 5000 samples per multisensor | N/A | ;#Urine;# | 7 days at 2–8°C or stored frozen for up to 30 days at -20°C.24 hour urine collection for sodium, potassium, and chloride should be made without addition of preservatives. Store refrigerated at 2–8°C or frozen for delayed analysis. | 25 μL | No |  | Indirect Integrated Multisensor Technology (IMT) | 2 - 300 mmol/L | 2 - 300 mEq/L | 2-fold manual urine dilution | 2-fold manual urine dilution | Urine: 25–125 mmol/L/24 hr | Urine: 25–125 mEq/L/24 hr | N/A | N/A | N/A | A‑LYTE IMT Standard A; A‑LYTE IMT Standard B + Salt Bridge | 11099304 11099306 | 30 days onboard | Liquid | No preparation required | 3 | 2 | N/A | automatic calibration every 4 hours | automatic calibration every 4 hours |  | Yes |  | No | N/A | 11099305 |  | N/A | N/A | No | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 1.03 mmol/L | N/A | N/A | 1.03 mEq/L | 31.2 mmol/L | 31.2 mEq/L | 0.5 | 0.9 | 73.7 mmol/L | 73.7 mEq/L | 0.5 | 0.9 | 258 mmol/L | 258 mEq/L | 0.8 | 1.1 | ADVIA Chemistry 1800 y = 1.03x - 0.67 mmol/L r=0.998 n=105 | ADVIA Chemistry 1800 y = 1.03x - 0.67 mEq/L r=0.998 n=105 | <10% up to 500 mg/dL of hemoglobin | ≤ 10% change up to 60 mg/dL of bilirubin, conjugated and unconjugated | <10% change up to 2000 mg/dL of Intralipid® | No Biotin in assay formulation | N/A | N/A | N/A | N/A | flame emission spectrophotometry reference method | 1109447_EN Rev. 03, 2019-09 | IFU Document | 0.167 | 6.125 | 2.19 | 3.75 | 0.076 | 15.557 | 5.562 | 9.525
Atellica® CH | 11099315 |  | ge38 | 00630414596518 | Global | General Chemistry | No | Atellica CH A‑LYTE Integrated Multisensor (IMT Na K Cl) | Na (serum/plasma)Na (serum/plasma) | 1903 | 20000 | 4 x 5000 | A-LYTE IMT Standard A SMN 11099304 A‑LYTE IMT Standard B + Salt Bridge SMN 11099306 A‑LYTE IMT Diluent SMN11099305 A‑LYTE IMT Dilution Check SMN 11099325 | N/A | 18 sec | 14 days or 5000 samples per multisensor | N/A | ;#Serum;#Lithium Heparin Plasma;#Urine;# | up to 7 days at 2–8°C or Frozen for up to 30 days at -20°C. | 25 μL | No |  | Indirect Integrated Multisensor Technology (IMT) | 50 - 200 mmol/L | 50 - 200 mEq/L | N/A | N/A | Serum/ Plasma: 50–200 mmol/L | Serum-136–145 mEq/L Plasma-136–145 mEq/L | N/A | N/A | N/A | A‑LYTE IMT Standard A; A‑LYTE IMT Standard B + Salt Bridge | 11099304 11099306 | 30 days onboard | Liquid | No preparation required | 3 | 2 | N/A | automatic calibration every 4 hours | automatic calibration every 4 hours |  | Yes |  | No | N/A | 11099305 |  | N/A | N/A | No | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 42.4 mmol/L | N/A | N/A | 42.4 mEq/L | 73.6 mmol/L | 73.6 mEq/L | 0.5 | 0.8 | 116 mmol/L | 116 mEq/L | 0.6 | 0.1 | 156 mmol/L | 156 mEq/L | 0.5 | 0.9 | ADVIA Chemistry 1800 y = 1.00x - 2.00 mmol/L r=0.995 n=106 | ADVIA Chemistry 1800 y = 1.00x - 2.00 mEq/L r=0.995 n=106 | <10% up to 750 mg/dL of hemoglobin | <10% up to 60 mg/dL of bilirubin, conjugated and unconjugated | <10% up to 3000 mg/dL of Intralipid® | No Biotin in assay formulation | N/A | N/A | N/A | N/A | flame emission spectrophotometry reference method | 1109447_EN Rev. 03, 2019-09 | IFU Document | 0.167 | 6.125 | 2.19 | 3.75 | 0.076 | 15.557 | 5.562 | 9.525
Atellica® CH | 11099315 |  | ge38 | 00630414596518 | Global | General Chemistry | No | Atellica CH A‑LYTE Integrated Multisensor (IMT Na K Cl) | Na (urine)Na (urine) | 1903 | 20000 | 4 x 5000 | A-LYTE IMT Standard A SMN 11099304 A‑LYTE IMT Standard B + Salt Bridge SMN 11099306 A‑LYTE IMT Diluent SMN11099305 A‑LYTE IMT Dilution Check SMN 11099325 | N/A | 27 sec | 14 days or 5000 samples per multisensor | N/A | ;#Serum;#Lithium Heparin Plasma;#Urine;# | 7 days at 2–8°C or stored frozen for up to 30 days at -20°C.24 hour urine collection for sodium, potassium, and chloride should be made without addition of preservatives. Store refrigerated at 2–8°C or frozen for delayed analysis. | 25 μL | No |  | Indirect Integrated Multisensor Technology (IMT) | 10 - 300 mmol/L | 10 - 300 mEq/L | 2-fold manual urine dilution | 2-fold manual urine dilution | 40–220 mmol/L/24 hr | 40–220 mEq/L/24 hr | N/A | N/A | N/A | A‑LYTE IMT Standard A; A‑LYTE IMT Standard B + Salt Bridge | 11099304 11099306 | 30 days onboard | Liquid | No preparation required | 3 | 2 | N/A | automatic calibration every 4 hours | automatic calibration every 4 hours |  | Yes |  | No | N/A | 11099305 |  | N/A | N/A | No | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 5.35 mmol/L | N/A | N/A | 5.35 mEq/L | 32.2 mmol/L | 32.2 mEq/L | 1 | 1.5 | 177 mmol/L | 177 mEq/L | 0.7 | 1 | 284 mmol/L | 284 mEq/L | 0.6 | 1.1 | ADVIA Chemistry 1800 y = 1.01x + 0.97 mmol/L r=1.000 n=101 | ADVIA Chemistry 1800 y = 1.01x + 0.97 mEq/L r=1.000 n=101 | <10% up to 500 mg/dL of hemoglobin | <10% up to 60 mg/dL of bilirubin, conjugated and unconjugated | <10% up to 2000 mg/dL of Intralipid® | No Biotin in assay formulation | N/A | N/A | N/A | N/A | flame emission spectrophotometry reference method | 1109447_EN Rev. 03, 2019-09 | IFU Document | 0.167 | 6.125 | 2.19 | 3.75 | 0.076 | 15.557 | 5.562 | 9.525
Atellica® IM | 10995441 |  | on02 | 00630414597706 | OUS | Oncology | No | Atellica IM Alpha Fetoprotein (AFP) | Atellica IM AFP | AFP | 100 | 1x100 | Atellica IM CAL D | Atellica IM Multi‑Diluent 2 Atellica IM AFP MCM (SMN 10995443) | 10 | 28 | Rabbit polyclonal anti-AFP antibody | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;#Amniotic fluid;# | RT 8 hours; 2–8°C after 8 hours; Freeze at ≤ -20°C after 48 hours; Freeze only 1 time | 10 μL | No |  | 2‑site sandwich immunoassay | 1.3–1000.0 ng/mL | 1.08–830.00 IU/mL |  |  |  |  | 98.4% of the apparently healthy subjects had AFP values < 8.1 ng/mL (serum) |  |  | Atellica IM CAL D | 10995509 (2-pack) 10995510 (6-pack) | Recon at 2–8°C 28 days Recon at RT 4 hrs | Lyophilized | Add 2.0 mL of special reagent water. Let the vials stand for 15–20 minutes at room temperature | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of alphafetoprotein (human) and carcinoembryonic (human) | 35 | 28 |  | No |  | Yes | Atellica IM Multi‑Diluent 2 | 10995644 |  | 1.3 - 200,000 ng/mL | 1.08–166,000 IU/mL | No | NA | NA | 1.3 ng/mL | 1.08 IU/mL | NP | NP | 1.8 ng/mL | 1.2 ng/mL | 4.5 ng/mL | 1.00 IU/mL | 1.49 IU/mL | 3.74 IU/mL | 8.2 ng/mL (serum) | 6.80 IU/mL (serum) | 6.3 | 9 | 67.5 ng/mL | 56.03 IU/mL | 3.2 | 4.5 | 182.7 ng/mL | 151.66 IU/mL | 3.1 | 3.8 | ADVIA Centaur AFP (Serum) y = 0.958x + 0.121 ng/mL r= 0.997 ADVIA Centaur AFP (Amniotic Fluid) y = 1.05x + 0.027 μg/mL r= 0.99 | ADVIA Centaur AFP (Serum) y = 0.958x + 0.100 IU/mL r= 0.997 | NSI up to 500 mg/dL of hemoglobin | NSI up to 20 mg/dL of bilirubin | NSI up to 1000 mg/dL of triglycerides | No Biotin in assay formulation | 1,000,000 ng/mL | 830,000 IU/mL |  |  | WHO, human AFP (72/225) using highly purified AFP | 10995270_EN Rev. 03, 2020-06 | IFU Document | 0.211 | 6.969 | 2.087 | 1.457 | 0.096 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995442 |  | on02 | 00630414597713 | OUS | Oncology | No | Atellica IM Alpha Fetoprotein (AFP) | Atellica IM AFP | AFP | 500 | 5 x 100 | Atellica IM CAL D | Atellica IM Multi‑Diluent 2, Atellica IM AFP MCM | 10 | 28 | Rabbit polyclonal anti-AFP antibody | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;#Amniotic fluid;# | RT 8 hours; 2–8°C after 8 hours; Freeze at ≤ -20°C after 48 hours; Freeze only 1 time | 10 μL | No |  | 2‑site sandwich immunoassay | 1.3–1000.0 ng/mL | 1.08–830.00 IU/mL |  |  |  |  | 98.4% of the apparently healthy subjects had AFP values < 8.1 ng/mL (serum) |  |  | Atellica IM CAL D | 10995509 (2-pack) 10995510 (6-pack) | Recon at 2–8°C 28 daysRecon at RT 4 hrs | Lyophilized | Add 2.0 mL of special reagent water. Let the vials stand for 15–20 minutes at room temperature | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of alphafetoprotein (human) and carcinoembryonic (human) | 35 | 28 |  | No |  | Yes | Atellica IM Multi‑Diluent 2 | 10995644 |  | 1.3 - 200,000 ng/mL | 1.08–166,000 IU/mL | No | NA | NA | 1.3 ng/mL | 1.08 IU/mL | NP | NP | 1.8 ng/mL | 1.2 ng/mL | 4.5 ng/mL | 1.00 IU/mL | 1.49 IU/mL | 3.74 IU/mL | 8.2 ng/mL (serum) | 6.80 IU/mL (serum) | 6.4 | 9.1 | 67.5 ng/mL | 56.03 IU/mL | 3.2 | 4.5 | 182.7 ng/mL | 151.66 IU/mL | 3.1 | 3.8 | ADVIA Centaur AFP (Serum) y = 0.958x + 0.121 ng/mL r= 0.997 ADVIA Centaur AFP (Amniotic Fluid) y = 1.05x + 0.027 μg/mL r= 0.99 | ADVIA Centaur AFP (Serum) y = 0.958x + 0.100 IU/mL r= 0.997 | NSI up to 500 mg/dL of hemoglobin | NSI up to 20 mg/dL of bilirubin | NSI up to 1000 mg/dL of triglycerides | No Biotin in assay formulation | 1,000,000 ng/mL | 830,000 IU/mL |  |  | WHO, human AFP (72/225) using highly purified AFP | 10995270_EN Rev. 03, 2020-06 | IFU Document | 0.845 | 6.969 | 2.087 | 4.37 | 0.384 | 17.7 | 5.3 | 11.1
Atellica® IM | 10995471 |  | ca01 | 00630414597980 | OUS | Cardiac | No | Atellica IM B‑Type Natriuretic Peptide (BNP) | Atellica IM BNP | BNP | 100 |  | Atellica IM BNP CAL (SMN 10995473) | Atellica IM BNP QC (SMN 10995475) Atellica IM BNP MCM (SMN 10995474) | 10 | 42 |  | ;#EDTA Plasma;# | RT not recommended, at 2–8°C up to 24 hours; at ≤ -20°C after 24 hours up to 9 months, test within 8 hours after thawing | 100 μL | No |  |  | 2.0–5000.0 pg/mL | 0.58–1445.00 pmol/L |  |  |  |  | <100 pg/mL (w/o heart failure) | See IFU for expected results by age | See IFU for expected results by NYHA class | Atellica IM BNP CAL | 10995473 | Recon at 2–8°C 5 days, Recon at ≤ -20°C 60 days; thaw 1 time, Recon at RT 8 hrs | Lyophilized | 2.0 mL/vial | 3 for lot calibration, 2 for pack recalibration | 2 |  | 61 | 42 |  | No |  | Yes | Atellica IM Multi‑Diluent 15 | 11200798 |  | 2.0–50000.0 pg/mL | 0.58–14450.00 pmol/L | Yes | Atellica IM BNP QC | 10995475 | 1.4 pg/mL | 0.40 pmol/L | <1.5 pg/mL | <0.43 pmol/L | 2.4 pg/mL | 1.1 pg/mL | NP | 0.32 pmol/L | 0.69 pmol/L | NP | 12.2 pg/mL | 3.5 pmol/L | 2.6 | 3.5 | 654.6 pg/mL | 189.2 pmol/L | 1.8 | 2.8 | 4325.2 pg/mL | 1250.0pg/mL | 1.5 | 2.5 | ADVIA Centaur BNP y = 0.965x - 0.3 pg/mL r=0.997 | ADVIA Centaur BNP y = 0.965x - 0.091 pmol/L r=0.997 | ≤ 10% change up to 100 mg/dL of hemoglobin | ≤ 5% change up to 25 mg/dL of unconjugated bilirubin≤ 7% change up 25 mg/dL conjugated bilirubin | ≤ 5% change up to 800 mg/dL of triglycerides,1000 mg/dL of cholesterol | ≤ 10% change up to 30 ng/mL of biotin | 100,000 pg/mL | 28,900 pmol/L |  |  | traceable to internal standard, synthetic human BNP (amino acid 77–108) | 10995281_EN Rev. 04, 2019-07 | IFU Document | 0.253 | 6.969 | 2.087 | 1.457 | 0.115 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995472 |  | ca01 | 00630414597997 | OUS | Cardiac | No | Atellica IM B‑Type Natriuretic Peptide (BNP) | Atellica IM BNP | BNP | 500 |  | Atellica IM BNP CAL (SMN 10995473) | Atellica IM BNP QC (SMN 10995475) Atellica IM BNP MCM (SMN 10995474) | 10 | 42 |  | ;#EDTA Plasma;# | RT not recommended, at 2–8°C up to 24 hours; at ≤ -20°C after 24 hours up to 9 months, test within 8 hours after thawing | 100 μL | No |  |  | 2.0–5000.0 pg/mL | 0.58–1445.00 pmol/L |  |  |  |  | <100 pg/mL (w/o heart failure) | See IFU for expected results by age |  | Atellica IM BNP CAL | 10995473 | Recon at 2–8°C 5 days, Recon at ≤ -20°C 60 days; thaw 1 time, Recon at RT 8hrs | Lyophilized | 2.0 mL/vial | 3 for lot calibration, 2 for pack recalibration | 2 |  | 61 | 42 |  | No |  | Yes | Atellica IM Multi‑Diluent 15 | 11200798 |  | 2.0–50000.0 pg/mL | 0.58–14450.00 pmol/L | Yes | Atellica IM BNP QC | 10995475 | 1.4 pg/mL | 0.40 pmol/L | <1.5 pg/mL | <0.43 pmol/L | 2.4 pg/mL | 1.1 pg/mL | NP | 0.32 pmol/L | 0.69 pmol/L | NP | 12.2 pg/mL | 3.5 pmol/L | 2.6 | 3.5 | 654.6 pg/mL | 189.2 pmol/L | 1.8 | 2.8 | 4325.2 pg/mL | 1250.0pg/mL | 1.5 | 2.5 | ADVIA Centaur BNP y = 0.965x - 0.3 pg/mL r=0.997 | ADVIA Centaur BNP y = 0.965x - 0.091 pmol/L r=0.997 | ≤ 10% change up to 100 mg/dL of hemoglobin | ≤ 5% change up to 25 mg/dL of unconjugated bilirubin≤ 7% change up 25 mg/dL conjugated bilirubin | ≤ 5% change up to 800 mg/dL of triglycerides,1000 mg/dL of cholesterol | ≤ 10% change up to 30 ng/mL of biotin | 100,000 pg/mL | 28,900 pmol/L |  |  | traceable to internal standard, synthetic human BNP (amino acid 77–108) | 10995281_EN Rev. 04, 2019-07 | IFU Document | 0.891 | 6.969 | 2.087 | 4.37 | 0.405 | 17.7 | 5.3 | 11.1
Atellica® IM | 10995523 |  | on06 | 00630414598505 | Global | Oncology | No | Atellica IM Carcinoembryonic Antigen (CEA) | Atellica IM CEA | CEA | 100 | 1 x 100 | Atellica IM CAL D | Atellica IM CEA DIL (SMN 10995525) Atellica IM CEA MCM (SMN 10995526) | 14 | 14 | Rabbit polyclonal anti-CEA antibody (~400 ng/mL) | ;#Serum;# | RT 8 hrs 2–8°C after 8 hours; ≤ -20°C after 48 hours; Freeze only 1 time | 50 μL | No |  | 2-site sandwich immunoassay | 0.50–100.00 ng/mL | 0.50–100.00 μg/L |  |  |  |  | See IFU |  |  | Atellica IM CAL D | 10995509 (2-pack) 10995510 (6-pack) | Recon at 2–8°C 28 daysRecon at RT 4 hrs | Lyophilized | 2.0 mL of special reagent water. Let the vials stand for 15–20 minutes at room temperature | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of alphafetoprotein (human) and carcinoembryonic antigen (human) | 29 | 21 |  | No |  | Yes | Atellica IM CEA DIL | 10995525 |  | 0.50-10,000 ng/mL | 0.50-10,000 μg/L | No | NA | NA | 0.26 ng/mL | 0.26 μg/L | NP | NP | 0.88 pg/mL | 0.13 ng/mL | NP | 0.13 μg/L | 0.88 μg/L | NP | 3.12 ng/mL | 3.12 μg/L | 5 | 6 | 33.35 pg/mL | 33.35 μg/L | 1.9 | 2.9 | 84.81 ng/mL | 84.81 μg/L | 1.7 | 2.6 | ADVIA Centaur CEA y = 1.020x - 0.112 ng/mL r=0.997 | ADVIA Centaur CEA y = 1.020x - 0.112 μg/L r=0.997 | ≤ 5% change up to 500 mg/dL of hemoglobin | ≤ 5% change up to 20 mg/dL of unconjugated/conjugated bilirubin | ≤ 5% change up to 1000 mg/dL of triglycerides | No Biotin in assay formulation | 100,000 ng/mL | 100,000 μg/L |  |  | traceable to an internal standard manufactured using highlypurified material | 10995312_EN Rev. 03, 2020-02 | IFU Document | 0.209 | 6.969 | 2.087 | 1.457 | 0.095 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995524 |  | on06 | 00630414598512 | Global | Oncology | No | Atellica IM Carcinoembryonic Antigen (CEA) | Atellica IM CEA | CEA | 500 | 5 x 100 | Atellica IM CAL D | Atellica IM CEA DIL, Atellica IM CEA MCM (SMN 10995526) | 14 | 14 | Rabbit polyclonal anti-CEA antibody (~400 ng/mL) | ;#Serum;# | RT 8 hrs 2–8°C after 8 hours; ≤ -20°C after 48 hours; Freeze only 1 time | 50 μL | No |  | 2-site sandwich immunoassay | 0.50–100.00 ng/mL | 0.50–100.00 μg/L |  |  |  |  | See IFU |  |  | Atellica IM CAL D | 10995509 (2-pack) 10995510 (6-pack) | Recon at 2–8°C 28 days Recon at RT 4 hrs | Lyophilized | 2.0 mL of special reagent water. Let the vials stand for 15–20 minutes at room temperature | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of alphafetoprotein (human) and carcinoembryonic antigen (human) | 29 | 21 |  | No |  | Yes | Atellica IM CEA DIL | 10995525 |  | 0.50-10,000 ng/mL | 0.50-10,000 μg/L | No | NA | NA | 0.26 ng/mL | 0.26 μg/L | NP | NP | 0.88 pg/mL | 0.13 ng/mL | NP | 0.13 μg/L | 0.88 μg/L | NP | 3.12 ng/mL | 3.12 μg/L | 5 | 6 | 33.35 pg/mL | 33.35 μg/L | 1.9 | 2.9 | 84.81 ng/mL | 84.81 μg/L | 1.7 | 2.6 | ADVIA Centaur CEA y = 1.020x - 0.112 ng/mL r=0.998 | ADVIA Centaur CEA y = 1.020x - 0.112 μg/L r=0.998 | ≤ 5% change up to 500 mg/dL of hemoglobin | ≤ 5% change up to 20 mg/dL of unconjugated/conjugated bilirubin | ≤ 5% change up to 1000 mg/dL of triglycerides | No Biotin in assay formulation | 100,000 ng/mL | 100,000 μg/L |  |  | traceable to an internal standard manufactured using highlypurified material | 10995312_EN Rev. 03, 2020-02 | IFU Document | 0.843 | 6.969 | 2.087 | 4.37 | 0.383 | 17.7 | 5.3 | 11.1
Atellica® IM | 10995544 |  | on01 | 00630414598680 | OUS | Oncology | No | Atellica IM Complexed Prostate‑Specific Antigen(cPSA) | Atellica IM cPSA | cPSA | 100 | 1X100 | Atellica IM CAL Y Atellica IM cPSA PRE (SMN 10995546) | Atellica IM cPSA QC, Atellica IM Multi‑Diluent 2, Atellica IM cPSA MCM (SMN 10995545) | 14 | 28 | Goat polyclonal anti‑PSA antibody | ;#Serum;# | RT 8 hrs, 2–8°C after 8hrs, ≤ -20°C after 48 hours. Freeze only 1 time | 35 μL | No |  | 2-site sandwich immunoassay | 0.03–100.00 ng/mL | 0.03–100.00 μg/L |  |  |  |  | See IFU |  |  | Atellica IM CAL Y | 10995520 | Recon at 2–8°C 21 days, Recon at RT 8 hrs | Lyophilized | 2.0 mL of special reagent water, let the vials stand for 15–20 minutes at room temperature. | 3 for lot calibration, 2 for pack recalibration | 2 | PSA (human); goat serum | 63 | 28 |  | No |  | Yes | Atellica IM Multi‑Diluent 2 | 10995644 |  | 0.03- 500 ng/mL | 0.03-500 μg/L | Yes | Atellica IM cPSA QC | 10995547 | 0.03 ng/mL | 0.03 μg/L | NP | NP | 0.10 ng/mL | 0.05 ng/mL | 0.10 ng/mL | 0.05 μg/L | 0.10 μg/L | 0.10 μg/L | 1.07 ng/ml | 1.07 μg/L | 1.4 | 2.5 | 47.28 ng/mL | 47.28 μg/L | 1.2 | 2.5 | 86.8 ng/mL | 86.8 μg/L | 1.6 | 2.7 | ADVIA Centaur cPSA y = 0.97x + 0.182 ng/mL r=0.999 | ADVIA Centaur cPSA y = 0.97x + 0.182 μg/L r=0.999 | ≤ 5% change up to 500 mg/dL of hemoglobin | ≤ 5% change up to 40 mg/dL of bilirubin | ≤ 5% change up to 1000 mg/dL of triglycerides | No Biotin in assay formulation | 9137 ng/mL | 9137 μg/L |  |  | traceable to an internal standard manufactured using highlypurified PSA-ACT | 10995337_EN Rev. 02, 2019-10 | IFU Document | 0.222 | 6.969 | 2.087 | 1.457 | 0.101 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995561 |  | re04 | 00630414598857 | Global | Reproductive Endocrinology | No | Atellica IM Enhanced Estradiol (eE2) | Atellica IM eE2 | eE2 | 100 | 1 x 100 | Atellica IM CAL 30 | Atellica IM eE2 DIL, Atellica IM eE2 MCM (SMN 10995564) | 14 | 42 | Sheep monoclonal anti-estradiol antibody | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | RT 20 hours, 2–8°C after 20 hours, ≤ -20°C after 48 hours up to 6 months, Freeze only 1 time | 80 μL | No |  | Competitive assay | 11.80–3000.00 pg/mL | 43.60–11,010.00 pmol/L |  |  |  |  | See IFU |  |  | Atellica IM CAL 30 | 10995495 (2 pack) 10995497 (6 pack) | Recon at 2–8°C 14 days, Recon at RT 4hrs | Lyophilized | 2.0 mL of special reagent water. Let the vials stand for 15–20 minutes at room temperature | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of USP‑grade estradiol, testosterone, cortisol, and progesterone charcoal-stripped, defibrinated human plasma | 22 | 21 |  | No |  | Yes | Atellica IM eE2 DIL | 10995563 |  | 11.8-15.000 pg/mL | 43.60–11,010.00 pmol/L | No | NA | NA | NP | NP | 18.81 pg/mL | 69.03 pmol/L | 11.54 pg/mL | 7.33 pg/mL | NP | 26.91 pmol/L | 42.36 pmol/L | NP | 19.88 pg/mL | 72.96 pmol/L | 8.2 | 14.7 | 290.49 ng/mL | 1066.10 pmol/L | 1.6 | 3.6 | 2799.58 pg/mL | 10274.46 pmol/L | 1.9 | 2.7 | ADVIA Centaur eE2 (serum) y = 0.96x - 1.23 pg/mL r=0.998 | ADVIA Centaur eE2 (serum) y = 0.96x - 4.52 pmol/L r=0.998 | ≤ 10% change up to 250 mg/dL of hemoglobin | ≤ 10% change upto 20 mg/dL of conjugated/unconjugated bilirubin | ≤ 10% change up to 1000 mg/dL of triglycerides | ≤ 10% change up to 1500 ng/mL of biotin | NA | NA |  |  | standardized using internal standards manufactured analytically which are traceable to ID-GC/MS | (OUS )10995342_EN Rev. 04, 2019-07, (US) 11200383_EN Rev. 04, 2020-02 | IFU Document | 0.273 | 6.969 | 2.48 | 2.913 | 0.124 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995562 |  | re04 | 00630414598864 | Global | Reproductive Endocrinology | No | Atellica IM Enhanced Estradiol (eE2) | Atellica IM eE2 | eE2 | 500 | 5 x 100 | Atellica IM CAL 30 | Atellica IM eE2 DIL, Atellica IM eE2 MCM (SMN 10995564) | 14 | 42 | Sheep monoclonal anti-estradiol antibody | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | RT 20 hours, 2–8°C after 20 hours, ≤ -20°C after 48 hours up to 6 months, Freeze only 1 time | 80 μL | No |  | Competitive assay | 11.80–3000.00 pg/mL | 43.60–11,010.00 pmol/L |  |  |  |  | See IFU |  |  | Atellica IM CAL 30 | 10995495 (2 pack) 10995497 (6 pack) | Recon at 2–8°C 14 days, Recon at RT 4hrs | Lyophilized | 2.0 mL of special reagent water. Let the vials stand for 15–20 minutes at room temperature | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of USP‑grade estradiol, testosterone, cortisol, and progesterone charcoal-stripped, defibrinated human plasma | 22 | 21 |  | No |  | Yes | Atellica IM eE2 DIL | 10995563 |  | 11.8-15.000 pg/mL | 43.60–11,010.00 pmol/L | No | NA | NA | NP | NP | 18.81 pg/mL | 69.03 pmol/L | 11.54 pg/mL | 7.33 pg/mL | NP | 26.91 pmol/L | 42.36 pmol/L | NP | 19.88 pg/mL | 72.96 pmol/L | 8.2 | 14.7 | 290.49 ng/mL | 1066.10 pmol/L | 1.6 | 3.6 | 2799.58 pg/mL | 10274.46 pmol/L | 1.9 | 2.7 | ADVIA Centaur eE2 (serum) y = 0.96x - 1.23 pg/mL r=0.998 | ADVIA Centaur eE2 (serum) y = 0.96x - 4.52 pmol/L r=0.998 | ≤ 10% change up to 250 mg/dL of hemoglobin | ≤ 10% change upto 20 mg/dL of conjugated/unconjugated bilirubin | ≤ 10% change up to 1000 mg/dL of triglycerides | ≤ 10% change up to 1500 ng/mL of biotin | NA | NA |  |  | standardized using internal standards manufactured analytically which are traceable to ID-GC/MS | (OUS )10995342_EN Rev. 04, 2019-07, (US) 11200383_EN Rev. 04, 2019-06 | IFU Document | 1.052 | 10.748 | 2.48 | 6.063 | 0.478 | 27.3 | 6.3 | 15.4
Atellica® IM | 10995577 |  | on07 | 00630414599014 | OUS | Oncology | No | Atellica IM Free Prostate‑Specific Antigen (fPSA) | Atellica IM fPSA | fPSA | 50 | 1 x 50 | Atellica IM fPSA CAL | Atellica IM Multi‑Diluent 2 | 10 | 28 |  | ;#Serum;# | RT 8 hrs2–8°C after 8 hours ≤ -20°C after 48 hours;Freeze only 1 time | 30 μL | No |  |  | 0.01–25.00 ng/mL | 0.01–25.00 μg/L |  |  |  |  | See IFU |  |  | Atellica IM fPSA CAL | 10995578 | Recon at 2–8°C 21 days, Recon at RT 8hrs | Lyophilized | 2.0 mL of special reagent water. Let the vials stand for 15–20 minutes at room temperature | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of fPSA (human), goat serum | 54 | 28 |  | No |  | Yes | Atellica IM Multi‑Diluent 2 | 10995644 |  | 0.01-125 ng/mL | 0.01-125 μg/L | No | NA | NA | NP | NP | NP | NP | 0.01 ng/mL | 0.01 ng/mL | 0.01 ng/mL | 0.01 μg/L | 0.01 μg/L | 0.01 μg/L | 0.05 ng/mL | 0.05 μg/L | 4.6 | 6 | 3.6 ng/mL | 3.6 μg/L | 2.2 | 2.9 | 12.09 ng/mL | 12.09 μg/L | 2.1 | 3.3 | ADVIA Centaur fPSA y = 0.93x - 0.024 ng/mL (μg/L) r=0.999 | ADVIA Centaur fPSA y = 0.93x - 0.024 μg/L r=0.999 | ≤ 5% change up to 1000 mg/dL of hemoglobin | ≤ 5% change up to 25 mg/dL of conjugated/unconjugated bilirubin | ≤ 5% change up to 1000 mg/dL of triglycerides | ≤ 5% change in results up to up to 1500 ng/mL | 15,355 ng/mL | 15,355 μg/L |  |  | traceable to World Health Organization (WHO) InternationalStandard (96/668) | 10995345_EN Rev. 03, 2019-10 | IFU Document | 0.174 | 6.969 | 2.087 | 1.457 | 0.079 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995589 |  | ty02 | 00630414599137 | Global | N/AThyroid | No | Atellica IM Free Thyroxine (FT4) | Atellica IM FT4FT4 | FT4 | 50 | 1 x 50 | Atellica IM CAL A 2- pack SMN 10995500 Atellica IM CAL A 6- pack SMN 10995501 | Atellica IM FT4 MCM SMN 10995590 | 14 | 21 | acridinium-ester-labeled T4 and biotinylated rabbit poloclonal anti-T4 antibody | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | RT 8 hrs, 2–8°C after 8 hours; ≤ -20°C after 48 hours; Freeze only 1 time | 25 μL | No |  | Competitive immunoasay using direct chemiluminescent technology | 0.1 – 12.0 ng/dL | 1.3 – 154.8 pmol/L | NA | NA | 0.89- 1.76 ng/dL | 11.5 - 22.7 pmol/L | NA | NA | NA | Atellica IM CAL A | 10995500 (2-pack) 10995501 (6-pack) | 28 days recon at 2–8°C, 4 hours recon at RT | Lyophilized | Add 5.0 mL of reagent water, let the calibrators stand for 15–20 minutes at RT. Gently swirl and invert the vials until homogeneous | 3 for lot calibration, 2 for pack recalibration | 2 | human plasma with sodium azide, preservatives and protein stabilizers, | 21 | 7 |  | No |  | No | NA | NA |  | NA | NA | No | Use commercially available Controls | NA | 0.1 ng/dL | 1.3 pmol/L | NP | NP | 0.2 ng/dL | 0.1 ng/dL | NP | 1.3 pmol/L | 2.6 pmol/L | NP | 0.4 ng/dL | 5.7 pmol/L | 4.7 | 6.8 | 5.1 ng/dL | 65.8 pmol/L | 2.1 | 4.1 | 10.7 ng/dL | 138.3 pmol/L | 2.5 | 3.5 | ADVIA Centaur FT4 y = 1.018x - 0.02 ng/dL r = 0.981 | ADVIA Centaur FT4 y = 1.018x - 2.6 pmol/L r = 0.981 | NSI up to 300 mg/dL of hemoglobin | NSI up to 20 mg/dL of bilirubin | NSI up to 1000 mg/dL of triglycerides | NSI up to 3500 ng/mL of biotin | NA | NA | NA | NA | Internal standard manufactured using U.S.P.(United States Pharmacopeia) material | 10995348_EN Rev. 06, 2020-11 | IFU Document | 0.187 | 6.969 | 2.087 | 1.457 | 0.085 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995588 |  | ty02 | 00630414599120 | Global | N/AThyroid | No | Atellica IM Free Thyroxine (FT4) | Atellica IM FT4FT4 | FT4 | 250 | 5 x 50 | Atellica IM CAL A 2- pack SMN 10995500 Atellica IM CAL A 6- pack SMN 10995501 | Atellica IM FT4 MCM (SMN 10995590) | 14 | 21 | acridinium-ester-labeled T4 and biotinylated rabbit polyclonal anti- T4 antibody | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | RT 8 hrs, 2–8°C after 8 hours; ≤ -20°C after 48 hours; Freeze only 1 time | 25 μL | No |  | Competitive immunoassay using direct chemiluminescent technology | 0.1 – 12.0 ng/dL | 1.3 – 154.8 pmol/L | NA | NA | 0.89- 1.76 ng/dL | 11.5 - 22.7 pmol/L | NA | NA | NA | Atellica IM CAL A | 10995500 (2-pack) 10995501 (6-pack) | 28 days recon at 2–8°C, 4 hours recon at RT | Lyophilized | Add 5.0 mL of reagent water, let the calibrators stand for 15–20 minutes at RT. Gently swirl and invert the vials until homogeneous | 3 for lot calibration, 2 for pack recalibration | 2 | human plasma with sodium azide, preservatives and protein stabilizers, | 21 | 7 |  | No |  | No | NA | NA |  | NA | NA | No | Use commercially available Controls | NA | 0.1 ng/dL | 1.3 pmol/L | NP | NP | 0.2 ng/dL | 0.1 ng/dL | NP | 1.3 pmol/L | 2.6 pmol/L | NP | 0.4 ng/dL | 5.7 pmol/L | 4.7 | 6.8 | 5.1 ng/dL | 65.8 pmol/L | 2.1 | 4.1 | 10.7 ng/dL | 138.3 pmol/L | 2.5 | 3.5 | ADVIA Centaur FT4 y = 1.018x - 0.02 ng/dL r = 0.981 | ADVIA Centaur FT4 y = 1.018x - 2.6 pmol/L r = 0.981 | NSI up to 300 mg/dL of hemoglobin | NSI up to 20 mg/dL of bilirubin | NSI up to 1000 mg/dL of triglycerides | NSI up to 3500 ng/mL of biotin | NA | NA | NA | NA | Internal standard manufactured using U.S.P.(United States Pharmacopeia) material | 10995348_EN Rev. 06, 2020-11 | IFU Document | 0.737 | 6.969 | 2.087 | 4.37 | 0.335 | 17.7 | 5.3 | 11.1
Atellica® IM | 10995690 |  | re01 | 00630414600031 | OUS | Reproductive Endocrinology | No | Atellica IM Total hCG (ThCG) | Atellica IM ThCG | ThCG | 90 | 1 x 90 | Atellica IM CAL B | Atellica IM ThCG DIL, Atellica IM ThCG MCM (SMN 10995694) | 10 | 21 | Goat polyclonal anti-hCG antibody and Mouse monoclonal anti-hCG antibody | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | RT 8 hours;2–8°C after 8 hours; ≤ -20°C after 48 hours; Freeze only 1 time | 50 μL | No |  | 2-site sandwich immunoassay using direct chemiluminescent | 2.0–1000.0 mIU/mL | 2.0–1000.0 IU/L |  |  |  |  | non-pregnant females 1.5–4.2 mIU/mL (IU/L) | postmenopausal females 1.8–10.1 mIU/mL (IU/L) |  | Atellica IM CAL B | 10995503 (2-pack) 10995504 (6-pack) | Recon at 2–8°C 28 days Recon at RT 4hrs | Lyophilized | Add 5.0 mL of special reagent water into each vial Let the vials stand for 15–20 min at RT | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of digoxin,follicle-stimulating hormone (FSH), luteinizinghormone (LH), prolactin, human chorionic gonadotropin (hCG), and TSH; equine serum | 34 | 28 |  | No |  | Yes | Atellica IM ThCG DIL | 10995691 - 2 pack10995692 - 6 pack10995693 - vial |  | 2.0-200,000.0 mIU/mL | 2.0-200,000.0 IU/L | No | NA | NA | 1.9 mIU/mL | 1.9 IU/L | NA | NA | 2.8 mIU/mL | 1.5 mIU/mL | 2.8 mIU/mL | 1.5 IU/L | 2.8 IU/L | 2.8 IU/L | 2.6 mIU/mL | 2.6 IU/L | 8.5 | 13.4 | 13.0 mIU/mL | 13.0 IU/L | 2.1 | 4.2 | 791.2 mIU/mL | 791.2 IU/L | 1.6 | 2.2 | ADVIA Centaur Total hCG y = 0.994x + 2.973 mIU/mL r=0.999 | ADVIA Centaur Total hCG y = 0.994x + 2.973 IU/L r=0.999 | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up 40 mg/dL conjugated/unconjugated bilirubin | ≤ 10% change up to 3000 mg/dL Intralipid | No Biotin in assay formulation | 400,000 mIU/mL | 400,000 IU/L |  |  | World Health Organization (WHO) 4th IS 75/589 reference material | 10995426_EN Rev. 03, 2020-02 | IFU Document | 0.205 | 6.969 | 2.087 | 1.457 | 0.093 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995689 |  | re01 | 00630414600024 | OUS | Reproductive Endocrinology | No | Atellica IM Total hCG (ThCG) | Atellica IM ThCG | ThCG | 450 | 5 x 90 | Atellica IM CAL B | Atellica IM ThCG DIL Atellica IM ThCG MCM (SMN 10995694) | 10 | 21 | Goat polyclonal anti-hCG antibody and Mouse monoclonal anti-hCG antibody | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | RT 8 hours;2–8°C after 8 hours; ≤ -20°C after 48 hours; Freeze only 1 time | 50 μL | No |  | 2-site sandwich immunoassay using direct chemiluminometric | 2.0–1000.0 mIU/mL | 2.0–1000.0 IU/L |  |  |  |  | non-pregnant females 1.5–4.2 mIU/mL (IU/L) | postmenopausal females 1.8–10.1 mIU/mL (IU/L) |  | Atellica IM CAL B | 10995503 (2-pack) 10995504 (6-pack) | Recon at 2–8°C 28 days Recon at RT 4hrs | Lyophilized | Add 5.0 mL of special reagent water into each vial Let the vials stand for 15–20 min at RT | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of digoxin,follicle-stimulating hormone (FSH), luteinizinghormone (LH), prolactin, human chorionic gonadotropin (hCG), and TSH; equine serum | 34 | 28 |  | No |  | Yes | Atellica IM ThCG DIL | 10995691 - 2 pack10995692 - 6 pack10995693 - vial |  | 2.0-200,000.0 mIU/mL | 2.0-200,000.0 IU/L | No | NA | NA | 1.9 mIU/mL | 1.9 IU/L | NA | NA | 2.8 mIU/mL | 1.5 mIU/mL | 2.8 mIU/mL | 1.5 IU/L | 2.8 IU/L | 2.8 IU/L | 2.6 mIU/mL | 2.6 IU/L | 8.5 | 13.4 | 13.0 mIU/mL | 13.0 IU/L | 2.1 | 4.2 | 791.2 mIU/mL | 791.2 IU/L | 1.6 | 2.2 | ADVIA Centaur Total hCG y = 0.994x + 2.973 mIU/mL (IU/L) r=0.999 | ADVIA Centaur Total hCG y = 0.994x + 2.973 IU/L r=0.999 | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up 40 mg/dL conjugated/unconjugated bilirubin | ≤ 10% change up to 3000 mg/dL Intralipid | No Biotin in assay formulation | 400,000 mIU/mL | 400,000 IU/L |  |  | World Health Organization (WHO) 4th IS 75/589 reference material | 10995426_EN Rev. 03, 2020-02 | IFU Document | 0.821 | 6.969 | 2.087 | 4.37 | 0.373 | 17.7 | 5.3 | 11.1
Atellica® IM | 10995696 |  | ca02 | 00630414600093 | OUS | Cardiac | No | Atellica IM Troponin I Ultra (TnI‑Ultra) | Atellica IM TnI‑Ultra | TnI_UL | 100 |  | Atellica IM APW3 (SMN 10998580) | Atellica IM Multi‑Diluent 11 (SMN 10995642) Atellica IM TnI‑Ultra MCM (SMN 10995698) | 10 | 28 | 1 Goat polyclonal anti-troponin I antibody / 2 Mouse monoclonal | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | 4 hrs at RT, 24 hrs at 2–8°C, 1 month at ≤ -20°C, Freeze only 1 time | 100 μL | No |  | direct chemiluminescent technology | 0.006–30.000 ng/mL | 0.006–30.000 μg/L |  |  |  |  | 99th percentile 0.02 ng/mL (μg/L) (Plasma) | See IFU for expected results by gender |  | Atellica IM TnI-Ultra Cal | Provided in kit | 4 hrs recon at RT, 24 hrs recon at 2–8°C, 30 days recon at ≤ -20°C; thaw 1 time | Lyophilized | Reconstituted 2.0 mL/vial | 3 for lot calibration, 2 for pack recalibration | 2 | bovine cTnI, goat serum | 66 | 28 |  | No |  | Yes | Atellica IM Multi‑Diluent 11 | 10995642 2-pack 1099564 1-vial |  | 0.006–300.000 ng/mL | 0.006–300.000 μg/L | No | NA | NA | 0.006 ng/ml | 0.006 μg/L | 0.015 ng/ml | 0.015 μg/L | 0.015 ng/mL | 0.007 ng/mL | NP | 0.007 μg/L | 0.015 μg/L | NP | 0.074 ng/mL (plasma) | 0.074 μg/L (plasma) | 3 | 4.7 | 0.533 ng/mL (plasma) | 0.533 μg/L (plasma) | 1.6 | 4.3 | 21.270 ng/mL(plasma) | 21.270 μg/L (plasma) | 1.6 | 2.2 | ADVIA Centaur TnI‑Ultra (Lithium Heparin) y = 1.05x + 0.022 ng/mL r=1.00 | ADVIA Centaur TnI‑Ultra (Lithium Heparin) y = 1.05x + 0.022 μg/L r=1.00 | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 40 mg/dL of conjugated bilirubin, upto 60 mg/dL of unconjugated bilirubin | ≤ 10% change up to 2000 mg/dL of Intralipid | ≤ 10% change in results up to 10 ng/mL. | 7606 ng/mL | 7606 μg/L |  |  | NIST SRM 2921 | 10995428_EN Rev. 05, 2019-08 | IFU Document | 0.323 | 6.969 | 2.48 | 2.913 | 0.147 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995697 |  | ca02 | 00630414600109 | OUS | Cardiac | No | Atellica IM Troponin I Ultra (TnI‑Ultra) | Atellica IM TnI‑Ultra | TnI_UL | 500 |  | Atellica IM APW3 (SMN 10998580) | Atellica IM Multi‑Diluent 11 (SMN 10995642) Atellica IM TnI‑Ultra MCM (SMN 10995698) | 10 | 28 | 1 Goat polyclonal anti-troponin I antibody / 2 Mouse monoclonal | ;#Serum;#Sodium Heparin Plasma;#Heparinized Plasma;#EDTA Plasma;# | 4 hrs at RT, 24 hrs at 2–8°C, 1 month at ≤ -20°C, Freeze only 1 time | 100 μL | No |  | direct chemiluminescent technology | 0.006–30.000 ng/mL | 0.006–30.000 μg/L |  |  |  |  | Combined gender 99th percentile using lithium heparin sample type - 0.02 ng/mL(μg/L) | See IFU for expected results by gender |  | Atellica IM TnI-Ultra Cal | Provided in Kit | 4 hrs recon at RT, 24 hrs recon at 2–8°C, 30 days recon at ≤ -20°C; thaw 1 time | Lyophilized | Reconstituted to 2.0 mL/vial | 3 for lot calibration, 2 for pack recalibration | 2 | bovine cTnI, goat serum | 66 | 28 |  | No |  | Yes | Atellica IM Multi‑Diluent 11 | 10995642 - 2 pack 10995641 - vial |  | 0.006–300.000 ng/mL | 0.006–300.000 μg/L | No | NA | NA | 0.006 ng/ml | 0.006 μg/L | 0.015 ng/ml | 0.015 μg/L | 0.015 ng/mL | 0.007 ng/mL | NP | 0.007 μg/L | 0.015 μg/L | NP | 0.074 ng/mL (plasma) | 0.074 μg/L (plasma) | 3 | 4.7 | 0.533 ng/mL (plasma) | 0.533 μg/L (plasma) | 1.6 | 4.3 | 21.270 ng/mL (plasma) | 21.270 μg/L (plasma) | 1.6 | 2.2 | ADVIA Centaur TnI‑Ultra (Lithium Heparin) y = 1.05x + 0.022 ng/mL r=1.00 | ADVIA Centaur TnI‑Ultra (Lithium Heparin) y = 1.05x + 0.022 μg/L r=1.00 | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 40 mg/dL of conjugated bilirubin, up to 60 mg/dL of unconjugated bilirubin | ≤ 10% change up to 2000 mg/dL of Intralipid | ≤ 10% change in results up to 10 ng/mL. | 7606 ng/mL | 7606 μg/L |  |  | NIST SRM 2921 | 10995428_EN Rev. 05, 2019-08 | IFU Document | 1.08 | 6.969 | 2.126 | 6.811 | 0.491 | 17.7 | 5.4 | 17.3
Atellica® IM | 10995714 |  | an01 | 00630414600277 | Global | N/AAnemia | No | Atellica IM Vitamin B12 (VB12) | VB12 | VB12 | 100 | 1 X100 | Atellica IM VB12 DTT/REL (SMN 0995718) Atellica IM T3/T4/VB12 ANC (2pk) (SMN 10995682) Atellica IM T3/T4/VB12 ANC (6pk) (SMN 10995683) | Atellica IM VB12 MCM (SMN 10995717) | 14 | 18 | purified hog intrinsic factor | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | 8 hours at RT 2–8°C after 8 hours; ≤ -20°C after 48 hours | 100 μL | No |  | competitive immunoassay using direct chemiluminescent technology | 45–2000 pg/mL | 33–1476 pmol/L | 45 – 20,000 pg/mL | 33 – 14,760 pmol/L | 211 to 911 pg/mL | 156 to 672 pmol/L | NA | NA | NA | Atellica IM CAL C | 10995506 (2-pack) 10995507 (6-pack) | 28 days recon at 2–8°C 4 hours recon at RT | Lyophilized | Add 5.0 mL of reagent water, let the calibrators stand for 15–20 minutes at RT | 3 for lot calibration, 2 for pack recalibration | 2 | buffered HSA with sodium azide, preservatives and protein stabilizers, | 30 | 18 |  | No |  | Yes | Atellica IM VB12 DIL | 10995716 |  | 45–20000 pg/mL | 33-14760 pmol/L | No | NA | NA | 38 pg/mL | 28 pmol/L | NA | NA | 54 pg/mL | 38 pg/mL | NP | 28 pmol/L | 40 pmol/L | NP | 146 pg/mL (LI heparin) | 108 pmol/L (LI heparin) | 2.5 | 9.2 | 368 pg/mL (serum) | 271 pmol/L (serum) | 2 | 6.4 | 486 pg/mL (serum) | 358 pmol/L (serum) | 2.3 | 5.6 | ADVIA Centaur VB12 y = 1.00x + 7 pg/mL r=0.994 | ADVIA Centaur VB12 y = 1.00x + 5 pmol/L r=0.994 | NSI up to 150 mg/dL of hemoglobin | NSI up to 20 mg/dL of bilirubin | NSI up to 3000 mg/dL of triglycerides | No Biotin in assay formulation | NA | NA | N/A | N/A | Traceable to an internal standard manufactured using U.S.P. (United States Pharmacopeia) material | 10995437_EN Rev. 02, 2019-08 | IFU Document | 0.235 | 6.969 | 2.087 | 1.457 | 0.107 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995715 |  | an01 | 00630414600284 | Global | N/AAnemia | No | Atellica IM Vitamin B12 (VB12) | VB12 | VB12 | 500 | 5 x 100 | Atellica IM VB12 DTT/REL (SMN 0995718) Atellica IM T3/T4/VB12 ANC (2pk) (SMN 10995682) Atellica IM T3/T4/VB12 ANC (6pk) (SMN 10995683) | Atellica IM VB12 MCM (SMN 10995717) | 14 | 18 | purified hog intrinsic factor | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | 8 hours at RT 2–8°C after 8 hours; ≤ -20°C after 48 hours | 100 μL | No |  | competitive immunoassay using direct chemiluminescent technology | 45–2000 pg/mL | 33–1476 pmol/L | 45 – 20,000 pg/mL | 33 – 14,760 pmol/L | 211 to 911 pg/mL | 156 to 672 pmol/L | NA | NA | NA | Atellica IM CAL C | 10995506 (2-pack) 10995507 (6-pack) | 28 days recon at 2–8°C, 4 hours recon at RT | Lyophilized | Add 5.0 mL of reagent water, let the calibrators stand for 15–20 minutes at RT | 3 for lot calibration, 2 for pack recalibration | 2 | buffered HSA with sodium azide, preservatives and protein stabilizers, | 30 | 18 |  | No |  | Yes | Atellica IM VB12 DIL | 10995716 |  | 45–20000 pg/mL | 33-14760 pmol/L | No | NA | NA | 38 pg/mL | 28 pmol/L | NA | NA | 54 pg/mL | 38 pg/mL | NP | 28 pmol/L | 40 pmol/L | NP | 146 pg/mL (LI heparin) | 108 pmol/L (LI heparin) | 2.5 | 9.2 | 368 pg/mL (serum) | 271 pmol/L (serum) | 2 | 6.4 | 486 pg/mL (serum) | 358 pmol/L (serum) | 2.3 | 5.6 | ADVIA Centaur VB12 y = 1.00x + 7 pg/mL r=0.994 | ADVIA Centaur VB12 y = 1.00x + 5 pmol/L r=0.994 | NSI up to 150 mg/dL of hemoglobin | NSI up to 20 mg/dL of bilirubin | NSI up to 3000 mg/dL of triglycerides | No Biotin in assay formulation | NA | NA | N/A | N/A | Traceable to an internal standard manufactured using U.S.P. (United States Pharmacopeia) material | 10995437_EN Rev. 02 2019-08 | IFU Document | 0.957 | 6.969 | 2.087 | 4.37 | 0.435 | 17.7 | 5.3 | 11.1
Atellica® IM | 10995444 |  | id01 | 00630414597737 | Global | Hepatitis | No | Atellica IM Hepatitis A IgM (aHAVM) | Atellica IM aHAVMaHAVM | aHAVM | 100 | 1 x 100 | Atellica IM Multi-Diluent 2 (SMN 10995644) | Atellica IM aHAVM QC | 46 | 41 | Mouse monoclonal anti-HAV antibody F(ab’)2 fragment | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#EDTA Plasma;# | RT not recommended2–8°C for up to 7 days; ≤ -20°C for up to 180 days; | 20 μL | No |  | IgM-capture immunoassay using a 2-pass format | 0.02-7.00 S/CO | 0.02-7.00 S/CO | NA | NA | NA | NA | Nonreactive <0.8 S/CO | Reactive >1.20 S/CO | Equivocal | Atellica IM aHAVM Cal | Provided in kit | at 2–8°C until expiration, on board at RT 8hrs | Liquid | N/A | 3 for lot calibration, 2 for pack recalibration | 2 | Processed human plasma positive for IgM antibodies to HAV; preservatives | 91 | 28 |  | No |  | No | NA | NA |  | NA | NA | Yes | Atellica IM aHAVM QC | 10995445 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 0.49 S/CO (plasma) | 0.49 S/CO (plasma) | 3.9 | 5.5 | 2.98 S/CO (plasma) | 2.98 S/CO (plasma) | 3.3 | 5 | 5.23 S/CO (plasma) | 5.23 S/CO (plasma) | 4.3 | 5.2 | NA | NA | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 60 mg/dL of conjugated bilirubin, and up to 40 mg/dL of unconjugated bilirubin | ≤ 10% change upto 3000 mg/dL of triglycerides | Insignificant effect on the assay up to 500 ng/mL | NA | NA |  |  | Standardization based upon relative clinical agreement with commercially available anti-HAV IgM assays | 10995271_EN Rev. 03, 2019-07 (OUS) 10995272_EN Rev. 03, 2019-06 (US) | IFU Document | 0.359 | 6.969 | 2.48 | 2.913 | 0.163 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995446 |  | id02 | 00630414597751 | OUS | Hepatitis | No | Atellica IM Hepatitis A Total (aHAVT) | Atellica IM aHAVTaHAVT | aHAVT | 100 | 1 x 100 | Atellica IM APW1 (SMN 10995458) | Atellica IM aHAVT QC | 54 | 41 | Mouse monoclonal anti-HAV (~1.0 μg/mL) labeled with acridinium ester and biotinylated mouse monoclonal anti- HAV Fab fragment | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#EDTA Plasma;# | RT not recommended 2–8°C for up to 7 days; ≤ -20°C for up to 30 days | 20 μL | No |  | competitive immunoassay using direct chemiluminescent technology. | 0.00–100.00 mIU/mL | 0.00–100.00 mIU/mL | NA | NA | NA | NA | Nonreactive < 20 mIU/mL | Reactive > 20 mIU/mL |  | Atellica IM aHAVT Cal | Provided in kit | Unopened at 2–8°C until expiration Opened at 2–8°C 60 days on board at RT 8hrs | Liquid | N/A | 3 for lot calibration, 2 for pack recalibration | 2 | Processed human plasma positive for anti-HAV antibodies | 41 | 28 |  | No |  | No | NA | NA |  | NA | NA | Yes | Atellica IM aHAVT QC | 10995448 | NA | NA | NA | NA | NA | N/A | N/A | NA | N/A | NA | 16.17 mIU/mL (plasma) | 16.17 mIU/mL (plasma) | 4.9 | 6.6 | 36.75 mIU/mL (plasma) | 36.75 mIU/mL (plasma) | 2.5 | 3.5 | 80.47 mIU/mL (plasma) | 80.47 mIU/mL (plasma) | 1.6 | 2.4 | NA | NA | NSI up to 500 mg/dL of hemoglobin | NSI - up to 60 mg/dL of conjugated bilirubin, up to 40 mg/dL of unconjugated bilirubin | NSI up to 3000 mg/dL of triglycerides | Insignificant effect on the assay up to 25 ng/mL | NA | NA |  |  | WHO Second International Standard for Anti-Hepatitis A Immunoglobulin (97/646) | 10995273_EN Rev. 03, 2019-07 | IFU Document | 0.387 | 6.969 | 2.48 | 2.913 | 0.176 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995451 |  | id04 | 00630414597812 | OUS | Hepatitis | No | Atellica IM Hepatitis B e antibody (aHBe) | Atellica IM aHBeaHBe | aHBe | 100 | 1 x 100 | NA | Atellica IM aHBe QC | 52 | 26 | Acridinium-ester-labeled mouse monoclonal anti-HBe ; bovine serum albumin; mouse IgG;Biotinylated mouse monoclonal anti-HBe | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#EDTA Plasma;# | RT for up to 24 hrs; 2–8°C for up to 7 days; ≤ -20°C for longer storage | 100 μL | No |  | competitive/neutralizing immunoassay using a 2‑step, singlewash format. | 0.05–4.50 Index | 0.05–4.50 Index | NA | NA | NA | NA | Nonreactive < 0.80 Index | Reactive > 1.20 Index | Equivocal ≥0.80 Index and <1.20 Index | Atellica IM aHBe Cal | Provided in kit | at 2–8°C until expirationon board at RT 8hrs | Liquid | N/A | 3 for lot calibration, 2 for pack recalibration | 2 | Processed human plasma positive for antibodies to HBe antigen | 51 | 21 |  | No |  | No | No | No |  | No | No | Yes | Atellica IM aHBe QC | 10995452 | NA | NA | NA | NA | NA | N/A | N/A | NA | N/A | NA | 0.46 Index (serum) | 0.46 Index (serum) | 4 | 7.7 | 1.41 Index (serum) | 1.41 Index (serum) | 2.3 | 3.6 | 3.61 Index (serum) | 3.61 Index (serum) | 1.2 | 2.1 | NA | NA | ≤ 20% change up to 500 mg/dL of hemoglobin | ≤ 20% change up to 40 mg/dL of conjugated bilirubin, up to 20 mg/dL of unconjugated bilirubin | ≤ 20% change up to 500 mg/dL of triglycerides, and up to 400 mg/dL of cholesterol | Insignificant effect on the assay up to 1500 ng/mL | NA | NA |  |  | Standarization based upon relative clinical agreement withcommercially available anti‑HBe assays. | 10995276_EN Rev. 03, 2019-07 | IFU Document | 0.251 | 6.969 | 2.48 | 2.913 | 0.114 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995481 |  | on03 | 00630414598086 | Global | Oncology | No | Atellica IM CA 125II (CA 125II) | Atellica IM CA 125II | CA125 | 100 | 1x100 | Atellica IM CA 125II CAL | Atellica IM Multi‑Diluent 1 Atellica IM CA 125II MCM (SMN 10995484) | 28 | 28 | Mouse monoclonal anti-M11 antibody (~0.15 μg/mL) | ;#Serum;# | RT: 8 hours; 2–8°C after 8 hours; ≤ -20°C after 24 hours | 50 μL | No |  | 2-site sandwich immunoassay | 2.0–600.0 U/mL | NA |  |  |  |  | <35 U/mL |  |  | Atellica IM CA 125II CAL, | 10995483 | 7 days recon at 2–8°C 8 hours recon at RT | Lyophilized | 2.0 mL of special reagent water. Let the vials stand for 15–20 minutes at room temperature | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of CA 125 | 91 | 28 |  | No |  | Yes | Atellica IM Multi‑Diluent 1 | 10995637 - 2 pack;10995638 - 6 pack;10995639 - vial |  | 2.0–12000.0 U/mL | NA | No | NA | NA | 1.2 U/mL | NA | NP | NA | 3.0 U/mL | 1.1 U/mL | NP | NA | NA | NA | 15.1 U/mL | NA | 1.5 | 2.8 | 172 U/mL | NA | 1.4 | 2.8 | 519 U/mL | NA | 1.4 | 3.2 | ADVIA Centaur CA 125II y = 0.99x + 1.7 U/mL r=1.00 | NA | ≤ 10% change up to 1000 mg/dL of hemoglobin | ≤10% change up 20 mg/dL conjugated/unconjugated bilirubin | ≤10% change upto 900 mg/dL of triglycerides | No Biotin in assay formulation | 70,000 U/mL | NA |  |  | traceable to an internal standard manufactured using highlypurified material | 10995283_EN Rev. 03, 2019-08 | IFU Document | 0.22 | 6.969 | 2.087 | 1.457 | 0.1 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995482 |  | on03 | 00630414598093 | Global | Oncology | No | Atellica IM CA 125II (CA 125II) | Atellica IM CA 125II | CA125 | 500 | 5 x 100 | Atellica IM CA 125II CAL | Atellica IM Multi‑Diluent 1, Atellica IM CA 125II MCM (SMN 10995484) | 28 | 28 | Mouse monoclonal anti-M11 antibody (~0.15 μg/mL) | ;#Serum;# | RT 8 hours; 2–8°C after 8 hours; ≤ -20°C 24hours | 50 μL | No |  | 2-site sandwich immunoassay | 2.0–600.0 U/mL | NA |  |  |  |  | <35 U/mL |  |  | Atellica IM CA 125II CAL, | 10995484 | 7 days recon at 2–8°C 8 hours recon at RT | Lyophilized | 2.0 mL of special reagent water. Let the vials stand for 15–20 minutes at room temperature | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of CA 125 | 91 | 28 |  | No |  | Yes | Atellica IM Multi‑Diluent 1 | 10995637 - 2 pack;10995638 - 6 pack;10995639 - vial |  | 2.0–12000.0 U/mL | NA | No | NA | NA | 1.2 U/mL | NA | NP | NA | 3.0 U/mL | 1.1 U/mL | NP | NA | NA | NA | 15.1 U/mL | NA | 1.5 | 2.8 | 172 U/mL | NA | 1.4 | 2.8 | 519 U/mL | NA | 1.4 | 3.2 | ADVIA Centaur CA 125II y = 0.99x + 1.7 U/mL r=1.01 | NA | ≤ 10% change up to 1000 mg/dL of hemoglobin | ≤10% change up 20 mg/dL conjugated/unconjugated bilirubin | ≤10% change upto 900 mg/dL of triglycerides | No Biotin in assay formulation | 70,000 U/mL | NA |  |  | traceable to an internal standard manufactured using highly purified material | 10995283_EN Rev. 03, 2019-08 | IFU Document | 0.9 | 6.969 | 2.087 | 4.37 | 0.409 | 17.7 | 5.3 | 11.1
Atellica® IM | 10995485 |  | on04 | 00630414598123 | Global | Oncology | No | Atellica IM CA 15‑3 (CA 15‑3) | Atellica IM CA 15‑3 | CA15_3 | 100 | 1 x 100 | Atellica IM CA 15‑3 CAL | Atellica IM Multi‑Diluent 1 Atellica IM CA 15‑3 MCM (SMN 10995488) | 28 | 28 | Mouse monoclonal anti-DF3 antibody (~ 2.0 μg/mL) | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | 2–8°C after 8 hours;≤ -20°C if the sample is not assayed within 24 hours | 20 μL | No |  | 2-step sandwich immunoassay | 0.5–200.0 U/mL | NA |  |  |  |  | <32.4 U/mL |  |  | Atellica IM CA 15‑3 CAL | 10995487 | 28 days recon at 2–8°C 8 hours recon at RT | Lyophilized | 2.0 mL of special reagent water. Let the vials stand for 15–20 minutes at room temperature | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of CA 15‑3 (human) | 91 | 28 |  | No |  | Yes | Atellica IM Multi‑Diluent 1 - | 10995637 - 2 pack;10995638 - 6 pack;10995639 - vial |  | 0.5–4000.0 U/mL | NA | No | NA | NA | 0.0 U/mL | NA | NP | NA | 1.9 U/mL | 0.0 U/mL | NP | NA | NA | NA | 9.6 U/mL | NA | 3.1 | 6.2 | 65.9 U/mL | NA | 1.6 | 3.7 | 165.8 U/mL | NA | 1.4 | 3.4 | ADVIA Centaur CA 15‑3 y = 1.07x + 0.146 U/mL r=0.99 | NA | ≤ 10% change up to 1000 mg/dL of hemoglobin | ≤10% change up 25 mg/dL conjugated bilirubin | ≤10% change upto 900 mg/dL of triglycerides | No Biotin in assay formulation | NA | NA |  |  | traceable to an internal standard manufactured using highly purified material | 10995284_EN Rev. 03, 2019-08 | IFU Document | 0.224 | 6.969 | 2.087 | 1.457 | 0.102 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995486 |  | on04 | 00630414598130 | Global | Oncology | No | Atellica IM CA 15‑3 (CA 15‑3) | Atellica IM CA 15‑3 | CA15_3 | 500 | 5 x 100 | Atellica IM CA 15‑3 CAL | Atellica IM Multi‑Diluent 1, Atellica IM CA 15‑3 MCM | 28 | 28 | Mouse monoclonal anti-DF3 antibody (~ 2.0 μg/mL) | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | 2–8°C after 8 hours;≤ -20°C if the sample is not assayed within 24 hours | 20 μL | No |  | 2-step sandwich immunoassay | 0.5–200.0 U/mL | NA |  |  |  |  | <32.4 U/mL |  |  | Atellica IM CA 15‑3 CAL | 10995487 | 28 days recon at 2–8°C 8 hours recon at RT | Lyophilized | 2.0 mL of special reagent water. Let the vials stand for 15–20 minutes at room temperature | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of CA 15‑3 (human) | 91 | 28 |  | No |  | Yes | Atellica IM Multi‑Diluent 1 - | 10995637 - 2 pack;10995638 - 6 pack;10995639 - vial |  | 0.5–4000.0 U/mL | NA | No | NA | NA | 0.0 U/mL | NA | NP | NA | 1.9 U/mL | 0.0 U/mL | NP | NA | NA | NA | 9.6 U/mL | NA | 3.1 | 6.2 | 65.9 U/mL | NA | 1.6 | 3.7 | 165.8 U/mL | NA | 1.4 | 3 | ADVIA Centaur CA 15‑3 y = 1.07x + 0.146 U/mL r=0.99 | NA | ≤ 10% change up to 1000 mg/dL of hemoglobin | ≤10% change up 25 mg/dL conjugated bilirubin | ≤10% change upto 900 mg/dL of triglycerides | No Biotin in assay formulation | NA | NA |  |  | traceable to an internal standard manufactured using highly purified material | 10995284_EN Rev. 03, 2019-08 | IFU Document | 0.913 | 6.969 | 2.087 | 4.37 | 0.415 | 17.7 | 5.3 | 11.1
Atellica® IM | 10995662 |  | on01 | 00630414599809 | OUS | Oncology | No | Atellica IM Prostate‑Specific Antigen (PSA) | Atellica IM PSA | PSA | 100 | 1x100 | Atellica IM CAL Q | Multi‑Diluent 2, Atellica IM PSA MCM | 10 | 28 | Goat polyclonal anti‑PSA antibody (~77 ng/mL) | ;#Serum;# | 8hrs at RT, 48 hrs at 2-8°C, after 48 hrs at ≤ -20°C | 35 μL | No |  | 2‑site sandwich immunoassay | 0.01–100.00 ng/mL | 0.01–100.00 μg/L |  |  |  |  | See IFU |  |  | Atellica IM CAL Q | 10995517 (2pk) 10995518 (6pk) | 21 days recon at 2–8°C 8 hours recon at RT | Lyophilized | 2.0 mL of special reagent water. Let the vials stand for 15–20 minutes at room temperature | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of PSA | 29 | 28 |  | No |  | Yes | Atellica IM Multi‑Diluent 2 | 10995644 |  | 0.01–10000.00 ng/mL | 0.01–10000.00 μg/L | No | NA | NA | 0.01 ng/mL | 0.01 μg/L | NP | NP | 0.02 ng/mL | 0.01 ng/mL | 0.02 ng/mL | 0.01 μg/L | 0.02 μg/L | 0.02 μg/L | 0.06 ng/mL | 0.06 μg/L | 6.2 | 8.5 | 19.63 ng/mL | 19.63 μg/L | 1.8 | 2.8 | 78.67 ng/mL | 78.67 μg/L | 2.2 | 4 | ADVIA Centaur PSA y = 0.93x + 0.238 ng/mL r=.996 | ADVIA Centaur PSA y = 0.93x + 0.238 μg/L r=.996 | ≤ 5% change up to 500 mg/dL of hemoglobin | ≤ 5% change up to 40 mg/dL of bilirubin | ≤ 5% change up to 1000 mg/dL of triglycerides | No Biotin in assay formulation | 2,500 ng/mL | 2,500 μg/L |  |  | (WHO) International Standard (96/670) | 10995416_EN Rev. 03, 2019-09 | IFU Document | 0.222 | 6.969 | 2.087 | 1.457 | 0.101 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995663 |  | on01 | 00630414599816 | OUS | Oncology | No | Atellica IM Prostate‑Specific Antigen (PSA) | Atellica IM PSA | PSA | 500 | 5 x 100 | Atellica IM CAL Q | Multi‑Diluent 2, Atellica IM PSA MCM | 10 | 28 | Goat polyclonal anti‑PSA antibody (~77 ng/mL) | ;#Serum;# | 8hrs at RT, 48 hrs at 2-8°C, after 48 hrs at ≤ -20°C | 35 μL | No |  | 2‑site sandwich immunoassay | 0.01–100.00 ng/mL | 0.01–100.00 μg/L |  |  |  |  | See IFU |  |  | Atellica IM CAL Q | 10995517 (2pk) 10995518 (6pk) | 21 days recon at 2–8°C 8 hours recon at RT | Lyophilized | 2.0 mL of special reagent water. Let the vials stand for 15–20 minutes at room temperature | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of PSA | 29 | 28 |  | No |  | Yes | Atellica IM Multi‑Diluent 3 | 10995645 |  | 0.01–10000.00 ng/mL | 0.01–10000.00 μg/L | No | NA | NA | 0.01 ng/mL | 0.01 μg/L | NP | NP | 0.02 ng/mL | 0.01 ng/mL | 0.02 ng/mL | 0.01 μg/L | 0.02 μg/L | 0.02 μg/L | 0.06 ng/mL | 0.06 μg/L | 6.2 | 8.5 | 19.63 ng/mL | 19.63 μg/L | 1.8 | 2.8 | 78.67 ng/mL | 78.67 μg/L | 2.2 | 4 | ADVIA Centaur PSA y = 0.93x + 0.238 ng/mL r=.996 | ADVIA Centaur PSA y = 0.93x + 0.238 μg/L r=.996 | ≤ 5% change up to 500 mg/dL of hemoglobin | ≤ 5% change up to 40 mg/dL of bilirubin | ≤ 5% change up to 1000 mg/dL of triglycerides | No Biotin in assay formulation | 2,500 ng/mL | 2,500 μg/L |  |  | (WHO) International Standard (96/670) | 10995416_EN Rev. 03, 2019-09 | IFU Document | 0.9 | 6.969 | 2.087 | 4.37 | 0.409 | 17.7 | 5.3 | 11.1
Atellica® IM | 10995527 |  | id07 | 00630414598543 | OUS | HIV | No | Atellica IM HIV Ag/Ab Combo (CHIV) | Atellica IM CHIVCHIV | CHIV | 100 | 1x100 | Atellica IM PW3 | Atellica IM CHIV QC | 39 | 42 | recombinant HIV antigens/Streptavidin-coated paramagnetic microparticles preformed with biotinylated HIV antigens, recombinant HIV ag labelled with AE | ;#Serum;#EDTA Plasma;# | RT not recommended; at 2–8°C up to 14 days; at ≤ -20°C up to 8 months | 100 μL | No |  | antigen-bridging microparticle chemiluminometric immunoassay | 0.05–12.00 Index | 0.05–12.00 Index | N/A | N/A | see IFU | see IFU | see IFU | see IFU | see IFU | IM CHIV CAL | Provided in Kit | at 2–8°C until expiration 8 hours at RT | Liquid | NA | 3 for lot calibration, 2 for pack recalibration | 2 | Processed human plasma negative for antibodies to HIV spiked with antibodies to HIV‑1 | 24 | 21 |  |  |  | No | NA | N/A |  | N/A | N/A | Yes | Atellica IM CHIV QC | 11200769 | ≤ 2.0 IU/mL | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 1.03 Index | 1.03 Index | 2.2 | 3 | 2.81 Index | 2.81 Index | 1.4 | 2.2 | 5.54 Index | 5.54 Index | 1.6 | 2.8 | N/A | N/A | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 40 mg/dL of conjugated bilirubin, and up to 40 mg/dL of unconjugated bilirubin | ≤10% change upto 1000 mg/dL of triglycerides | Insignificant effect on the assay&nbsp; up to 3500 ng/mL | N/A | N/A | N/A | N/A | Standardization is based upon relative clinical agreement with commercially available anti-HIV combo assays | 10995315_EN Rev. 04, 2019-07 | IFU Document | 0.264 | 6.969 | 2.48 | 2.913 | 0.12 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995597 |  | id08 | 00630414599205 | OUS | Hepatitis | No | Atellica IM Anti‑Hepatitis B core Total (HBcT) | Atellica IM HBcTHBcT | HBcT | 200 | 1 x 200 | Atellica IM APW1; Atellica IM PW3 | Atellica IM HBcT QC | 46 | 28 | Recombinant hepatitis B core antigen | ;#Serum;#EDTA Plasma;# | RT not recommended; 2–8°C up to 7 days; ≤ -20°C for longer storage | 50 μL | No |  | 2‑wash antigen sandwich immunoassay | NA | NA | NA | NA | NA | NA | Nonreactive < 0.50 Index | Reactive ≥ 0.50 Index |  | Atellica IM HBcT CAL | Provided in Kit | at 2–8°C until expiration; RT at 8hrs on board | Liquid | N/A | 3 for lot calibration, 2 for pack recalibration | 2 | Processed human plasma positive for HBc antibodies | 33 | 14 |  |  |  | No | N/A | N/A |  | N/A | N/A | Yes | Atellica IM HBcT QC | 10995598 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 0.38 Index | 0.38 Index | 4.8 | 6.9 | 2.07 Index | 2.07 Index | 4.8 | 7.7 | 5.27 Index | 5.27 Index | 7.2 | 9.5 | NA | NA | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 60 mg/dL of conjugated bilirubin; 40 mg/dL of unconjugated bilirubin | ≤ 10% change up to 1000 mg/dL of intralipids | Insignificant effect on the assay up to 1500 ng/mL | N/A | N/A |  |  | based upon the relative clinical agreement with commercially available anti-HBc Total assays | 10995353_EN Rev. 03, 2019-07 | IFU Document | 0.394 | 6.969 | 2.48 | 2.913 | 0.179 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995449 |  | id03 | 00630414597799 | OUS | Hepatitis | No | Atellica IM Hepatitis B core IgM (aHBcM) | Atellica IM aHBcMaHBcM | aHBcM | 100 | 1 x 100 | N/A | Atellica IM aHBcM QC | 46 | 41 | monoclonal mouse anti-HBc | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#EDTA Plasma;# | RT not recommended 2–8°C up to 7 days 7 days at ≤ -20°C | 15 μL | No |  | IgM capture immunoassay using a 2-step format | 0.05–9.00 Index | 0.05–9.00 Index | NA | NA | NA | NA | Negative <0.80 Index | Equivocal 0.80-0.99 Index | Positive ≥ 1.0 Index | Atellica IM aHBcM CAL | Provided in Kit | at 2–8°C until expiration RT at 8hrs on board | Liquid | N/A | 3 for lot calibration, 2 for pack recalibration | 2 | processed human plasma positive for IgM antibodies to HBc antigen with preservatives | 92 | 27 |  |  |  | No | N/A | N/A |  | N/A | N/A | Yes | Atellica IM aHBcM QC | 10995450 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 0.39 Index | 0.39 Index | 1.01 | 1.55 | 2.11 Index | 2.11 Index | 2.18 | 2.87 | 7,86 | 7,86 | 5.09 | 5.99 | NA | NA | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 60 mg/dL of conjugated bilirubin 40 mg/dL of unconjugated bilirubin | ≤ 10% change up to 1000 mg/dL of triglycerides | Insignificant effect on the assay up to 150 ng/mL | N/A | N/A |  |  | assay standardization is based upon relative clinical agreement with commercially available anti-HBc IgM assays | 10995274_EN Rev. 04, 2019-07 | IFU Document | 0.374 | 6.969 | 2.48 | 2.913 | 0.17 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995604 |  | id05 | 00630414599274 | OUS | Hepatitis | No | Atellica IM Hepatitis B surface Antigen II (HBsII) | Atellica IM HBsIIHBsII | HBsII | 200 | 1 x 200 | Atellica IM PW3 (SMN 10995666) | Atellica IM HBsII QC | 26 | 60 |  | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#EDTA Plasma;#Sodium Citrate Plasma;# | RT not recommended; 2–8°C up to 7 days ; ≤ -20°C for longer storage | 100 μL | No |  |  | 0.1–1000.0 Index | 0.1–1000.0 Index | NA | NA | NA | NA | Non Reactive <1.0 Index | Reactive ≥ 1.00 Index |  | Atellica IM HBsII CAL | Provided in Kit | at 2–8°C until expiration RT at 8hrs on board | Liquid | N/A | 3 for lot calibration, 2 for pack recalibration | 2 |  | 48 | 21 |  |  |  | No | N/A | N/A |  | N/A | N/A | Yes | Atellica IM HBsII QC | 10995605 | 0.018 IU/mL | NA | NA | NA | NA | NA | NA | NA | NA | NA | 0.9 Index | 0.9 Index | 7.2 | 7.5 | 50.2 Index | 50.2 Index | 1.9 | 2.8 | 811.6 Index | 811.6 Index | 1.5 | 2.2 | N/A | NA | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 40 mg/dL of conjugated bilirubin, and up to 40 mg/dL of unconjugated bilirubin | ≤ 10% change up to 1000 mg/dL of triglycerides | ≤ 10% change up to 10 ng/mL | 3.36 mg/mL | N/A |  |  | based upon the relative clinical agreement with commercially available HBsAg assays | 10995358_EN Rev. 03, 2019-07 | IFU Document | 0.394 | 6.969 | 2.48 | 2.913 | 0.179 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995603 |  | id10 | 00630414599267 | OUS | Hepatitis | No | Atellica IM Hepatitis B surface Antigen II Confirmatory (HBsII Conf) | Atellica IM HBsII ConfConf (HBsII) | Conf (HBsII) | 98 | 1 x 98 | NA | Atellica IM HBsII QC | 53 | 41 | human plasma positive for HBs antibodies | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#EDTA Plasma;#Citrate Plasma;# | RT: 24 hrs 2-8C: 7 days, -20C:for longer storage | 200 μL | No |  |  | NA | NA | NA | NA | NA | NA | Not confirmed | Confirmed |  | Atellica IM HBsII CAL | Provided in Kit | at 2–8°C until expiration, RT at 8hrs on board | Liquid | N/A | 3 for lot calibration, 2 for pack recalibration | 2 |  | 43 | 21 |  |  |  | No | N/A | N/A |  | N/A | N/A | Yes | Atellica IM HBsII QC | 10995605 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA |  |  | NA | NA |  |  | NA | NA |  |  | NA | NA | NA | NA | NA | NA | N/A | N/A |  |  | N/A | 10995268_EN Rev. 03, 2019-10 | IFU Document | 0.101 | 4.37 | 2.48 | 2.52 | 0.046 | 11.1 | 6.3 | 6.4
Atellica® IM | 10995569 |  | an03 | 00630414598932 | OUS | N/AAnemia | No | Atellica IM Ferritin (Fer) | Atellica IM FerFer | Fer | 90 | 1 x 90 | N/A | Atellica IM Fer MCM (SMN 10995570) | 14 | 28 | Goat polyclonal and mouse monoclonal anti-ferritin antibody. | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | up to 8 hrs at RT; 48 hrs at 2–8°C. Freeze samples at ≤ -20°C if the sample is not assayed within 48 hours. | 10 μL | No |  | 2-site sandwich immunoassay using direct chemiluminometric | 0.5–1650.0 ng/mL | 1.1–3630.0 pmol/L | 0.5 – 16,500 ng/mL | 1.1 - 36,300 pmol/L | Normal Males 22–322 ng/mL; Normal Females 10–291 ng/mL | Normal Males 48–708 pmol/L; Normal Females 22–640 pmol/L |  |  |  | Atellica IM Calibrator C (CAL C) | 10995506 (2‑pack) 10995507 (6‑pack) | 4 hours at RT; 28 days at 2–8°C | Lyophilized | Reconstitute with 5.0 mL of special reagent water. | 3 for lot calibration, 2 for pack recalibration | 2 | Vitamin B12 and ferrtin; buffered HSA; protein stabilizers; sodium azide | 50 | 28 |  |  |  | Yes | Atellica IM Multi‑Diluent 1 | 10995637 (2‑pack) 10995638 (6‑pack) 10995639 (vial) |  | 0.5 – 16,500 ng/mL | 1.1 - 36,300 pmol/L | No | NA | NA | 0.0 ng/mL | 0.0 pmol/L | NA | NA | 0.7 ng/mL | 0.3 ng/mL | 0.9 ng/mL | 0.7 pmol/L | 1.5 pmol/L | 2.0 pmol/L | 8.2 ng/mL | 18.0 pmol/L | 1.8 | 5.4 | 487.1 ng/mL | 1071.6 pmol/L | 1.6 | 4.2 | 1453.6 ng/mL | 3197.9 pmol/L | 3.4 | 6.3 | ADVIA Centaur FER y = 1.03x - 0.6 ng/mL r= 1.0 n=106 | ADVIA Centaur FER y = 1.03x - 1.3 pmol/L r=1.0 | ≤ 10% change up to 900 mg/dL (0.56 mmol/L) of hemoglobin | ≤ 10% change up to 60 mg/dL conjugated bilirubin (1021 μmol/L) 60 mg/dL unconjugated bilirubin (1021 μmol/L) | ≤ 10% change up to 2000 mg/dL (22.6 mmol/L) | No Biotin in assay formulation | 80,000 ng/mL | 176,000 pmol/L | 80,000 ng/mL | 176,000 pmol/L | Traceable to World Health Organization 2nd International Standard (WHO 80/578) | 10995343_EN Rev. 03, 2019-07 | IFU Document | 0.229 | 6.969 | 2.087 | 1.457 | 0.104 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995568 |  | an03 | 00630414598925 | OUS | Anemia | No | Atellica IM Ferritin (Fer) | Atellica IM FerFer | Fer | 450 | 5 x 90 | N/A | Atellica IM Fer MCM (SMN 10995570) | 14 | 28 | Goat polyclonal and mouse monoclonal anti-ferritin antibody. | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | up to 8 hrs at RT; 48 hrs at 2–8°C. Freeze samples at ≤ -20°C if the sample is not assayed within 48 hours. | 10 μL | No |  | 2-site sandwich immunoassay using direct chemiluminometric | 0.5–1650.0 ng/mL | 1.1–3630.0 pmol/L | 0.5 – 16,500 ng/mL | 1.1 - 36,300 pmol/L | Normal Males 22–322 ng/mL; Normal Females 10–291 ng/mL | Normal Males 48–708 pmol/L; Normal Females 22–640 pmol/L |  |  |  | Atellica IM Calibrator C (CAL C) | 10995506 (2‑pack) 10995507 (6‑pack) | 4 hours at RT; 28 days at 2–8°C | Lyophilized | Reconstitute with 5.0 mL of special reagent water | 3 for lot calibration, 2 for pack recalibration | 2 | Vitamin B12 and ferrtin; buffered HSA; protein stabilizers; sodium azide | 50 | 28 |  |  |  | Yes | Atellica IM Multi‑Diluent 1 | 10995637 (2‑pack) 10995638 (6‑pack) 10995639 (vial) |  | 0.5 – 16,500 ng/mL | 1.1 - 36,300 pmol/L | No | NA | NA | 0.0 ng/mL | 0.0 pmol/L | NA | NA | 0.7 ng/mL | 0.3 ng/mL | 0.9 ng/mL | 0.7 pmol/L | 1.5 pmol/L | 2.0 pmol/L | 8.2 ng/mL | 18.0 pmol/L | 1.8 | 5.4 | 487.1 ng/mL | 1071.6 pmol/L | 1.6 | 4.2 | 1453.6 ng/mL | 3197.9 pmol/L | 3.4 | 6.3 | ADVIA Centaur FER y = 1.03x - 0.6 ng/mL r= 1.0 n=106 | ADVIA Centaur FER y = 1.03x - 1.3 pmol/L r=1.0 | ≤ 10% change up to 900 mg/dL (0.56 mmol/L) of hemoglobin | ≤ 10% change up to 60 mg/dL conjugated bilirubin (1021 μmol/L) 60 mg/dL unconjugated bilirubin (1021 μmol/L) | ≤ 10% change up to 2000 mg/dL (22.6 mmol/L) | No Biotin in assay formulation | 80,000 ng/mL | 176,000 pmol/L | 80,000 ng/mL | 176,000 pmol/L | Traceable to World Health Organization 2nd International Standard (WHO 80/578) | 10995343_EN Rev. 03, 2019-07 | IFU Document | 0.834 | 6.969 | 2.087 | 4.37 | 0.379 | 17.7 | 5.3 | 11.1
Atellica® IM | 10995490 |  | on05 | 00630414598178 | OUS | Oncology | No | Atellica IM CA 19‑9™ (CA 19‑9™) | Atellica IM CA 19‑9 | CA19_9 | 50 | 1 X 50 | N/A | Atellica IM CA 19‑9 DIL (SMN 10995491)Atellica IM CA 19‑9 MCM (SMN 10995492) | 28 | 28 | Mouse monoclonal anti-CA 19‑9 antibody | ;#Serum;# | up to 8 hrs at RT, 2–8°C up to 48 hrs, after 48 hrs at ≤ -20°C. Freeze once | 75 μL | No |  | 2-step sandwich immunoassay using direct chemiluminescent technology | 1.20–700.00 U/mL | NA |  |  |  |  | See IFU |  |  | Atellica IM CA 19‑9 CAL | Provided in Kit | at 2–8°C 14 days RT at 8hrs | Lyophilized | Add 2.0 mL of special reagent water, let the vials stand for 15–20 minutes at room temperature | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of CA 19‑9 (human) | 38 | 27 |  |  |  | Yes | Atellica IM CA 19‑9 DIL | 10995491 |  | 1.20–140,000.00 U/mL | N/A | No | NA | NA | 0.00 U/mL | NA | NA | NA | 3.99 U/mL | 0.00 U/mL | NA | NA | NA | NA | Control 1 25.56 U/mL | NA | 4 | 6 | Control 2 125.24 U/mL | NA | 3.9 | 6.2 | Control 3 301.12 U/mL | NA | 4.4 | 5.7 | ADVIA Centaur CA 19‑9 (x) y = 1.03x - 2.03 U/mL r=0.99 | NA | ≤ 4.0% change up to 1200 mg/dL of hemoglobin | ≤ 4.0% change up to 50 mg/dL of bilirubin | ≤ 4.0% change up to 3000 mg/dL of triglycerides | No Biotin in assay formulation | N/A | N/A |  |  | traceable to internal standard manufactured using highly purified material | 10995285_EN Rev. 03, 2019-07 | IFU Document | 0.317 | 6.969 | 2.48 | 2.913 | 0.144 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995489 |  | on05 | 00630414598161 | OUS | Oncology | No | Atellica IM CA 19‑9™ (CA 19‑9™) | Atellica IM CA 19‑9 | CA19_9 | 250 | 1 X 250 | N/A | Atellica IM CA 19‑9 DIL (SMN 10995491) Atellica IM CA 19‑9 MCM (SMN 10995492) | 28 | 28 | Mouse monoclonal anti-CA 19‑9 antibody | ;#Serum;# | up to 8 hrs at RT, 2–8°C up to 48 hrs, after 48 hrs at ≤ -20°C. Freeze once | 75 μL | No |  | 2-step sandwich immunoassay using direct chemiluminescent technology | 1.20–700.00 U/mL | NA |  |  |  |  | See IFU |  |  | Atellica IM CA 19‑9 CAL | Provided in Kit | at 2–8°C 14 days RT at 8hrs | Lyophilized | Add 2.0 mL of special reagent water, let the vials stand for 15–20 minutes at room temperature | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of CA 19‑9 (human) | 38 | 27 |  |  |  | Yes | Atellica IM CA 19‑9 DIL | 10995491 |  | 1.20–140,000.00 U/mL | N/A | No | NA | NA | 0.00 U/mL | NA | NA | NA | 3.99 U/mL | 0.00 U/mL | NA | NA | NA | NA | Control 1 25.56 U/mL | NA | 4 | 6 | Control 2 125.24 U/mL | NA | 3.9 | 6.2 | Control 3 301.12 U/mL | NA | 4.4 | 5.7 | ADVIA Centaur CA 19‑9 y = 1.03x - 2.03 U/mL r=0.99 | NA | ≤ 4.0% change up to 1200 mg/dL of hemoglobin | ≤ 4.0% change up to 50 mg/dL of bilirubin | ≤ 4.0% change up to 3000 mg/dL of triglycerides | No Biotin in assay formulation | N/A | N/A |  |  | traceable to internal standard manufactured using highly purified material | 10995285_EN Rev. 03, 2019-07 | IFU Document | 0.999 | 6.969 | 2.126 | 6.811 | 0.454 | 17.7 | 5.4 | 17.3
Atellica® IM | 10995585 |  | ty01 | 00630414599090 | Global | N/AThyroid | No | Atellica IM Free Triiodothyronine (FT3) | Atellica IM FT3FT3 | FT3 | 60 | 1 x 60 | Atellica IM CAL A | Atellica IM FT3 MCM (SMN 10995586) | 14 | 28 | acridinium -ester labeled mouse monoclonal anti-T3 antibody | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | up to 8hrs at RT, 8 hrs at 2–8°C, 48hrs at ≤ -20°C | 35 | No |  | Competitive immunoassay using direct chemilumimescent technology | 0.20–20.00 pg/mL | 0.31–30.80 pmol/L | NA | NA | 2.3 - 4.2 pg/mL | 3.5 - 6.5 pmol/L | NA | NA | NA | Atellica IM CAL A | ;2- pack (SMN 10995500) 6- pack (SMN 10995501) | 28 days recon at 2–8°C; 4 hrs recon at RT | Lyophilized | Reconstitute with 5 mL of distilled or deionized water . Let stand 15 - 20 minutes . Gently swirl or mix by inversion | 3 for lot calibration, 2 for pack recalibration | 2 | NA | 82 | 28 |  |  |  | No | NA | N/A |  | N/A | N/A | No | Use commercially available controls | NA | 0.06 pg/mL | 0.09 pmol/L | NA | NA | 0.02 pg/mL | 0.00 pg/mL | NA | 0.00 pmol/L | 0.03 pmol/L | NA | 2.87 pg/mL | 4.42 pmol/L | 0.94 | 1.78 | 4.96 pg/mL | 7.64 pmol/L | 0.67 | 1.07 | 8.71 pg/mL | 13.41 pmol/L | 0.78 | 1.13 | ADVIA Centaur FT3 y = 0.99x + 0.007 pg/mL; r=0.99 | ADVIA Centaur FT3 y = 0.99x + 0.011 pmol/L r=0.99 | ≤ 10% change up to 500 mg/dL (0.3 mmol/L) of hemoglobin | ≤ 10% change up to 20 mg/dL (341 μmol/L) conjugated bilirubin 20 mg/dL (341 μmol/L) unconjugated bilirubin | ≤ 10% change up to 1000 mg/dL (11.3 mmol/L) intralipids | No Biotin in assay formulation | N/A | N/A | N/A | N/A | Internal standard manufactured using U.S.P.(United States Pharmacopeia) material | 10995347_EN Rev. 02, 2020-02 | IFU Document | 0.205 | 6.969 | 2.087 | 1.457 | 0.093 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995584 |  | ty01 | 00630414599083 | Global | N/AThyroid | No | Atellica IM Free Triiodothyronine (FT3) | Atellica IM FT3FT3 | FT3 | 300 | 5 x 60 | Atellica IM CAL A | Atellica IM FT3 MCM (SMN 10995586) | 14 | 28 | acridinium-ester-labeled mouse monoclonal anti-T3 antibody | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | up to 8hrs at RT, 8 hrs at 2–8°C, 48hrs at ≤ -20°C | 35 | No |  | Competitive immunoassay using direct chemiluminescent technology | 0.20–20.00 pg/mL | 0.31–30.80 pmol/L | NA | NA | 2.3 - 4.2 pg/mL | 3.5 - 6.5 pmol/L | NA | NA | NA | Atellica IM CAL A | ;2- pack (SMN 10995500) 6- pack (SMN 10995501) | 28 days recon at 2-8°C; 4 hrs recon at RT | Lyophilized | Reconstitute with 5 mL of distilled or deionized water . Let stand 15 - 20 minutes . Gently swirl or mix by inversion | 3 for lot calibration, 2 for pack recalibration | 2 | NA | 82 | 28 |  |  |  | No | NA | N/A |  | N/A | N/A | No | Use commercially available controls | NA | 0.06 pg/mL | 0.09 pmol/L | NA | NA | 0.02 pg/mL | 0.00 pg/mL | NA | 0.00 pmol/L | 0.03 pmol/L | NA | 2.87 pg/mL | 4.42 pmol/L | 0.94 | 1.78 | 4.96 pg/mL | 7.64 pmol/L | 0.67 | 1.07 | 8.71 pg/mL | 13.41 pmol/L | 0.78 | 1.13 | ADVIA Centaur FT3 y = 0.99x + 0.007 pg/mL; r=0.99 | ADVIA Centaur FT3 y = 0.99x + 0.011 pmol/L r=0.99 | ≤ 10% change up to 500 mg/dL (0.3 mmol/L) of hemoglobin | ≤ 10% change up to 20 mg/dL (341 μmol/L) conjugated bilirubin 20 mg/dL (341 μmol/L) unconjugated bilirubin | ≤ 10% change up to 1000 mg/dL (11.3 mmol/L) intralipids | No Biotin in assay formulation | N/A | N/A | N/A | N/A | Internal standard manufactured using U.S.P.(United States Pharmacopeia) material | 10995347_EN Rev. 02, 2020-02 | IFU Document | 0.823 | 6.969 | 2.087 | 4.37 | 0.374 | 17.7 | 5.3 | 11.1
Atellica® IM | 10995679 |  | ty08 | 00630414599922 | Global | N/AThyroid | No | Atellica IM Total Triiodothyronine (T3) | Atellica IM T3T3 | T3 | 120 | 1 x 120 | Atellica IM CAL A | Atellica IM T3 DIL (SMN 10995680) Atellica IM T3 MCM (SMN 10995681) | 14 | 28 | acridinium-ester-labeled mouse monoclonal anti-T3 antibody | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | up to 8hrs at RT, 8 hrs at 2–8°C, 48hrs at ≤ -20°C | 30 | No |  | Competitive immunoassay using direct chemiluminescent technology | 0.10–8.00 ng/mL | 0.15–12.30 nmol/L | NA | NA | 0.6 - 1.81 ng/mL | 0.92 - 2.79 nmol/L | NA | NA | NA | Atellica IM CAL A | 2-pack (SMN 10995500) 6-pack (SMN 10995501) | 28 days recon at 2–8°C; 4 hrs recon at RT | Lyophilized | Reconstitute with 5 mL of distilled or deionized water . Let stand 15 - 20 minutes . Gently swirl or mix by inversion | 3 for lot calibration, 2 for pack recalibration | 2 | human plasma | 56 | 21 |  |  |  | Manual | Atellica IM T3 DIL | 10995680 |  | N/A | N/A | No | Use commercially available controls | NA | 0.10 ng/mL | 0.15 nmol/L | NA | NA | 0.38 ng/mL | 0.00 ng/mL | NA | 0.00 nmol/L | 0.59 nmol/L | NA | 0.74 ng/mL | 1.14 nmol/L | 2.8 | 5.2 | 1.23 ng/mL | 1.89 nmol/L | 2.5 | 6.7 | 7.93 ng/mL | 12.21 nmol/L | 1.8 | 9.3 | ADVIA Centaur T3; y = 1.08x - 0.02 ng/mL r=0.99 | ADVIA Centaur T3; y = 1.08x - 0.02 ng/mL r=0.99 | ≤ 10% change up to 250 mg/dL (0.15 mmol/L) hemoglobin | ≤ 10% change up to 20 mg/dL (341 μmol/L) conjugated bilirubin 20 mg/dL (341 μmol/L) unconjugated bilirubin | ≤ 10% change up to1000 mg/dL (11.3 mmol/L) intralipids | No Biotin in assay formulation | N/A | N/A | N/A | N/A | Internal standard manufactured using U.S.P.(United States Pharmacopeia) material | 10995424_EN Rev. 03, 2020-06 | IFU Document | 0.216 | 6.969 | 2.087 | 1.457 | 0.098 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995678 |  | ty08 | 00630414599915 | Global | N/AThyroid | No | Atellica IM Total Triiodothyronine (T3) | Atellica IM T3T3 | T3 | 600 | 5 x 120 | Atellica IM CAL A | Atellica IM T3 DIL (SMN 10995680) Atellica IM T3 MCM (SMN 10995681) | 14 | 28 | acridinium-ester-labeled mouse monoclonal anti-T3 antibody | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | up tp 8hrs at RT 8 hrs at 2–8°C 48hrs at ≤ -20°C | 30 | No |  | Competitive immunoassay using direct chemiluminescent technology | 0.10–8.00 ng/mL | 0.15–12.30 nmol/L | NA | NA | 0.6 - 1.81 ng/mL | 0.92 - 2.79 nmol/L | NA | NA | NA | Atellica IM CAL A | 2-pack (SMN 10995500) 6-pack (SMN 10995501) | 28 days recon at 2–8°C; 4 hrs recon at RT | Lyophilized | Reconstitute with 5 mL of distilled or deionized water . Let stand 15 - 20 minutes . Gently swirl or mix by inversion | 3 for lot calibration, 2 for pack recalibration | 2 | human plasma | 56 | 21 |  |  |  | Manual | Atellica IM T3 DIL | 10995680 |  | N/A | N/A | No | Use commercially available controls | NA | 0.10 ng/mL | 0.15 nmol/L | NA | NA | 0.38 ng/mL | 0.00 ng/mL | NA | 0.00 nmol/L | 0.59 nmol/L | NA | 0.74 ng/mL | 1.14 nmol/L | 2.8 | 5.2 | 1.23 ng/mL | 1.89 nmol/L | 2.5 | 6.7 | 7.93 ng/mL | 12.21 nmol/L | 1.8 | 9.3 | ADVIA Centaur T3; y = 1.08x - 0.02 ng/mL r=0.99 | ADVIA Centaur T3; y = 1.08x - 0.02 ng/mL r=0.99 | ≤ 10% change up to 250 mg/dL (0.15 mmol/L) hemoglobin | ≤ 10% change up to 20 mg/dL (341 μmol/L) conjugated bilirubin 20 mg/dL (341 μmol/L) unconjugated bilirubin | ≤ 10% change up to1000 mg/dL (11.3 mmol/L) intralipids | No Biotin in assay formulation | N/A | N/A | N/A | N/A | Internal standard manufactured using U.S.P.(United States Pharmacopeia) material | 10995424_EN Rev. 03, 2020-06 | IFU Document | 0.865 | 6.969 | 2.087 | 4.37 | 0.393 | 17.7 | 5.3 | 11.1
Atellica® IM | 10995600 |  | id09 | 00630414599236 | OUS | Hepatitis | No | Atellica IM Hepatitis B e Antigen (HBeAg) | Atellica IM HBeAgHBeAg | HBeAg | 50 | 1 x 50 | Atellica IM PW3 (SMN 10995666) | Atellica IM HBeAg QC Atellica IM Multi‑Diluent 10 (SMN10995640) | 39 | 60 |  | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#EDTA Plasma;# | RT not recommended; at 2–8°C up to 3 days (primary tubes), up to 7 days (separated samples); ≤ -20°C for longer storage | 100 μL | No |  |  | 0.05–1000.00 Index | 0.05–1000.00 Index | NA | NA | NA | NA | See IFU |  |  | Atellica IM HBeAg CAL | Provided in Kit | at 2–8°C until expiration 8 hours at RT | Liquid | N/A | 3 for lot calibration, 2 for pack recalibration | 2 |  | 74 | 30 |  |  |  | Yes | Atellica IM Multi‑Diluent 10 | 10995640 |  | na | N/A | Yes | Atellica IM HBeAg QC | 10995601 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 0.91 Index | 0.91 Index | 6.9 | 8.3 | 219.56 Index | 219.56 Index | 1.3 | 1.7 | 770.11 Index | 770.11 Index | 1.1 | 1.8 | NA | NA | NA | NA | NA | Insignificant effect on the assay up to 1500 ng/mL | N/A | N/A |  |  | standardization is based upon relative clinical agreement with commercially available HBeAg assays | 10995355_EN Rev. 03, 2019-07 | IFU Document | 0.264 | 6.969 | 2.48 | 2.913 | 0.12 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995610 |  | id11 | 00630414599335 | OUS | HIV | No | Atellica IM HIV 1/O/2 Enhanced (EHIV) | Atellica IM EHIVEHIV | EHIV | 200 | 1x200 | Atellica IM PW3 | Atellica IM EHIV QC | 39 | 28 | recombinant HIV antigens/Streptavidin-coated paramagnetic microparticles preformed with biotinylated HIV antigens, recombinant HIV ag labelled with AE | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#EDTA Plasma;# | at 2–8°C up to 7 days; ≤ -20°C for longer storage | 50 μL | No |  | antigen-bridging microparticle chemiluminometric immunoassay | 0.05–50.00 Index | 0.05–50.00 Index | N/A | N/A | see IFU | see IFU | see IFU | see IFU | see IFU | Atellica IM EHIV CAL | Provided in Kit | at 2–8°C until expiration RT at 8hrs on board | Liquid | NA | 3 for lot calibration, 2 for pack recalibration | 2 | Processed human plasma negative for antibodies to HIV spiked with antibodies to HIV‑1 | 29 | 14 |  |  |  | No | NA | N/A |  | N/A | N/A | Yes | Atellica IM EHIV QC | 10995612 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 0.55 Index | 0.55 Index | 5.1 | 10.5 | 6.37 | 6.37 | 2.6 | 2.8 | 25.66 | 25.66 | 4.3 | 5.4 | NA | NA | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 30 mg/dL of conjugated bilirubin, 20 mg/dL of unconjugated bilirubin | ≤ 10% change up to 3000 mg/dL of triglycerides | Insignificant effect on the assay up to 3500 ng/mL | N/A | N/A | N/A | N/A | standardization is based upon relative clinical agreement with commercially available anti-HIV assays | 110995364_EN Rev. 04, 2020-11 | IFU Document | 0.323 | 6.969 | 2.48 | 2.913 | 0.147 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995656 |  | re06 | 00630414599748 | Global | Reproductive Endocrinology | No | Atellica IM Prolactin (PRL) | Atellica IM PRL | PRL | 50 | 1 x 50 | Atellica IM CAL B | Atellica IM Multi‑Diluent 1 Atellica IM PRL MCM (SMN 10995657) | 14 | 28 | goat polyclonal anti-prolactin antibody | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | up to 8hrs at RT, 8 hrs at 2–8°C, 48 hrs at ≤ -20°C | 25 μL | No |  | 2-site sandwich immunoassay | 0.30–200.00 ng/mL | 6.36–4240.00 μIU/mL |  |  |  |  | See IFU |  |  | Atellica IM CAL B | 10995503 (2‑pack) 10995504 (6‑pack) | at 2–8°C 28 days | Lyophilized | Add 5.0 mL of special reagent water into each vial Let the vials stand for 15–20 min at RT | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of digoxin,follicle-stimulating hormone (FSH), luteinizinghormone (LH), prolactin, human chorionic gonadotropin (hCG), and TSH; equine serum | 34 | 28 |  |  |  | Yes | Atellica IM Multi‑Diluent 1 | 10995637 (2‑pack) 10995638 (6‑pack) 10995639 (vial) |  | 0.30–1000.00 ng/mL | 6.36–21200 μIU/mL | No | NA | NA | 0.20 ng/mL | 4.24 μIU/mL | NA | NA | 0.35 ng/mL | 0.06 ng/mL | NA | 1.27 μIU/mL | 7.42 μIU/mL | NA | Control 1 16.69 ng/mL | Control 1 353.83 μIU/mL | 1.4 | 3.3 | Control 2 37.23 ng/mL | Control 2 789.28 μIU/mL | 1.5 | 2.3 | Control 3 182.25 ng/mL | Control 3 3863.70 μIU/mL | 3.1 | 4.9 | ADVIA Centaur PRL y = 0.94x + 0.26 ng/mL r=1.0 | ADVIA Centaur PRL y = 0.94x + 5.49 μIU/mL r=1.0 | insignificant effect up to 500 mg/dL of hemoglobin | insignificant effect up to 20 mg/dL of bilirubin | insignificant effect up to 1000 mg/dL of triglycerides | No Biotin in assay formulation | 30,000 ng/mL | 636,000 μIU/mL |  |  | traceable to the World Health Organization (WHO) 3rd IRP for human Prolactin (84/500) | 10995414_EN Rev. 03, 2020-02 (OUS), 11200387_EN Rev. 03, 2020-02 (US) | IFU Document | 0.205 | 6.969 | 2.087 | 1.457 | 0.093 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995655 |  | re06 | 00630414599731 | Global | Reproductive Endocrinology | No | Atellica IM Prolactin (PRL) | Atellica IM PRL | PRL | 250 | 5 x 50 | Atellica IM CAL B | Atellica IM Multi‑Diluent 1 Atellica IM PRL MCM (SMN 10995657) | 14 | 28 | goat polyclonal anti-prolactin antibody | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | up to 8hrs at RT, 8 hrs at 2–8°C, 48 hrs at ≤ -20°C | 25 μL | No |  | 2-site sandwich immunoassay | 0.30–200.00 ng/mL | 6.36–4240.00 μIU/mL |  |  |  |  | See IFU |  |  | Atellica IM CAL B | 10995503 (2‑pack) 10995504 (6‑pack) | at 2–8°C 28 days | Lyophilized | Add 5.0 mL of special reagent water into each vial Let the vials stand for 15–20 min at RT | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of digoxin,follicle-stimulating hormone (FSH), luteinizinghormone (LH), prolactin, human chorionic gonadotropin (hCG), and TSH; equine serum | 34 | 28 |  |  |  | Yes | Atellica IM Multi‑Diluent 1 | 10995637 (2‑pack) 10995638 (6‑pack) 10995639 (vial) |  | 0.30–1000.00 ng/mL | 6.36–21200 μIU/mL | No | NA | NA | 0.20 ng/mL | 4.24 μIU/mL | NA | NA | 0.35 ng/mL | 0.06 ng/mL | NA | 1.27 μIU/mL | 7.42 μIU/mL | NA | Control 1 16.69 ng/mL | Control 1 353.83 μIU/mL | 1.4 | 3.3 | Control 2 37.23 ng/mL | Control 2 789.28 μIU/mL | 1.5 | 2.3 | Control 3 182.25 ng/mL | Control 3 3863.70 μIU/mL | 3.1 | 4.9 | ADVIA Centaur PRL y = 0.94x + 0.26 ng/mL r=1.0 | ADVIA Centaur PRL y = 0.94x + 5.49 μIU/mL r=1.0 | insignificant effect up to 500 mg/dL of hemoglobin | insignificant effect up to 20 mg/dL of bilirubin | insignificant effect up to 1000 mg/dL of triglycerides | No Biotin in assay formulation | 30,000 ng/mL | 636,000 μIU/mL |  |  | traceable to the World Health Organization (WHO) 3rd IRP for human Prolactin (84/500) | 10995414_EN Rev. 03, 2020-02 (OUS), 11200387_EN Rev. 03, 2020-02 (US) | IFU Document | 0.818 | 6.969 | 2.087 | 4.37 | 0.372 | 17.7 | 5.3 | 11.1
Atellica® IM | 10995660 |  | re07 | 00630414599786 | Global | Reproductive Endocrinology | No | Atellica IM Progesterone (PRGE) | Atellica IM PRGE | PRGE | 90 | 1 x 90 | Atellica IM CAL E | Atellica IM Multi‑Diluent 3 Atellica IM PRGE MCM (SMN 10995661) | 14 | 28 | mouse monoclonal anti-progesterone antibody | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | up to 8hrs at RT, 8 hrs at 2–8°C, 48 hrs at ≤ -20°C, Freeze samples only 1 time and mix thoroughly after thawing. | 10 μL | No |  | competitive immunoassay | 0.21–60.00 ng/mL | 0.67–190.80 nmol/L |  |  |  |  | See IFU for Normals, Pregnant and Postmenopausal Females | Males 0.28–1.22 ng/mL 0.89–3.88 nmol/L | See IFU for Pediatric Reference Intervals by Age | Atellica IM CAL E | 10995512 (2‑pack) 10995513 (6‑pack) | at 2–8°C 28 days | Lyophilized | Add 2.0 mL of special reagent water into each vial. Let vials stand for 15-20 minutes at room temperature. | 3 for lot calibration, 2 for pack recalibration | 2 | After reconsitution, low and high levels of cortisol, progesterone, and testosterone; human serum | 90 | 28 |  |  |  | Yes | Atellica IM Multi‑Diluent 3 | 10995645 |  | 0.21–600.00 ng/mL | 0.67–1908.0 nmol/L | No | NA | NA | 0.10 ng/mL | 0.32 nmol/L | NA | NA | 0.31 ng/mL | 0.18 ng/mL | NA | 0.57 nmol/L | 0.98 nmol/L | NA | Control 1 0.64 ng/mL | Control 1 2.04 nmol/L | 6.1 | 12.9 | Control 3 18.63 ng/mL | Control 3 59.24 nmol/L | 1.9 | 2.9 | Serum 53.84 ng/mL | Serum 171.21 nmol/L | 1.6 | 3.8 | ADVIA Centaur Progesterone y = 0.980x + 0.014 ng/mL r=0.99 | ADVIA Centaur Progesterone y = 0.980x + 0.044 nmol/L r=0.99 | ≤ 10% change up to 250 mg/dL (0.15 mmol/L) hemoglobin | ≤ 10% change up to 20 mg/dL (341 μmol/L) conjugated bilirubin 20 mg/dL (341 μmol/L) unconjugated bilirubin | ≤ 10% change up to 500 mg/dL (5.65 mmol/L) intralipids | No Biotin in assay formulation | N/A | N/A |  |  | standardized using internal standards manufactured analytically which are traceable to gas chromatography–mass spectroscopy (GCMS) | 10995415_EN Rev. 03, 2020-02 (OUS), 11200386_EN Rev.03, 2020-02 (US) | IFU Document | 0.246 | 6.969 | 2.48 | 2.913 | 0.112 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995659 |  | re07 | 00630414599779 | Global | Reproductive Endocrinology | No | Atellica IM Progesterone (PRGE) | Atellica IM PRGE | PRGE | 450 | 5 x 90 | Atellica IM CAL E | Atellica IM Multi‑Diluent 3 Atellica IM PRGE MCM (SMN 10995661) | 14 | 28 | mouse monoclonal anti-progesterone antibody | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | up to 8hrs at RT, 8 hrs at 2–8°C, 48 hrs at ≤ -20°C | 10 μL | No |  | competitive immunoassay | 0.21–60.00 ng/mL | 0.67–190.80 nmol/L |  |  |  |  | See IFU for Normals, Pregnant and Postmenopausal Females | Males 0.28–1.22 ng/mL 0.89–3.88 nmol/L | See IFU for Pediatric Reference Intervals by Age | Atellica IM CAL E | 10995512 (2‑pack) 10995513 (6‑pack) | at 2–8°C 28 days | Lyophilized | Add 2.0 mL of special reagent water into each vial. Let vials stand for 15-20 minutes at room temperature. | 3 for lot calibration, 2 for pack recalibration | 2 | After reconsitution, low and high levels of cortisol, progesterone, and testosterone; human serum | 90 | 28 |  |  |  | Yes | Atellica IM Multi‑Diluent 3 | 10995645 |  | 0.21–600.00 ng/mL | 0.67–1908.0 nmol/L | No | NA | NA | 0.10 ng/mL | 0.32 nmol/L | NA | NA | 0.31 ng/mL | 0.18 ng/mL | NA | 0.57 nmol/L | 0.98 nmol/L | NA | Control 1 0.64 ng/mL | Control 1 2.04 nmol/L | 6.1 | 12.9 | Control 3 18.63 ng/mL | Control 3 59.24 nmol/L | 1.9 | 2.9 | Serum 53.84 ng/mL | Serum 171.21 nmol/L | 1.6 | 3.8 | ADVIA Centaur Progesterone y = 0.980x + 0.014 ng/mL r=0.99 | ADVIA Centaur Progesterone y = 0.980x + 0.044 nmol/L r=0.99 | ≤ 10% change up to 250 mg/dL (0.15 mmol/L) hemoglobin | ≤ 10% change up to 20 mg/dL (341 μmol/L) conjugated bilirubin 20 mg/dL (341 μmol/L) unconjugated bilirubin | ≤ 10% change up to 500 mg/dL (5.65 mmol/L) intralipids | No Biotin in assay formulation | N/A | N/A |  |  | standardized using internal standards manufactured analytically which are traceable to gas chromatography–mass spectroscopy (GCMS) | 10995415_EN Rev. 03, 2020-02 (OUS), 11200386_EN Rev.03, 2020-02 (US) | IFU Document | 0.957 | 10.748 | 2.48 | 6.063 | 0.435 | 27.3 | 6.3 | 15.4
Atellica® IM | 10995708 |  | re10 | 00630414600222 | Global | Reproductive Endocrinology | No | Atellica IM Testosterone II (TSTII) | Atellica IM TSTII | TSTII | 100 | 1 x 100 | Atellica IM APW1 (SMN 10995458) | Atellica IM Multi‑Diluent 3 Atellica IM TSTII MCM (SMN 10995709) | 14 | 18 | sheep monoclonal antibody | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#EDTA Plasma;# | up to 48 hours at RT, 7 days at 2–8ºC, 7 days at ≤ -20°C | 20 μL | No |  | competitive immunoassay using direct chemiluminescent technology | 7.00–1500.00 ng/dL | 0.24–52.05 nmol/L |  |  |  |  | See IFUs |  |  | Atellica IM TSTII CAL | Provided in Kit | 14 days recon at 2–8°C 6 hrs recon at RT | Lyophilized | 2.0 mL of special reagent water. Let the vials stand for 30 minutes at room temperature | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of USP-grade estradiol, testosterone, cortisol, and progesterone spiked in charcoal-stripped, defibrinated human plasma;ked in | 91 | 28 |  |  |  | Yes | Atellica IM Multi‑Diluent 3 | 10995645 |  | 7.00–3000 ng/dL | 0.24–104.1 nmol/L | No | NA | NA | NA | NA | NA | NA | 4.43 ng/dL | 1.81 ng/dL | 4.43 ng/dL | 0.06 nmol/L | 0.15 nmol/L | 0.15 nmol/L | 17.81 ng/dL | 0.62 nmol/L | 2.2 | 5.9 | 290.14 ng/dL | 10.07 nmol/L | 1.8 | 2.9 | 1014.48 ng/dL | 35.2 nmol/L | 2.7 | 4.7 | ADVIA Centaur TSTII y = 0.97x + 2.04 ng/dL r=1.0 | ADVIA Centaur TSTII y = 0.97x + 0.07 nmol/L r=1.0 | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 15 mg/dL of conjugated bilirubin 20 mg/dL of unconjugated bilirubin | ≤ 10% change up to 1000 mg/dL of triglycerides | ≤ 10% change in results up to 30 ng/mL of biotin | N/A | N/A |  |  | traceable to the primary testosterone standard National Measurement Institute (NMI) M914 | 10995435_EN Rev. 04, 2019-08 (OUS), 11200389_EN Rev. 04, 2019-06 (US) | IFU Document | 0.381 | 6.969 | 2.48 | 2.913 | 0.173 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995707 |  | re10 | 00630414600215 | Global | Reproductive Endocrinology | No | Atellica IM Testosterone II (TSTII) | Atellica IM TSTII | TSTII | 500 | 1 x 500 | Atellica IM APW1 (SMN 10995458) | Atellica IM Multi‑Diluent 3 Atellica IM TSTII MCM (SMN 10995709) | 14 | 18 | sheep monoclonal antibody | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#EDTA Plasma;# | up to 48 hours at RT, 7 days at 2–8ºC, 7 days at ≤ -20°C | 20 μL | No |  | competitive immunoassay using direct chemiluminescent technology | 7.00–1500.00 ng/dL | 0.24–52.05 nmol/L |  |  |  |  | See IFUs |  |  | Atellica IM TSTII CAL | Provided in Kit | 14 days recon at 2–8°C 6 hrs recon at RT | Lyophilized |  | 3 for lot calibration, 2 for pack recalibration | 2 |  | 91 | 28 |  |  |  | Yes | Atellica IM Multi‑Diluent 3 | 10995645 |  | 7.00–3000 ng/dL | 0.24–104.1 nmol/L | No | NA | NA | NA | NA | NA | NA | 4.43 ng/dL | 1.81 ng/dL | 4.43 ng/dL | 0.06 nmol/L | 0.15 nmol/L | 0.15 nmol/L | 17.81 ng/dL | 0.62 nmol/L | 2.2 | 5.9 | 290.14 ng/dL | 10.07 nmol/L | 1.8 | 2.9 | 1014.48 ng/dL | 35.2 nmol/L | 2.7 | 4.7 | ADVIA Centaur TSTII y = 0.97x + 2.04 ng/dL r=1.0 | ADVIA Centaur TSTII y = 0.97x + 0.07 nmol/L r=1.0 | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 15 mg/dL of conjugated bilirubin 20 mg/dL of unconjugated bilirubin | ≤ 10% change up to 1000 mg/dL of triglycerides | ≤ 10% change in results up to 30 ng/mL of biotin | N/A | N/A |  |  | traceable to the primary testosterone standard National Measurement Institute (NMI) M914 | 10995435_EN Rev. 04, 2019-08(OUS), 11200389_EN Rev. 04, 2019-06 (US) | IFU Document | 1.415 | 6.969 | 2.126 | 6.811 | 0.643 | 17.7 | 5.4 | 17.3
Atellica® CH | 11097637 |  | di03 | 00630414595580 | Global | Diabetes | No | Atellica CH Fructosamine (Fruc) | FrucFruc | 0628 | 400 | 2 x 200 | N/A | N/A | 10 | 30 | Proteinase-K; 4-aminoantipyrine: R2: fructosaminase; peroxidase (horseradish) N-Ethyl-N-Sulphohydroxypropyl-m-toluidine (TOOS) | ;#Serum;#Lithium Heparin Plasma;#Potassium EDTA Plasma;# | Up to 3 days at 20-25C; Up to 14 days at 4-8C; Up to 60 days frozen at -20C | 16 μL | No |  | Colorimetric end-point reaction | 30–1000 μmol/L | 30–1000 μmol/L | N/A | N/A | 122–236 μmol/L | 122–236 μmol/L | N/A | N/A | N/A | Atellica CH Fruc CAL | 11099432 | 28 days | Lyophilized | Let the vials stand for 10–15 minutes at room temperature; Add 1.0 mL of room temperature deionized water: Let the vials stand for 5 minutes at room temperature; , swirl and gently invert the vials at least 5 times | 3 | 2 | human serum albumin | 180 | 30 |  | No |  | No | N/A | 11099300 |  | N/A | N/A | No | N/A | N/A | N/A | N/A | N/A | N/A | 10 μmol/L | 6 μmol/L | 28 μmol/L | N/A | N/A | N/A | 35 μmol/L | 35 μmol/L | 4 | 5.5 | 324 μmol/L | 324 μmol/L | 0.7 | 0.9 | 636 μmol/L | 636 μmol/L | 0.4 | 1.4 | ADVIA Chemistry 1800 y = 1.02x + 3 μmol/L r=0.994 n= 112 | ADVIA Chemistry 1800 y = 1.02x + 3 μmol/L r=0.994 n= 112 | ≤ 3% change up to 750 mg/dL of hemoglobin. | ≤ 8% change up to 5 mg/dL of conjugated bilirubin; ≤ 8% change up to 7.5 mg/dL of unconjugated bilirubin | ≤ 6% change up to 1000 mg/dL of Trig. | No Biotin in assay formulation | N/A | N/A | N/A | N/A | Internal standard | 11110200_EN Rev. 02, 2019-07 | IFU Document | 0.341 | 2.875 | 3.5 | 2.75 | 0.155 | 7.302 | 8.89 | 6.985
Atellica® CH | 11318376 |  | pl33 | 05055273215564 | OUS | Specific Proteins | Yes | Soluble Transferrin Receptors (Randox Alliance Application) | sTfRsTfR | 8092 | 85 | 1 x 85 | Atellica CH Empty Reagent Packs (11097534) | N/A | 9 | 14 | Antibody-latex Reagent; Mouse monoclonal anti-STFR antibodies | ;#Serum;#Lithium Heparin Plasma;# | If not analysed immediately, serum should be stored at -20°C. | 8.9 | No |  | Turbidimetry | 0.50 – 11.77 mg/L | 0.50 – 11.77 mg/L | N/A | N/A | 0.65 - 1.88 mg/L | 0.65 - 1.88 mg/L | N/A | N/A | N/A | Randox STFR Calibrator | 11306493 | 28 days | Lyophilized | 1. Open vial very carefully, avoiding any loss of material. 2. Reconstitute each vial with 1 ml of double deionised water. 3. Replace stopper, close vial and mix carefully for 30 minutes. 4. Ensure contents are completely dissolved before use. | 3 | 6 | Human Source | 6 | 6 |  | No |  | No | N/A | 11099300 |  | N/A | N/A | No | Randox STFR CONTROL | 11309043 | N/A | N/A | N/A | N/A | 0.13 mg/L | 0.7 mg/L | 0.50 mg/L | 0.7 mg/L | 0.13 mg/L | 0.50 mg/L | 0.61 mg/L | 0.61 mg/L | 8.8 | 9.9 | 1.95 mg/L | 1.95 mg/L | 2.5 | 5 | 4.75 mg/L | 4.75 mg/L | 1.5 | 2.7 | Roche cobas 501 y=0.97x -0.023 mg/L r=0.997 | Roche cobas 501 y=0.97x -0.023 mg/L r=0.997 | ≤ 10% change up to 175 mg/dL of hemoglobin. | ≤ 10% change up to 60 mg/dL of free and conjugate bilirubin. | ≤ 10% change up to 400 mg/dL of Intralipd | No Biotin in assay formulation | N/A | N/A | N/A | N/A | WHO Reference Reagent Recombinant Soluble Transferrin Receptor (NIBSC code:07/202) | IFU: Revised 27 Feb 18 ml and App. Sheet: (s): Revised 28 Mar 18 ml | IFU Document | 0.1201 | 4.8 | 1.96 | 1.88 | 0.0545 | 12.2 | 5 | 4.8
Atellica® IM | 10997840 |  | ca02 | 00630414006703 | OUS | Cardiac | No | Atellica IM High Sensitivity Troponin I (TnIH) | Atellica IM TnIH | TnIH | 100 | 1X100 | Atellica IM APW3 ( SMN 10998580) | Atellica IM Multi‑Diluent 11 (10995642)Atellica IM TnIH MCM (SMN 10997842) | 10 | 31 | Bovine serum albumin (BSA) conjugated to a recombinant monoclonal (sheep) Fab anti-human cTnI | ;#Serum;#Lithium Heparin Plasma;# | RT for up to 8 hours 2–8°C for 24 hours ≤ -20°C for up to 40 days ≤ -70°C for up to 1 year | 100 μL | No |  | direct chemiluminescent technology | 2.51–25,000.00 pg/mL | 2.51–25,000.00 ng/L |  |  |  |  | Combined gender and lithium heparin sample type were used to determine the overall 99th percentile of 45.20 pg/mL (ng/L) |  |  | Atellica IM TnIH CAL | Provided in Kit | 4 hrs recon at 2–8°C 30 days recon at ≤ -20°C 4 hrs onboard at RT | Lyophilized | Low Calibrator is liquid ready to use. High Calibrator - Reconstituted to 1.00 mL/vial | 3 for lot calibration, 2 for pack recalibration | 2 | HEPES buffer; bovine serum albumin (BSA) | 47 | 31 |  |  |  | Yes | Atellica IM Multi‑Diluent 11 | 10995642 (2‑pack) |  | 2.50 - 125,000 pg/mL | 2.50 - 125,000 ng/L | No | NA | NA | NA | NA | NA | NA | 1.27 pg/mL | 0.58 pg/mL | 2.51 pg/mL | 0.58 ng/L | 1.27 ng/L | 2.51 ng/L | 9.54 | N/A | 5 | 7.5 | 147.37 | N/A | 1.7 | 3 | 21 392.14 | N/A | 1.9 | 2.5 | See IFU | See IFU | <10% change up to 500 mg/dL | <10% change up to 40 mg/dL of conjugated bilirubin . 60 mg/dL of unconjugated bilirubin | <10% change up to 2000 mg/dL of triglycerides | ≤ 10% change in results up to 1500 ng/mL of biotin. | 618,000 pg/mL | 618,000 ng/L |  |  | Traceable to an internal standard manufactured using human heart homogenate | 11200497_EN Rev. 06, 2019-08 | IFU Document | 0.286 | 6.875 | 2.125 | 2.438 | 0.13 | 17.463 | 5.398 | 6.193
Atellica® IM | 10997841 |  | ca02 | 00630414006758 | OUS | Cardiac | No | Atellica IM High Sensitivity Troponin I (TnIH) | Atellica IM TnIH | TnIH | 500 | 5 X 100 | Atellica IM APW3 (SMN 10998580) | Atellica IM Multi‑Diluent 11 (SMN 10995642) Atellica IM TnIH MCM (SMN 10997842) | 10 | 28 | Bovine serum albumin (BSA) conjugated to a recombinant monoclonal (sheep) Fab anti-human cTnI | ;#Serum;#Lithium Heparin Plasma;# | RT for up to 8 hours 2–8°C for 24 hours ≤ -20°C for up to 40 days ≤ -70°C for up to 1 year | 100 μL | No |  | direct chemiluminescent technology | 2.51–25,000.00 pg/mL | 2.51–25,000.00 ng/L |  |  |  |  | Combined gender and lithium heparin sample type were used to determine the overall 99th percentile of 45.20 pg/mL (ng/L) |  |  | Atellica IM TnIH CAL | Provided in Kit | 4 hrs recon at 2–8°C 30 days recon at ≤ -20°C 4 hrs onboard at RT | Lyophilized | Low calibrator is liquid ready to use. High calibrator - Reconstituted to 1.00 mL/vial | 3 for lot calibration, 2 for pack recalibration | 2 |  | 47 | 31 |  |  |  | Yes | Atellica IM Multi‑Diluent 11 | 10995642 (2‑pack) |  | 2.50 - 125,000 pg/mL | 2.50 - 125,000 ng/L | No | NA | NA | NA | NA | NA | NA | 1.27 pg/mL | 0.58 pg/mL | 2.51 pg/mL | 0.58 ng/L | 1.27 ng/L | 2.51 ng/L | 9.54 | N/A | 5 | 7.5 | 147.37 | N/A | 1.7 | 3 | 21 392.14 | N/A | 1.9 | 2.5 | See IFU | See IFU | <10% change up to 500 mg/dL | <10% change up to 40 mg/dL of conjugated bilirubin . 60 mg/dL of unconjugated bilirubin | <10% change up to 2000 mg/dL of triglycerides | ≤ 10% change in results up to 1500 ng/mL of biotin. | 618,000 pg/mL | 618,000 ng/L |  |  | Traceable to an internal standard manufactured using human heart homogenate | 11200497_EN Rev. 06, 2019-08 | IFU Document | 0.953 | 6.875 | 2.125 | 6.75 | 0.433 | 17.463 | 5.398 | 17.145
Atellica® IM | 10995456 |  | id06 | 00630414597867 | OUS | Hepatitis | No | Atellica IM Hepatitis C Antibody (aHCV) | Atellica IM aHCVaHCV | aHCV | 200 | 1 x 200 | Atellica IM APW1 (SMN 10995458) | Atellica IM aHCV QC | 46 | 41 | anti-human IgG monoclonal antibody | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#EDTA Plasma;# | 2–8°C for up to 7 days, ≤ -20°C for longer storage | 10 μL | No |  | indirect two wash sandwich immunoassay | 0.02–11.00 Index | 0.02–11.00 Index | NA | NA | NA | NA | nonreactive <0.80 Index | reactive ≥ 1.00 Index | equivocal ≥ 0.80 Index and <1.00 Index | Atellica IM aHCV CAL | Provided in Kit | at 2–8°C until expiration 8 hours at RT | Liquid | N/A | 3 for lot calibration, 2 for pack recalibration | 2 | processed human plasma negative and positive for anti-HCV with preservatives | 60 | 28 |  |  |  | No | NA | NA |  | NA | NA | Yes | Atellica IM aHCV QC | 10995457 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 0.80 Index | 0.80 Index | 3.4 | 11.5 | 3.97 Index | 3.97 Index | 4.7 | 7.2 | 8.61 Index | 8.61 Index | 6.3 | 7.2 | NA | NA | ≤ 10% change up to up to 500 mg/dL of hemoglobin | ≤ 10% change up to 60 mg/dL of conjugated bilirubin 40 mg/dL of unconjugated bilirubin | ≤ 10% change up to 1000 mg/dL of triglycerides | Insignificant effect on the assay up to 1500 ng/mL | NA | NA |  |  | Based upon relative clinical agreement with commercially available anti-HCV assays | 10995360_EN Rev. 03, 2019-07 | IFU Document | 0.389 | 6.969 | 2.48 | 2.913 | 0.177 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995580 |  | re05 | 00630414599045 | Global | Reproductive Endocrinology | No | Atellica IM Follicle Stimulating Hormone (FSH) | Atellica IM FSH | FSH | 190 | 1 x 190 | Atellica IM CAL B | Atellica IM Multi‑Diluent 1 Atellica IM FSH MCM(SMN 10995582) | 14 | 28 | mouse monoclonal anti-FSH antibody | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | RT up to 8 hours > 8 hours at 2–8°C > 48 hr at ≤ -20°C Freeze 1 time | 50 μL | No |  | 2-site samdwich immonoassay | 0.30–200.00 mIU/mL | 0.30–200.00 IU/L |  |  |  |  | See IFU |  |  | Atellica IM CAL B | 10995503 (2‑pack) 10995504 (6‑pack) | 28 days recon at 2–8°C 4 hours at RT | Lyophilized | Add 5.0 mL of special reagent water into each vial Let the vials stand for 15–20 min at RT | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of digoxin,follicle-stimulating hormone (FSH), luteinizinghormone (LH), prolactin, human chorionic gonadotropin (hCG), and TSH; equine serum | 36 | 14 |  |  |  | Yes | Atellica IM Multi‑Diluent 1 | 10995637 (2‑pack) 10995638 (6‑pack) 10995639 (vial) |  | 0.30–400.00 mIU/mL | 0.30–400.00 IU/L | No | NA | NA | 0.02 mIU/mL | 0.02 IU/L | NA | NA | 0.55 mIU/mL | 0.20 mIU/mL | Not Performed | 0.20 IU/L | 0.55 IU/L | Not Performed | 1.75 mIU/mL | 1.75 IU/L | 6.3 | 8.2 | 50.78 mIU/mL | 50.78 IU/L | 1.5 | 2.9 | 169.96 mIU/mL | 169.96 IU/L | 2 | 2.9 | ADVIA Centaur FSH y = 0.98x + 1.92 mIU/mL r=1.00 | ADVIA Centaur FSH y = 0.98x + 1.92 IU/L r=1.00 | ≤ 10% change up to up to 150 mg/dL of hemoglobin | ≤ 10% change up to 20 mg/dL of conjugated bilirubin 20 mg/dL of unconjugated bilirubin | ≤ 10% change up to 500 mg/dL of lipemia | No Biotin in assay formulation | 1000 mIU/mL | 1000 IU/L |  |  | Traceable to World Health Organization (WHO) 2nd International Standard for human FSH (IS 94/632) | 10995346_EN Rev. 03, 2020-02 (OUS) 11200384_EN Rev. 03, 2020-02 (US) | IFU Document | 0.205 | 6.969 | 2.087 | 1.457 | 0.093 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995581 |  | re05 | 00630414599052 | Global | Reproductive Endocrinology | No | Atellica IM Follicle Stimulating Hormone (FSH) | Atellica IM FSH | FSH | 950 | 5 x 190 | Atellica IM CAL B | Atellica IM Multi‑Diluent 1 Atellica IM FSH MCM (SMN 10995582) | 14 | 28 | mouse monoclonal anti-FSH antibody | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | RT up to 8 hours > 8 hours at 2–8°C >48 hr at ≤ -20°C Freeze 1 time | 50 μL | No |  | 2-site samdwich immonoassay | 0.30–200.00 mIU/mL | 0.30–200.00 IU/L |  |  |  |  | See IFU |  |  | Atellica IM CAL B | 10995503 (2‑pack) 10995504 (6‑pack) | 28 days recon at 2–8°C 4 hours at RT | Lyophilized | Add 5.0 mL of special reagent water into each vial Let the vials stand for 15–20 min at RT | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of digoxin,follicle-stimulating hormone (FSH), luteinizinghormone (LH), prolactin, human chorionic gonadotropin (hCG), and TSH; equine serum | 36 | 14 |  |  |  | Yes | Atellica IM Multi‑Diluent 1 | 10995637 (2‑pack) 10995638 (6‑pack) 10995639 (vial) |  | 0.30–400.00 mIU/mL | 0.30–400.00 IU/L | No | NA | NA | 0.02 mIU/mL | 0.02 IU/L | NA | NA | 0.20 mIU/mL | 0.20 mIU/mL | Not Performed | 0.20 IU/L | 0.20 IU/L | Not Performed | 1.75 mIU/mL | 1.75 IU/L | 6.3 | 8.2 | 50.78 mIU/mL | 50.78 IU/L | 1.5 | 2.9 | 169.96 mIU/mL | 169.96 IU/L | 2 | 2.9 | ADVIA Centaur FSH y = 0.98x + 1.92 mIU/mL r=1.00 | ADVIA Centaur FSH y = 0.98x + 1.92 IU/L r=1.00 | ≤ 10% change up to up to 150 mg/dL of hemoglobin | ≤ 10% change up to 20 mg/dL of conjugated bilirubin 20 mg/dL of unconjugated bilirubin | ≤ 10% change up to 500 mg/dL of lipemia | No Biotin in assay formulation | 1000 mIU/mL | 1000 IU/L |  |  | Traceable to World Health Organization (WHO) 2nd International Standard for human FSH (IS 94/632) | 10995346_EN Rev. 03, 20120-02 (OUS), 11200384_EN Rev.03, 2020-02 (US) | IFU Document | 0.812 | 6.969 | 2.087 | 4.37 | 0.369 | 17.7 | 5.3 | 11.1
Atellica® IM | 10995635 |  | re02 | 00630414599533 | Global | Reproductive Endocrinology | No | Atellica IM Luteinizing Hormone (LH) | Atellica IM LH | LH | 110 | 1 x 110 | Atellica IM CAL B | Atellica IM Multi‑Diluent 1 Atellica IM LH MCM (SMN 10995636) | 14 | 28 | mouse monoclonal anti-LH antibody | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | RT up to 8 hours > 8 hours at 2–8°C >48 hr at ≤ -20°C Freeze 1 time | 25 μL | No |  | 2-site sandwich immunoassay | 0.07–200.00 mIU/mL | 0.07–200.00 IU/L |  |  |  |  | See IFU |  |  | Atellica IM CAL B | 10995503 (2‑pack) 10995504 (6‑pack) | 28 days recon at 2–8°C 4 hours at RT | Lyophilized | Add 5.0 mL of special reagent water into each vial Let the vials stand for 15–20 min at RT | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of digoxin,follicle-stimulating hormone (FSH), luteinizinghormone (LH), prolactin, human chorionicgonadotropin (hCG), and TSH; equine serum | 65 | 28 |  |  |  | Yes | Atellica IM Multi‑Diluent 1 | 10995637 (2‑pack) 10995638 (6‑pack) 10995639 (vial) |  | 0.07–400.00 mIU/mL | 0.07–400.00 IU/L | No | NA | NA | 0.06 mIU/mL | 0.06 IU/L | NA | NA | 1.44 mIU/mL | 0.07 mIU/mL | Not Performed | 0.07 IU/L | 1.44 IU/L | Not Performed | 18.74 mIU/mL | 18.74 IU/L | 1.8 | 3.4 | 63.64 mIU/mL | 63.64 IU/L | 1.4 | 2.2 | 177.25 mIU/mL | 177.25 IU/L | 1.1 | 2.1 | ADVIA Centaur LH y = 1.06x + 2.06 mIU/mL r=1.00 | ADVIA Centaur LH y = 1.06x + 2.06 IU/L r=1.00 | ≤ 10% change up to up to 500 mg/dL of hemoglobin | ≤ 10% change up to 20 mg/dL of conjugated bilirubin 20 mg/dL of unconjugated bilirubin | ≤ 10% change up to 3000mg/dL of lipemia | No Biotin in assay formulation | 18,000 mIU/mL | 18,000 IU/L |  |  | Traceable to World Health Organization (WHO) 2nd International Standard for human LH (IS 80/552). | 10995369_EN Rev. 03, 2020-02 (OUS) 11200385_EN Rev. 03, 2020-02 (US) | IFU Document | 0.213 | 6.969 | 2.087 | 1.457 | 0.097 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995634 |  | re02 | 00630414599526 | Global | Reproductive Endocrinology | No | Atellica IM Luteinizing Hormone (LH) | Atellica IM LH | LH | 550 | 5 x 110 | Atellica IM CAL B | Atellica IM Multi‑Diluent 1 Atellica IM LH MCM (SMN 10995636) | 14 | 28 | mouse monoclonal anti-LH antibody | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | RT up to 8 hours > 8 hours at 2–8°C >48 hr at ≤ -20°C Freeze 1 time | 25 μL | No |  | 2-site sandwich immunoassay | 0.07–200.00 mIU/mL | 0.07–200.00 IU/L |  |  |  |  | See IFU |  |  | Atellica IM CAL B | 10995503 (2‑pack) 10995504 (6‑pack) | 28 days recon at 2–8°C 4 hours at RT | Lyophilized | Add 5.0 mL of special reagent water into each vial Let the vials stand for 15–20 min at RT | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of digoxin,follicle-stimulating hormone (FSH), luteinizinghormone (LH), prolactin, human chorionic gonadotropin (hCG), and TSH; equine serum | 65 | 28 |  |  |  | Yes | Atellica IM Multi‑Diluent 1 | 10995637 (2‑pack) 10995638 (6‑pack) 10995639 (vial) |  | 0.07–400.00 mIU/mL | 0.07–400.00 IU/L | No | NA | NA | 0.06 mIU/mL | 0.06 IU/L | NA | NA | 1.44 mIU/mL | 0.07 mIU/mL | Not Performed | 0.07 IU/L | 1.44 IU/L | Not Performed | 18.74 mIU/mL | 18.74 IU/L | 1.8 | 3.4 | 63.64 mIU/mL | 63.64 IU/L | 1.4 | 2.2 | 177.25 | 177.25 IU/L | 1.1 | 2.1 | ADVIA Centaur LH y = 1.06x + 2.06 mIU/mL r=1.00 | ADVIA Centaur LH y = 1.06x + 2.06 IU/L r=1.00 | ≤ 10% change up to up to 500 mg/dL of hemoglobin | ≤ 10% change up to 20 mg/dL of conjugated bilirubin 20 mg/dL of unconjugated bilirubin | ≤ 10% change up to 3000mg/dL of lipemia | No Biotin in assay formulation | 18,000 mIU/mL | 18,000 IU/L |  |  | Traceable to World Health Organization (WHO) 2nd International Standard for human LH (IS 80/552). | 10995369_EN Rev. 03, 2020-02 (OUS) 11200385_EN Rev. 03, 2020-02 (US) | IFU Document | 0.851 | 6.969 | 2.087 | 4.37 | 0.387 | 17.7 | 5.3 | 11.1
Atellica® IM | 10995572 |  | an02 | 00630414598963 | OUS | Anemia | No | Atellica IM Folate (Fol) | Atellica IM FolFol | Fol | 140 | 1 x 140 | Atellica IM Fol DTT/REL (SMN 10995576) Atellica IM APW1 (SMN 10995458) | Atellica IM RBC Fol (SMN 10995440) Atellica IM Fol MCM (SMN 10995575) | 14 | 14 | Bovine and human serum albumi | ;#Serum;#Whole blood;# | Serum < 8 hrs at RT > 8 hrs at 2–8°C >48 hrs at ≤ -20°C | 100 μL | No |  | Competitive immunoassay using direct chemiluminescent technology | 0.35–24.00 ng/mL | 0.79–54.36 nmol/L | 0.35 - 48.00 ng/mL | 0.79 - 108.72 nmol/L | Serum Folate; Deficient 0.35–3.37 ng/mL Indeterminate 3.38–5.38 ng/mL Normal > 5.38 ng/mL | Serum Folate; Deficient 0.79–7.63 nmol/L Indeterminate 7.64–12.19 nmol/L Normal > 12.19 nmol/L |  |  |  | Atellica IM Fol CAL | Provided in Kit | 7 days at 2–8°C<br>8 hours at RT<br>28 days ≤ -20°C | Lyophilized | Add 3.0 mL of reagent water into each calibrator vial. Let the calibrators stand for 15–20 minutes at room temperature. | 3 for lot calibration, 2 for pack recalibration | 2 | Human serum albumin | 14 | 7 |  |  |  | Yes | Atellica IM Fol DIL | 10995574 |  | 0.35 - 48.00 ng/mL | 0.79 - 108.72 nmol/L | No | NA | NA | < 0.01 ng/mL | < 0.02 nmol/L | NA | NA | 0.38 ng/mL | 0.19 ng/mL | NA | 0.43 nmol/L | 0.86 nmol/L | NA | 4.13 ng/mL | 9.35 nmol/L | 2.4 | 5.9 | 9.23 ng/mL | 20.91 nmol/L | 2.6 | 6.5 | 21.91 ng/mL | 49.63 nmol/L | 3.3 | 7.1 | ADVIA Centaur Folate y = 0.94x - 0.01 ng/mL r=0.99 | ADVIA Centaur Folate y = 0.94x - 0.02 nmol/L r=0.99 | NA | ≤ 10% change up to 20 mg/dL conjugated bilirubin, 20 mg/dL unconjugated bilirubin | ≤ 10% change up to 2000 mg/dL Lipemia | ≤ 10% change up to 38 ng/mL of biotin | NA | NA | NA | NA | Traceable to an internal standard manufactured using highly purified material (N‑5‑methyl tetrahydrofolate) | 10995344_EN Rev. 03, 2019-07 | IFU Document | 0.411 | 6.969 | 2.48 | 2.913 | 0.187 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995573 |  | an02 | 00630414598970 | OUS | Anemia | No | Atellica IM Folate (Fol) | Atellica IM FolFol | Fol | 700 | 5 x 140 | Atellica IM Fol DTT/REL (SMN 10995576) Atellica IM APW1 (SMN 10995458) | Atellica IM RBC Fol (SMN 10995440) Atellica IM Fol MCM (SMN 10995575) | 14 | 14 | Bovine and human serum albumi | ;#Serum;#Whole blood;# | Serum < 8 hrs at RT >8 hrs at 2–8°C >48 hrs at ≤ -20°C | 100 μL | No |  | Competitive immunoassay using direct chemiluminescent technology | 0.35–24.00 ng/mL | 0.79–54.36 nmol/L | 0.35 - 48.00 ng/mL | 0.79 - 108.72 nmol/L | Serum Folate; Deficient 0.35–3.37 ng/mL Indeterminate 3.38–5.38 ng/mL Normal > 5.38 ng/mL | Serum Folate; Deficient 0.79–7.63 nmol/L Indeterminate 7.64–12.19 nmol/L Normal > 12.19 nmol/L |  |  |  | Atellica IM Fol CAL | Provided in Kit | 7 days at 2–8°C<br>8 hours at RT<br>28 days ≤ -20°C | Lyophilized | Add 3.0 mL of reagent water into each calibrator vial. Let the calibrators stand for 15–20 minutes at room temperature. | 3 for lot calibration, 2 for pack recalibration | 2 | Human serum albumin | 14 | 7 |  |  |  | Yes | Atellica IM Fol DIL | 10995574 |  | 0.35-48.00 ng/mL | 0.79-108.72 nmol/L | No | NA | NA | < 0.01 ng/mL | < 0.02 nmol/L | NA | NA | 0.38 ng/mL | 0.19 ng/mL | NA | 0.43 nmol/L | 0.86 nmol/L | NA | 4.13 ng/mL | 9.35 nmol/L | 2.4 | 5.9 | 9.23 ng/mL | 20.91 nmol/L | 2.6 | 6.5 | 21.91 (ng/mL) | 49.63 (nmol/L) | 3.3 | 7.1 | ADVIA Centaur Folate y = 0.94x - 0.01 ng/mL r=0.99 | ADVIA Centaur Folate y = 0.94x - 0.02 nmol/L r=0.99 | NA | ≤ 10% change up to 20 mg/dL conjugated bilirubin, 20 mg/dL unconjugated bilirubin | ≤ 10% change up to 2000 mg/dL Lipemia | ≤ 10% change up to 38 ng/mL of biotin | NA | NA | NA | NA | Traceable to an internal standard manufactured using highly purified material (N‑5‑methyl tetrahydrofolate) | 10995344_EN Rev. 03, 2019-07 | IFU Document | 1.208 | 6.969 | 2.126 | 6.811 | 0.549 | 17.7 | 5.4 | 17.3
Atellica® IM | 10995621 |  | bo02 | 00630414599441 | OUS | Bone Metabolism | No | Atellica IM Intact Parathyroid Hormone (PTH) | Atellica IM PTH | PTH | 190 |  | NA | Atellica IM PTH QC (10995626) Atellica IM Multi‑Diluent 13 (10995643) Atellica IM PTH MCM (SMN 10995625) | 10 | 28 | Two mouse monoclonal anti-human PTH antibodies | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#EDTA Plasma;# | Serum: 25°C 8 hours 2–8°C 8 hours -20°C 1 month Plasma EDTA: 25°C 25 hours 2–8°C 14 days -20°C Not recommended Plasma Lithium Heparin & Sodium Heparin: 25°C 9 hours 2–8°C 72 hours -20°C Not recommended | 25 μL | No |  | 2-site sandwich immunoassay using direct chemiluminescent technology | 6.3–2000.0 pg/mL (US); 4.6-2000 pg/mL (OUS) | 0.7–212.0 pmol/L (US); 0.5-212.0 pmol/L (OUS) |  |  |  |  | Plasma: 18.4–80.1 pg/mL (1.95–8.49 pmol/L) | Serum: 18.5–88.0 pg/mL (1.96–9.33 pmol/L) |  | IM PTH CAL low | Provided in Kit | 60 days recon at ≤ -20°C 8 hours recon at RT | Lyophilized | Reconstitute to 1mL/vial | 3 for lot calibration, 2 for pack recalibration | 2 | Human plasma-based product | 91 | 14 |  |  |  | Yes | Atellica IM Muti-Diluent 13 | 10995643 |  | 6.3–10,000pg/mL (US); 4.6-10,000pg/mL (OUS) | 0.7 - 8480 pmol/L; 0.5-8480 pmol/L (OUS) | Yes | Atellica IM PTH QC | 10995626 (2x1.0mL) | NA | NA | NA | NA | 1.5 pg/mL | 1.4 pg/mL | 2.3 pg/mL | 0.15 pmol/L | 0.16 pmol/L | 0.24 pmol/L | 42.0 pg/mL (Control 1 | 4.4.pmol/L | 4.6 | 4.7 | 240.7 pg/mL (Control 2) | 25.5 pmol/L (Control 2) | 1.7 | 2.9 | 868.7 pmol/L (Control 3) | 92.1 pg/mL (Control 3) | 1.4 | 2.3 | ADVIA Centaur PTH y = 1.03x - 1.08 pg/mL r=1.00 | ADVIA Centaur PTH y = 1.03x - 0.11 pmol/L r=1.00 | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 60 mg/dL of conjugated bilirubin 60 mg/dL of unconjugated bilirubin | ≤ 10% change up to 1000 mg/dL of Lipemia | ≤ 10% change up to 1500 ng/mL | 100,000 pg/mL | 10,600 pmol/L |  |  | The Atellica IM PTH assay standardization is maintained with internal standards using purified human PTH (1-84) | 10995367_EN Rev. 04, 2021-01-27 | IFU Document | 0.308 | 6.969 | 2.087 | 1.457 | 0.14 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995622 |  | bo02 | 00630414599458 | OUS | Bone Metabolism | No | Atellica IM Intact Parathyroid Hormone (PTH) | Atellica IM PTH | PTH | 950 |  | NA | Atellica IM PTH QC (10995626) Atellica IM Multi‑Diluent 13 (10995643) Atellica IM PTH MCM (SMN 10995625) | 10 | 28 | Two mouse monoclonal anti-human PTH antibodies | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#EDTA Plasma;# | Serum: 25°C 8 hours 2–8°C 8 hours -20°C 1 month Plasma EDTA: 25°C 25 hours 2–8°C 14 days -20°C not recommended Plasma Lithium Heparin & Sodium Heparin: 25°C 9 hours 2–8°C 72 hours -20°C not recommended | 25 μL | No |  | 2-site sandwich immunoassay using direct chemiluminescent technology | 6.3–2000.0 pg/mL (US); 4.6-2000.0 pg/mL (OUS) | 0.7–212.0 pmol/L (US); 0.5-212 pmol/L (OUS) |  |  |  |  | Plasma: 18.4–80.1 pg/mL (1.95–8.49 pmol/L) | Serum: 18.5–88.0 pg/mL (1.96–9.33 pmol/L) |  | IM PTH CAL low | Provided in Kit | 60 days recon at ≤ -20°C 8 hours recon at RT | Lyophilized | Reconstitute to 1.0 mL/vial | 3 for lot calibration, 2 for pack recalibration | 2 | Human plasma-based product | 91 | 14 |  |  |  | Yes | Atellica IM Multi-Diluent 13 | 10995643 |  | 6.3–10,000pg/mL (US); 4.6-10,000 pg/mL (OUS) | 0.7 - 1060 pmol/L (US); 0.5-1060 pmol/L (OUS) | Yes | Atellica IM PTH QC | 10995626 (2x1.0mL) | NA | NA | NA | NA | 1.5 pg/mL | 1.4 pg/mL | 2.3 pg/mL | 0.15 pmol/L | 0.16 pmol/L | 0.24 pmol/L | 42.0 pg/mL (Control 1) | 4.4 pmol/L | 4.6 | 4.7 | 240.7 pg/mL (Control 2) | 25.5 pmol/L (Control) | 1.7 | 2.9 | 868.7 pmol/L (Control 3) | 92.1 pg/mL (Control) | 1.4 | 2.3 | ADVIA Centaur PTH y = 1.03x - 1.08 pg/mL r=1.00 | ADVIA Centaur PTH y = 1.03x - 0.11 pmol/L r=1.00 | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 60 mg/dL of conjugated bilirubin 60 mg/dL of unconjugated bilirubin | ≤ 10% change up to 1000 mg/dL of Lipemia | ≤ 10% change up to 1500 ng/mL | 100,000 pg/mL | 10,600 pmol/L |  |  | The Atellica IM PTH assay standardization is maintained with internal standards using purified human PTH (1-84) | 10995367_EN Rev. 04, 2021-01-27 | IFU Document | 0.999 | 6.969 | 2.087 | 4.37 | 0.454 | 17.7 | 5.3 | 11.1
Atellica® IM | 11202699 |  | se01 | 00630414246550 | OUS | Sepsis | No | Atellica IM BRAHMS Procalcitonin (PCT) | Atellica IM BRAHMS PCTPCT | PCT | 100 | 1 X 100 | NA | Atellica IM BRAHMS PCT MCM (SMN 10998265) | 18 | 60 | Mouse monoclonal anti-PCT antibody | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#EDTA Plasma;# | < 8 hrs at RT, > 8 hrs at 2–8°C, >48 hrs at ≤ -20°C, 5 freeze thaw cycles | 100 μL | No |  | 2-site sandwich immunoassay using direct chemiluminescent technology | 0.03–50.00 ng/mL | 0.03–50.00 ng/mL | 0.03–1000.00 ng/mL | 0.03–1000.00 ng/mL | < 0.05 ng/mL | < 0.05 ng/mL |  |  |  | Atellica IM BRAHMS PCT CAL | Included in the kit | 8 hours at RT; 24 hrs at 2–8°C, 60 days at ≤ -20°C; thaw up to 2x | Lyophilized | Recon with 2.0 mL of special reagent water, let stand for 15–20 minutes at RT | 3 for lot calibration, 2 for pack recalibration | 2 | Recombinant PCT; equine serum | 82 | 35 |  |  |  | Yes | Atellica IM Multi‑Diluent 1 | 10995637 (2‑pack) 10995638 (6‑pack) 10995639 (vial) |  | 0.03–1000.00 ng/mL | 0.03–1000.00 ng/mL | Yes | Atellica IM BRAHMS PCT QC | 11202700 | NA | NA | NA | NA | 0.03 ng/mL | 0.00 ng/mL | 0.04 ng/mL | 0.00 ng/mL | 0.03 ng/mL | 0.04 ng/mL | 0.05 ng/mL | 0.05 ng/mL | 9.7 | 12.1 | 1.52 ng/mL | 1.52 ng/mL | 1.4 | 2.6 | 19.14 ng/mL | 19.14 ng/mL | 1.5 | 2.2 | B·R·A·H·M·S PCT sensitive Kryptor y= 1.02 x - 0.02 ng/mL r= 0.98; N= 522 | B·R·A·H·M·S PCT sensitive Kryptor y= 1.02 x - 0.02 ng/mL r= 0.98; N= 522 | 500 mg/dL at conc 0.36 ng/mL -3.6% bias, at conc 1.73 ng/mL -3.4% bias | Conjugted Bili 40 mg/dL at conc 0.34 ng/mL -3.1% bias at conc 1.61 ng/mL -1.6% bias Unconjugated 40 mg/dL at conc 0.34 ng/mL 6.6% bias, at conc 1.67 ng/mL 3.8% bias | 1000 mg/dL at conc 0.35 ng/mL -2.9% bias, at conc 1.62 ng/mL 1.4% bias | No Biotin in assay formulation | 2000 ng/mL | 2000 ng/mL | 2000 ng/mL | 2000 ng/mL | Specified by BRAHMS with N-terminal amino acid sequencing (Edmann’s method) and Mass analysis | 11202687_EN Rev. 03, 2019-07 | IFU Document | 0.343 | 6.969 | 2.48 | 2.913 | 0.156 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995478 |  | on08 | 00630414598055 | Global | Oncology | No | Atellica IM BR 27.29 (BR) | Atellica IM BR | BR | 50 | 1x50 | Atellica IM BR CAL (SMN 10995515) Atellica IM BR PRE (SMN 10995480) | Atellica IM Multi‑Diluent 1 (SMN 10995637, 2-pack; SMN 10995638, 6-pack; SMN 10995639, vial) Atellica IM BR MCM (SMN 10995479) | 14 | 28 | Mouse monoclonal anti-CA 27.29 antibody | ;#Serum;#EDTA Plasma;# | < 8 hrs at RT > 8 hrs at 2–8°C >48 hrs at ≤ -20°C | 25 μL | Yes |  | Competitive immunoassay using direct chemiluminescent technology | 3.5–450.0 U/mL | NA |  |  |  |  | < 38.6 U/mL | See IFU for other results for other conditions |  | Atellica IM BR CAL | 10995515 (2‑pack) | 28 days recon at 2–8°C 32 hours recon at RT | Lyophilized | 2.0 mL of special reagent water | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of CA 27.29 (human) | 32 | 10 |  |  |  | Yes | Atellica IM Multi‑Diluent 1 | 10995637 (2‑pack) 10995638 (6‑pack) 10995639 (vial) |  | 3.5–9000 U/mL | NA | No | NA | NA | < 0.1 U/mL | NA | NA | NA | 4.1 U/mL | 1.0 U/mL | Not Performed | NA | NA | NA | 31.2 U/mL (Control) | NA | 4.5 | 7.3 | 154.6 U/mL (Control) | NA | 1.7 | 3.3 | 317.1 U/mL | NA | 2.1 | 4.1 | ADVIA Centaur BR y = 1.02x + 0.1 U/mL r=1.00 | NA | ≤ 5% change up to 500 mg/dL hemoglobin | ≤ 5% change up to 20 mg/dL conjugated bilirubin, 20 mg/dL unconjugated bilirubin | ≤ 5% change up to 3000 mg/dL Triglycerides | No Biotin in assay formulation | NA | NA |  |  | Traceable to an internal standard manufactured using highly purified CA 27.29 | 10995282_EN Rev. 03, 2019-08 | IFU Document | 0.218 | 6.969 | 2.087 | 1.457 | 0.099 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995477 |  | on08 | 00630414598048 | Global | Oncology | No | Atellica IM BR 27.29 (BR) | Atellica IM BR | BR | 250 | 5x50 | Atellica IM BR CAL (SMN 10995515) Atellica IM BR PRE (SMN 10995480) | Atellica IM Multi‑Diluent 1 (SMN 10995637, 2-pack; SMN 10995638, 6-pack; SMN 10995639, vial) Atellica IM BR MCM (SMN 10995479) | 14 | 28 | Mouse monoclonal anti-CA 27.29 antibody | ;#Serum;#EDTA Plasma;# | < 8 hrs at RT > 8 hrs at 2–8°C >48 hrs at ≤ -20°C | 25 μL | Yes |  | Competitive immunoassay using direct chemiluminescent technology | 3.5–450.0 U/mL | NA |  |  |  |  | < 38.6 U/mL | See IFU for other results for other conditions |  | Atellica IM BR CAL | 10995515 (2‑pack) | 28 days recon at 2–8°C 32 hours recon at RT | Lyophilized | 2.0 mL of special reagent water | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of CA 27.29 (human) | 32 | 10 |  |  |  | Yes | Atellica IM Multi‑Diluent 1 | 10995637 (2‑pack) 10995638 (6‑pack) 10995639 (vial) |  | 3.5–9000 U/mL | NA | No | NA | NA | < 0.1 U/mL | NA | NA | NA | 4.1 U/mL | 1.0 U/mL | Not Performed | NA | NA | NA | 31.2 U/mL (Control) | NA | 4.5 | 7.3 | 154.6 U/mL (Control) | NA | 1.7 | 3.3 | 317.1 U/mL | NA | 2.1 | 4.1 | ADVIA Centaur BR y = 1.02x + 0.1 U/mL r=1.00 | NA | ≤ 5% change up to 500 mg/dL hemoglobin | ≤ 5% change up to 20 mg/dL conjugated bilirubin, 20 mg/dL unconjugated bilirubin | ≤ 5% change up to 3000 mg/dL Triglycerides | No Biotin in assay formulation | NA | NA |  |  | Traceable to an internal standard manufactured using highly purified CA 27.29 | 10995282_EN Rev. 03, 2019-08 | IFU Document | 0.713 | 6.969 | 2.087 | 4.37 | 0.324 | 17.7 | 5.3 | 11.1
Atellica® IM | 10995649 |  | ca07 | 00630414599670 | Global | Cardiac | No | Atellica IM Myoglobin (MYO) | Atellica IM MYO | MYO | 50 | 1 X 100 | Atellica IM CAL U (SMN 10995519) | Atellica IM Multi‑Diluent 10 (SMN 10995640) Atellica IM MYO MCM (SMN 10995650) | 10 | 28 | Goat polyclonal anti-myoglobin antibody / mouse monoclonal anti-myoglobin antibody | ;#Serum;#Heparinized Plasma;# | Up to 8 hrs at RT, 48hr 2-8C, freeze 1 time | 10 μL | No |  | direct chemiluminescent technology | 3.00–1000.00 ng/mL | 3.00–1000.00 μg/L |  |  |  |  | A myoglobin value of >110 ng/mL (ug/L) is highly suggestive of myocardial infarct |  |  | Atellica IM CAL U | 10995519 (2‑pack) | 4 hrs at RT, 24 hrs recon at 2–8°C, 60 days recon at -80 to -60°C | Lyophilized | reconstitute 2.0mL / Vial | 3 for lot calibration, 2 for pack recalibration | 2 |  | 33 | 14 |  |  |  | Yes | Atellica IM Multi‑Diluent 10 | 10995640 |  | 3.00–20,000 ng/mL | 3.00-20,000.00 μg/L | No | NA | NA | 2.53 ng/mL | 2.53 μg/L | NA | NA | 3.57 ng/mL | 2.56 ng/mL | NA | 2.56 μg/L | 3.57 μg/L | NA | 155.95 mg/dL (Control) | 155.95 μg/L (Control) | 2 | 2.9 | 361.51 mg/dL (Control) | 361.51 μg/L (Control) | 2.1 | 3 | 832.09 mg/dL | 832.09 μg/L | 2.2 | 2.9 | ADVIA Centaur Myoglobin y = 0.96x + 2.80 ng/mL r=1.00 | ADVIA Centaur Myoglobin y = 0.96x + 2.80 μg/L r=1.00 | ≤ 5% change up to 1000 mg/dL of hemoglobin | ≤ 5% change up to 40 mg/dL of conjugated bilirubin 40 mg/dL of unconjugated bilirubin | ≤ 5% change up to 1000 mg/dL of triglycerides | No Biotin in assay formulation | 200,000 ng/mL | 200,000 μg/L |  |  | Traceable to an internal standard manufactured using highly purified material | 10995411_EN Rev. 02, 2019-07 | IFU Document | 0.176 | 6.969 | 2.087 | 1.457 | 0.08 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995555 |  | re09 | 00630414598796 | Global | Reproductive Endocrinology | No | Atellica IM DHEA‑SO4 (DHEAS) | Atellica IM DHEAS | DHEAS | 50 | 1 x 50 | NA | Atellica IM Multi‑Diluent 1 Atellica IM DHEAS MCM (SMN 10995556) | 14 | 16 | DHEAS (~280 pg/mL) labeled with acridinium ester, Biotinylated mouse monoclonal anti-DHEAS antibody | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | RT up to 4 hours > 4 hours, at 2–8°C > 6 days, at ≤ -20°C for 1 month, Freeze 1 time | 25 μL | No |  | Quantitative competitive immunoassay | 3.00–1500.00 μg/dL | 0.08–40.71 μmol/L |  |  |  |  | See IFUs |  |  | Atellica IM DHEAS CAL | Provided in Kit | at 2–8°C until expiration 4 hours at RT | Liquid |  | 3 for lot calibration, 2 for pack recalibration | 2 | DHEAS; human serum | 70 | 14 |  |  |  | Yes | Atellica IM Multi‑Diluent 1 | 10995637 (2‑pack) 10995638 (6‑pack) 10995639 (vial) |  | 3.00–3000.00 μg/dL | 0.08–81.42 μmol/L | No | NA | NA | NA | NA | NA | NA | 4.27 μg/dL | 2.35 μg/dL | NA | 0.06 μmol/L | 0.12 μmol/L | NA | 74.56 μg/dL | 2.02 μmol/L | 2.3 | 5.6 | 636.71 μg/dL | 17.28 μmol/L | 2.2 | 4.4 | 1236.97 μg/dL | 33.57 μmol/L | 2.5 | 4.5 | ADVIA Centaur DHEAS (x) y = 1.07x - 1.48 μg/dL r=0.99 | ADVIA Centaur DHEAS y = 1.07x - 0.04 μmol/L r=0.99 | ≤ 10% change up to 500 mg/dL hemoglobin | ≤ 10% change up to 5 mg/dL conjugated bilirubin 5 mg/dL unconjugated bilirubin | ≤ 10% change up to 1000 mg/dL triglycerides | ≤ 10% change in results up to 12.5 ng/mL of biotin | NA | NA |  |  | Traceable to an internal standard manufactured using purified DHEAS | 10995340_EN Rev. 04, 2019-07 (OUS); 11200382_EN Rev. 04, 2019-06 (US) | IFU Document | 0.317 | 6.969 | 2.48 | 2.913 | 0.144 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995591 |  | on09 | 00630414599151 | Global | Oncology | No | Atellica IM HER‑2/neu (H2n) | Atellica IM H2n | H2n | 50 | 1 x 50 | Atellica IM H2n CAL (SMN 10995592, 2-pack) | Atellica IM H2n QC (SMN 10995494) Atellica IM Multi‑Diluent 1 (SMN 10995637, 2-pack; SMN 10995638, 6-pack; SMN10995639, vial) Atellica IM H2n MCM (SMN 10995593) | 14 | 41 | Mouse monoclonal anti-human HER‑2/neu antibody | ;#Serum;# | Up to 8 hrs at RT > 8 hrs at 2–8°C >24 hrs at ≤ -20°C | 20 μL | No |  | 2-site sandwich immunoassay using direct chemiluminescent technology | 0.5–350.0 ng/mL | NA |  |  |  |  | ULN defined as the 95th percentile of the observed results, was 15.2 ng/mL. | See IFU for information on monitoring metastatic breast cancer | HER‑2/neu Changes and Clinical Status | Atellica IM H2n CAL | 10995592 (2‑pack) | 60 days open at 2–8°C 8 hours at RT | Liquid | N/A | 3 for lot calibration, 2 for pack recalibration | 2 | Low or high levels of HER‑2/neu | 58 | 14 |  |  |  | Yes | Atellica IM Multi‑Diluent 1 | 10995637 (2‑pack) 10995638 (6‑pack) 10995639 (vial) |  | 0.5–7000.0 ng/mL | NA | Yes | Atellica IM H2n QC | 10995594 | < 0.1 ng/mL | NA | NA | NA | 0.0 ng/mL | < 0.1 ng/mL | NA | NA | NA | NA | 15.2 ng/mL (Control) | NA | 2.1 | 2.9 | 97.8 ng/mL | NA | 1.7 | 2.9 | 333.9 ng/mL | NA | 1.3 | 3 | ADVIA Centaur H2n y = 0.98x - 0.11ng/mL r=1.00 | NA | Insignificant effect up to 1000 mg/dL hemoglobin | Insignificant effect up to 25 mg/dL conjugated bilirubin, 25 mg/dL unconjugated bilirubin | Insignificant effect up to 900 mg/dL Triglycerides | No Biotin in assay formulation | 20,000 ng/mL | NA |  |  | Traceable to an internal standard manufactured using highly purified material | 10995349_EN Rev. 02, 2019-07 | IFU Document | 0.183 | 6.969 | 2.087 | 1.457 | 0.083 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995548 |  | is01 | 00630414598727 | Global | Immunosupressant Drugs | No | Atellica IM Cyclosporine (CsA) | Atellica IM CsA | CsA | 50 |  | Atellica IM CsA CAL Atellica IM CsA PRE (10995552) | Atellica IM Multi‑Diluent 12 Atellica IM CsA MCM (10995551) | 14 | 42 |  | ;#Whole blood;# | 6hr at RT; 7 days at 2-8C; frozen at -20C up to 1 month | 100 μL | Yes |  |  | 30.00–1500.00 ng/mL | 24.95–1247.25 nmol/L |  |  |  |  | See IFU |  |  | Atellica IM CsA CAL | 10995549 | 21 weeks at 2–8°C 4 hours at RT | Liquid |  | 3 for lot calibration, 2 for pack recalibration | 2 |  | 32 | 28 |  |  |  | Yes | Atellica IM Multi‑Diluent 12 | 10995550 |  | 30.00–7500.00 ng/mL | 24.95–6236.25 nmol/L | No | NA | NA | NA | NA | 29.40 ng/mL | 24.45 nmol/L | 13.50 ng/mL | 8.60 ng/mL | NA | 7.15 nmol/L | 11.23 nmol/L | NA | 127.14 ng/mL | 105.72 nmol/L | 2.4 | 7.6 | 284.55 ng/mL | 236.60 nmol/L | 3.3 | 5.6 | 640.89 ng/mL | 532.90 nmol/L | 2.7 | 6.5 | ADVIA Centaur CsA y = 0.97x - 4.23 ng/mL r=0.99 | ADVIA Centaur CsA y = 0.97x - 3.52 nmol/L r=0.99 | NA | ≤ 10% change up to 60 mg/dL conjugated bilirubin 40 mg/dL unconjugated bilirubin | ≤ 10% change up to 900 mg/dL triglycerides 300 mg/dL cholesterol | ≤ 10% change up to 50 ng/mL | NA | NA |  |  | Traceable to an internal standard manufactured using highly purified cyclosporine (USP grade) | 10995338_EN Rev. 04, 2019-07 | IFU Document | 0.196 | 6.969 | 2.087 | 1.457 | 0.089 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995695 |  | li02 | 00630414600086 | OUS | Liver Fibrosis | No | Atellica IM Tissue Inhibitor of Matrix Metalloproteinase 1 (TIMP‑1) | Atellica IM TIMP-1TIMP-1 | TIMP-1 | 50 | 1 x 50 |  | N/A | 14 | 60 | Mouse monoclonal antibodies (MAb) | ;#Serum;# | 48 hr at RT, 7 days at 2-8C. Freeze samples at ≤ -20°C if the sample is not assayed within 7 days | 25 μL | No |  | 2-site sandwich immunoassay using direct chemiluminescent technology | 3.5–1300.0 ng/mL | 3.5–1300.0 ng/mL | 3.5–6500.0 ng/mL | 3.5–6500.0 ng/mL | Refer to the Atellica IM ELF Test instructions for use for ELF score expected values | Refer to the Atellica IM ELF Test instructions for use for ELF score expected values |  |  |  | Atellica IM ELF CAL | 10995565 | 8 hours at RT; 5 days at 2–8°C; 90 days at -20°C | Lyophilized | Add 2.0 mL of reagent water into each calibrator vial, Let the calibrators stand for 15–20 minutes at RT, Gently swirl | 3 for lot calibration, 2 for pack recalibration | 2 | low or high levels of HA, PIIINP (bovine), TIMP-1 with bovine serum albumin, buffer, sodium azide and preservatives | 64 | 60 |  |  |  | Yes | Atellica IM Multi‑Diluent 10 | 10995640 |  | 3.5–6500.0 ng/mL | 3.5–6500.0 ng/mL | Yes | Atellica IM ELF QC | 10995566 | NA | NA | NA | NA | 1.1 ng/mL | 0.4 ng/mL | NA | 0.4 ng/mL | 1.1 ng/mL | NA | 81.3 ng/mL | 81.3 ng/mL |  |  | 254.5 ng/mL | 254.5 ng/mL | 2.7 | 3.3 | 533.5 ng/mL | 533.5 ng/mL | 2.8 | 3.9 | ADVIA Centaur TIMP-1 y=0.94x + 4.6 ng/mL r=1.00; N=112 | ADVIA Centaur TIMP-1 y=0.94x + 4.6 ng/mL r=1.00; N=112 | < 10% change up to 500 mg/dL hemoglobin | < 10% change up to 40 mg/dL conjugated bilirubin 40 mg/dL unconjugated bilirubin | < 10% change up to 1000 mg/dL triglycerides | No Biotin in assay formulation | 30,000 ng/mL | 30,000 ng/mL | 30,000 ng/mL | 30,000 ng/mL | Traceable to an internal standard manufactured using highly purified material | 10995427_EN Rev. 02, 2019-08 | IFU Document | 0.202 | 6.969 | 2.087 | 1.457 | 0.092 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995653 |  | li03 | 00630414599717 | OUS | Liver Fibrosis | No | Atellica IM N‑terminal Propeptide of Type III Procollagen (PIIINP) | Atellica IM PIIINPPIIINP | PIIINP | 50 | 1 x 50 | Atellica IM PW3 (SMN 10995666) | N/A | 14 | 30 | Mouse monoclonal antibodies (MAb) | ;#Serum;# | 48 hr at RT, 7 days at 2-8C. Freeze samples at ≤ -20°C if the sample is not assayed within 7 days | 20 | No |  | 2-site sandwich immunoassay using direct chemiluminescent technology | 0.50–150.00 ng/mL | 0.50–150.00 ng/mL | 0.50–750.00 ng/mL | 0.50–750.00 ng/mL | Refer to the Atellica IM ELF Test instructions for use for ELF score expected values | Refer to the Atellica IM ELF Test instructions for use for ELF score expected values |  |  |  | Atellica IM ELF CAL | 10995565 | 8 hours at RT; 5 days at 2–8°C; 90 days at -20°C | Lyophilized | Add 2.0 mL of reagent water into each calibrator vial, Let the calibrators stand for 15–20 minutes at RT, Gently swirl | 3 for lot calibration, 2 for pack recalibration | 2 | low or high levels of HA, PIIINP (bovine), TIMP-1 with bovine serum albumin, buffer, sodium azide and preservatives | 63 | 30 |  |  |  | Yes | Atellica IM Multi‑Diluent 13 | 10995643 |  | 0.50–750.00 ng/mL | 0.50–750.00 ng/mL | Yes | Atellica IM ELF QC | 10995566 | NA | NA | NA | NA | 0.35 ng/mL | 0.08 ng/mL | NA | 0.08 ng/mL | 0.35 ng/mL | NA | 3.31 ng/mL | 3.31 ng/mL | 4.9 | 5.7 | 6.63 ng/mL | 6.63 ng/mL | 2.7 | 3.5 | 12.1 ng/mL | 12.1 ng/mL | 2.4 | 3.3 | ADVIA Centaur PIIINP y = 0.97x + 0.09 ng/mL ; r=1.00, N= 138 | ADVIA Centaur PIIINP y = 0.97x + 0.09 ng/mL ; r=1.00, N= 138 | < 10% change up to 500 mg/dL hemoglobin | < 10% change up to 40 mg/dL conjugated bilirubin 40 mg/dL unconjugated bilirubin | < 10% change up to 1000 mg/dL triglycerides | <10% change up to 50 ng/mL of biotin | 2500 ng/mL | 2500 ng/mL | 2500 ng/mL | 2500 ng/mL | Traceable to an internal standard manufactured using bovine PIIINP | 10995413_EN Rev. 03, 2019-07 | IFU Document | 0.2 | 6.969 | 2.087 | 1.457 | 0.091 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995595 |  | li01 | 00630414599199 | OUS | Liver Fibrosis | No | Atellica IM Hyaluronic Acid (HA) | Atellica IM HAHA | HA | 50 | 1 x 50 |  | N/A | 39 | 28 | Mouse monoclonal antibodies (MAb) | ;#Serum;# | 48 hr at RT, 7 days at 2-8C. Freeze samples at ≤ -20°C if the sample is not assayed within 7 days | 20 μL | No |  | 2-site sandwich assay using direct chemiluminescent technology | 1.60–1000.00 ng/mL | 1.60–1000.00 ng/mL | 1.60 – 5000.00 ng/mL | 1.60 – 5000.00 ng/mL | Refer to the Atellica IM ELF Test instructions for use for ELF score expected values | Refer to the Atellica IM ELF Test instructions for use for ELF score expected values |  |  |  | Atellica IM ELF CAL | 10995565 | 8 hours at RT; 5 days at 2–8°C; 90 days at -20°C | Lyophilized | Add 2.0 mL of reagent water into each calibrator vial, Let the calibrators stand for 15–20 minutes at RT, Gently swirl | 3 for lot calibration, 2 for pack recalibration | 2 | low or high levels of HA, PIIINP (bovine), TIMP-1 with bovine serum albumin, buffer, sodium azide and preservatives | 24 | 14 |  |  |  | Yes | Atellica IM Multi‑Diluent 13 | 10995643 |  | 1.60 – 5000.00 ng/mL | 1.60 – 5000.00 ng/mL | Yes | Atellica IM ELF QC | 10995566 | NA | NA | NA | NA | 1.10 ng/mL | 0.56 ng/mL | NA | 0.56 ng/mL | 1.10 ng/mL | NA | 21.66 ng/mL | 21.66 ng/mL | 3.1 | 4.3 | 54.92 ng/mL | 54.92 ng/mL | 3 | 3.8 | 194.17 ng/mL | 194.17 ng/mL | 3 | 4.6 | ADVIA Centaur HA y = 0.98x + 1.89 ng/mL r=1.00, N= 128 | ADVIA Centaur HA y = 0.98x + 1.89 ng/mL r=1.00, N= 128 | < 10% change up to 500 mg/dL hemoglobin | < 10% change up to 40 mg/dL conjugated bilirubin 40 mg/dL unconjugated bilirubin | < 10% change up to 1000 mg/dL triglycerides | No Biotin in assay formulation | 25,000 ng/mL | 25,000 ng/mL | 25,000 ng/mL | 25,000 ng/mL | Traceable to an internal standard manufactured using highly purified material | 10995350_EN Rev. 02, 2019-07 | IFU Document | 0.198 | 6.969 | 2.087 | 1.457 | 0.09 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995672 |  | re08 | 00630414599854 | Global | Reproductive Endocrinology | No | Atellica IM Sex Hormone Binding Globulin (SHBG) | Atellica IM SHBG | SHBG | 50 | 1 x 50 | Atellica IM SHBG CAL Atellica IM Multi‑Diluent 1 | Atellica IM Multi‑Diluent 1 (SMN 10995639) Atellica IM SHBG MCM (SMN 10995674) | 14 | 60 | Mouse monoclonal anti-SHBG antibody, | ;#Serum;#Lithium Heparin Plasma;# | RT up to 4 hours > 4 hours at 2–8°C > 6 days at ≤ -20°C for 1 month Freeze 1 time | 10 μL | No |  | Sandwich immunoassay | 1.60–180.00 nmol/L | 0.15–17.10 μg/mL |  |  |  |  | See IFU |  |  | Atellica IM SHBG CAL | 10995673(2‑pack) | 60 days recon at ≤ -20°C thaw 1 time | Lyophilized | Add 2.0 mL of special reagent water into each vial. Let the vials stand for 15–20 minutes at room temperature | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, low or high levels of SHBG (human); equine serum | 58 | 35 |  |  |  | Yes | Atellica IM Multi‑Diluent 1 | 10995637(2‑pack) 10995638(6‑pack) 10995639(vial) |  | 1.60–360.00 nmol/L | 0.15–34.20 μg/mL | No | NA | NA | NA | NA | NA | NA | 1.11 nmol/L | 0.59 nmol/L | 1.11 nmol/L | 0.06 μg/mL | 0.11 μg/mL | 0.11 μg/mL | 18.95 nmol/L | 1.80 μg/mL | 3.1 | 3.6 | 67.85 nmol/L | 6.45 μg/mL | 3.3 | 7 | 138.93 nmol/L | 13.20 μg/mL | 3.4 | 6.6 | ADVIA Centaur SHBG y = 0.99x - 1.93 nmol/L r=0.99 | ADVIA Centaur SHBG y = 0.99x - 0.18 μg/mL r=0.99 | < 10% change up to 500 mg/dL hemoglobin | < 10% change up to 20 mg/dL conjugated bilirubin 20 mg/dL unconjugated bilirubin | < 10% change up to 1000 mg/dL triglycerides | ≤ 10% change in results up to 250 ng/mL biotin | 1000 nmol/L | 95 μg/mL |  |  | Standardized to World Health Organization (WHO) 2nd International Standard for SHBG, NIBSC code 08/266 | 10995421_EN Rev. 04, 2019-08 (OUS) 11200388_EN Rev. 04, 2019-06 (US) | IFU Document | 0.174 | 6.969 | 2.087 | 1.457 | 0.079 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995719 |  | bo01 | 00630414600321 | OUS | Bone Metabolism | No | Atellica IM Vitamin D Total (VitD) | Atellica IM VitD | VitD | 100 |  | NA | Atellica IM VitD QC (10995724) Atellica IM VitD DIL (10995721) Atellica IM VitD MCM (10995723) | 14 | 28 | Mouse monoclonal anti-25(OH)vitamin D antibody | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#EDTA Plasma;# | < 24 hours at RT >24 hours at 2–8°C for up to 7 days &gt;7 days at ≤ -20°C | 20 μL | No |  | Competitive immunoassay using direct chemiluminescent technology | 4.20–150.00 ng/mL | 10.50–375.00 nmol/L |  |  |  |  | See IFU |  |  | Atellica IM VitD CAL | Provided in Kit | 28 days recon at 2–8°C 120 days recon at ≤ -20°C 10 hours recon at RT | Lyophilized | Reconstituted to 2.0 mL/vial | 3 for lot calibration, 2 for pack recalibration | 2 | Defibinated human plasma containing albumin | 28 | 28 |  |  |  | Yes | Atellica IM VitD DIL | 10995721 |  | 4.20–300.00 ng/mL | 10.50–750.00 nmol/L | Yes | Atellica IM VitD QC | 10995724 (3x2.0mL) | NA | NA | NA | NA | 2.73 ng/mL | 1.26 ng/mL | 2.73 ng/mL | 3.15 nmol/L | 6.83 nmol/L | 6.83 nmol/L | 24.76 ng/mL | 61.90 nmol/L | 3.4 | 6.9 | 64.33 ng/mL | 160.83 nmol/L | 3.7 | 6.2 | 122.18 ng/mL | 305.45 nmol/L | 2.2 | 2.8 | ADVIA Centaur VitD y = 1.05x + 2.16 ng/mL r=0.99 | ADVIA Centaur VitD y = 1.05x + 5.40 nmol/L r=0.99 | ≤ 10% change up to 155 mg/dL of hemoglobin | ≤ 10% change up to 40 mg/dL of conjugated bilirubin 40 mg/dL of unconjugated bilirubin | ≤ 10% change up to 540 mg/dL of triglycerides | No Biotin in assay formulation | NA | NA |  |  | Traceable to the ID LC/MS/MS 25(OH)vitamin D RMP | 10995438_EN Rev. 02, 2019-08 | IFU Document | 0.378 | 6.969 | 2.48 | 2.913 | 0.172 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995720 |  | bo01 | 00630414600338 | OUS | Bone Metabolism | No | Atellica IM Vitamin D Total (VitD) | Atellica IM VitD | VitD | 500 |  | NA | Atellica IM VitD QC (10995724) Atellica IM VitD DIL (10995721) Atellica IM VitD MCM ( SMN 10995723) | 14 | 28 | Mouse monoclonal anti-25(OH)vitamin D antibody | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#EDTA Plasma;# | < 24 hours at RT, >24 hours at 2–8°C for up to 7 days, >7 days at ≤ -20°C | 20 μL | No |  | Competitive immunoassay using direct chemiluminescent technology | 4.20–150.00 ng/mL | 10.50–375.00 nmol/L |  |  |  |  | See IFU |  |  | Atellica IM VitD CAL | Provided in Kit | 28 days recon at 2–8°C 120 days recon at ≤ -20°C 10 hours recon at RT | Lyophilized | Reconstitute to 2.0 mL/vial | 3 for lot calibration, 2 for pack recalibration | 2 | Defribrinated human plasma containing albumin | 28 | 28 |  |  |  | Yes | Atellica IM VitD DIL | 10995721 |  | 4.20–300.00 ng/mL | 10.50–750.00 nmol/L | Yes | Atellica IM VitD QC | 10995724 (3x2.0mL) | NA | NA | NA | NA | 2.73 ng/mL | 1.26 ng/mL | 2.73 ng/mL | 3.15 nmol/L | 6.83 nmol/L | 6.83 nmol/L | 24.76 ng/mL | 61.90 nmol/L | 3.4 | 6.9 | 64.33 ng/mL | 160.83 nmol/L | 3.7 | 6.2 | 122.18 ng/mL | 305.45 nmol/L | 2.2 | 2.8 | ADVIA Centaur VitD y = 1.05x + 2.16 ng/mL r=0.99 | ADVIA Centaur VitD y = 1.05x + 5.40 nmol/L r=0.99 | ≤ 10% change up to 155 mg/dL of hemoglobin | ≤ 10% change up to 40 mg/dL of conjugated bilirubin 40 mg/dL of unconjugated bilirubin | ≤ 10% change up to 540 mg/dL of triglycerides | No Biotin in assay formulation | NA | NA |  |  | Traceable to the ID LC/MS/MS 25(OH)vitamin D RMP | 10995438_EN Rev. 02, 2019-08 | IFU Document | 1.54 | 10.748 | 2.48 | 6.063 | 0.7 | 27.3 | 6.3 | 15.4
Atellica® IM | 10995712 |  | ty05 | 00630414600253 | Global | N/AThyroid | No | Atellica IM Thyroid Uptake (TUp) | Atellica IM TUpTUp | TUp | 50 | 1 x 50 | Atellica IM CAL A | NA | 14 | 4 | Acridinium ester labeled T3 -BBG (bovine gamma globulin) and unlabeled T3 | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | Up to 8 hrs at RT, Up to 48 hrs at 2-8°C; Frozen at <-20 °C | 20 | No |  | Competitive immunoasay using direct chemiluminescent technology | NA | NA | NA | NA | 0.75 - 1.23 TU Ratio | 0.75 - 1.23 TU Ratio | NA | NA | NA | Atellica IM CAL A | 10995500 10995501 | 28 days | Lyophilized | Add 5 mL of reagent grade water, let stand 15-20 minutes at RT , gently invert and swirl until homogeneous | 3 for lot calibration, 2 for pack recalibration | 2 | Human plasma | 33 | 4 |  |  |  | No | NA | NA |  | NA | NA | No | Use commercially available control material | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 1.14 T Uptake Ratio | 1.14 T Uptake Ratio | 2.1 | 3.7 | 1.44 T Uptake Ratio | 1.44 T Uptake Ratio | 1.6 | 3.4 | 1.79 T Uptake Ratio | 1.79 T Uptake Ratio | 1.4 | 3.2 | ADVIA Centaur TUp FTI y = 1.05x - 0.07 Index Units r = 0.98 | ADVIA Centaur TUp FTI y = 1.05x - 0.07 Index Units r = 0.98 | insignificant effect up to 1200 mg/dL of hemoglobin | insignificant effect up to 30 mg/dL of bilirubin | insignificant effect up to 2000 mg/dL of triglycerides | No Biotin in assay formulation | NA | NA | NA | NA | Internal Standard | 10995436_EN Rev. 04, 2020-06 | IFU Document | 0.202 | 6.969 | 2.087 | 1.457 | 0.092 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995711 |  | ty05 | 00630414600246 | Global | N/AThyroid | No | Atellica IM Thyroid Uptake (TUp) | Atellica IM TUpTUp | TUp | 250 | 5 x 50 | Atellica IM CAL A | NA | 14 | 4 | Acridinium ester labeled T3 -BBG (bovine gamma globulin) and unlabeled T3 | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | Up to 8 hrs at RT, Up to 48 hrs at 2-8°C; Frozen at <-20 °C | 20 | No |  | Competitive immunoassay using direct chemiluminescent technology | NA | NA | NA | NA | 0.75 - 1.23 TU Ratio | 0.75 - 1.23 TU Ratio | NA | NA | NA | Atellica IM CAL A | 10995500 10995501 | 28 days | Lyophilized | Add 5 mL of reagent grade water, let stand 15-20 minutes at RT , gently invert and swirl until homogeneous | 3 for lot calibration, 2 for pack recalibration | 2 | Human plasma | 33 | 4 |  |  |  | No | NA | NA |  | NA | NA | No | Use commercially available control material | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 1.14 T Uptake Ratio | 1.14 T Uptake Ratio | 2.1 | 3.7 | 1.44 T Uptake Ratio | 1.44 T Uptake Ratio | 1.6 | 3.4 | 1.79 T Uptake Ratio | 1.79 T Uptake Ratio | 1.4 | 3.2 | ADVIA Centaur TUp FTI y = 1.02x + 0.00 (ACS:180 Index Units) r = 0.99 | ADVIA Centaur TUp FTI y = 1.02x + 0.00 (ACS:180 Index Units) r = 0.99 | insignificant effect up to 1200 mg/dL of hemoglobin | insignificant effect up to 30 mg/dL of bilirubin | insignificant effect up to 2000 mg/dL of triglycerides | No Biotin in assay formulation | NA | NA | NA | NA | Internal Standard | 10995436_EN Rev. 04, 2020-06 | IFU Document | 0.812 | 6.969 | 2.087 | 4.37 | 0.369 | 17.7 | 5.3 | 11.1
Atellica® IM | 10995615 |  | in01 | 00630414599380 | Global | Inflamation | No | Atellica IM Total IgE (tIgE) | Atellica IM tlgE | tlgE | 50 |  | Atellica IM CAL 80 | Atellica IM tIgE DIL Atellica IM tIgE MCM (SMN 10995617) | 14 | 28 | Lite Reagent: Goat anti-human IgE antibody, Solid Phase: Mouse anti-human IgE antibody | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | Up to 8 hrs at RT, Up to 48 hrs at 2-8°C, Frozen at <-20 °C | 30 μL | No |  | 2-site sandwich immunoassay | 1.5–3000.0 IU/mL | NA |  |  |  |  | Observed range for normal samples was 0-378 IU/mL | See IFU for age related ranges |  | Atellica IM CAL 80 | 10995497 (2‑pack) | 60 days recon at 2–8°C, 8 hours recon at RT | Lyophilized | Reconstitue with 2.0mLof special reagent water. Let stand for 15-20 minutes at RT | 3 for lot calibration, 2 for pack recalibration | 2 | Human IgE: equine serum, preservatives | 50 | 28 |  |  |  | Yes | Atellica IM tIgE DIL | 10995616 |  | 1.5–15000.0 IU/mL | NA | No | NA | NA | NA | NA | NA | NA | 1.4 IU/mL | 1.0 IU/mL | NA | NA | NA | NA | 9.4 U/mL | NA | 2.6 | 3.5 | 329.2 IU/mL (Control) | NA | 1.4 | 2.1 | 2408.8 IU/mL | NA | 3.5 | 12.6 | ADVIA Centaur tIgE y = 1.04x + 0.48 IU/mL r=0.99 | NA | insignificant effect up to 500 mg/dL of hemoglobin | insignificant effect up to 20 mg/dL of bilirubin | insignificant effect up to 3000 mg/dL of triglycerides | No Biotin in assay formulation | 10,000 IU/mL | NA |  |  | traceable to the World Health Organization (WHO) International Reference Preparation (75/502) | 10995366_EN Rev. 02, 2020-02 | IFU Document | 0.207 | 6.969 | 2.087 | 1.457 | 0.094 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995614 |  | in01 | 00630414599373 | Global | Inflamation | No | Atellica IM Total IgE (tIgE) | Atellica IM tlgE | tlgE | 250 |  | Atellica IM CAL 80 | Atellica IM tIgE DIL Atellica IM tIgE MCM (SMN 10995617) | 14 | 28 | Lite Rgt: Goat anti-human IgE antibody, Solid Phase: Mouse anti-human IgE antibody | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | Up to 8 hrs at RT, Up to 48 hrs at 2-8°C, Frozen at <-20 °C | 30 μL | No |  | 2-site sandwich immunoassay | 1.5–3000.0 IU/mL | NA |  |  |  |  | Observed range for normal samples was 0–378 IU/mL | See IFU for age related ranges |  | Atellica IM CAL 80 | 10995497 (2‑pack) | 60 days recon at 2–8°C, 8 hours recon at RT | Lyophilized | Reconstitue with 2.0 mL of special reagent water. Let stand for 15–20 minutes at RT | 3 for lot calibration, 2 for pack recalibration | 2 | Human IgE;equine serum; preservatives | 50 | 28 |  |  |  | Yes | Atellica IM tIgE DIL | 10995616 |  | 1.5–15000.0 IU/mL | NA | No | NA | NA | NA | NA | NA | NA | 1.4 IU/mL | 1.0 IU/mL | NA | NA | NA | NA | 9.4 U/mL | NA | 2.6 | 3.5 | 329.2 IU/mL (Control) | NA | 1.4 | 2.1 | 2408.8 IU/mL | NA | 3.5 | 12.6 | ADVIA Centaur tIgE y = 1.04x + 0.48 IU/mL r=0.99 | NA | insignificant effect up to 500 mg/dL of hemoglobin | insignificant effect up to 20 mg/dL of bilirubin | insignificant effect up to 3000 mg/dL of triglycerides | No Biotin in assay formulation | 10,000 IU/mL | NA |  |  | traceable to the World Health Organization (WHO) International Reference Preparation (75/502) | 10995366_EN Rev. 02, 2020-02 | IFU Document | 0.827 | 6.969 | 2.087 | 4.37 | 0.376 | 17.7 | 5.3 | 11.1
Atellica® IM | 10995606 |  | me02 | 00630414599298 | Global | N/AMetabolic | No | Atellica IM Homocysteine (HCY) | Atellica IM HCYHCY | HCY | 100 | 1 x 100 | Atellica IM HCY CAL SMN 10995498 | Atellica IM HCY DIL SMN 10995498 Atellica IM HCY MCM SMN 10995609 | 14 | 28 | Mouse monoclonal anti-SAH labeled with acridinium ester | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | 48 hrs at 2-8°C, frozen at <-20 °C up to 13 weeks | 20 μL | No |  | Competitive immunoassay using direct chemiluminescent technology | 0.50–65.00 μmol/L | 0.50–65.00 μmol/L | 0.50– 260.00 μmol/L | 0.50– 260.00 μmol/L | 3.7 – 13.9 μmol/L | 3.7 – 13.9 μmol/L | NA | NA | NA | Atellica IM HCY CAL | 10995498 | 4 hours | Liquid | Ready to use Warm to RT and gently mix | 3 for lot calibration, 2 for pack recalibration | 2 | S-adenosylhomocysteine | 42 | 14 |  |  |  | Yes | Atellica IM HCY DIL | 10995608 |  | 0.50–130.00 μmol/L | 0.50–130.00 μmol/L | No | NA | NA | NA | NA | NA | NA | 0.67 μmol/L | < 0.01 μmol/L | NA | < 0.01 μmol/L | 0.67 μmol/L | NA | 5.53 μmol/L | 5.53 μmol/L | 2.9 | 3.3 | 11.55 μmol/L | 11.55 μmol/L | 1.6 | 2.7 | 26.38 μmol/L | 26.38 μmol/L | 1.4 | 2.5 | ADVIA Centaur HCY y = 1.06 + 0.22 μmol/L r=0.98 | ADVIA Centaur HCY y = 1.06 + 0.22 μmol/L r=0.98 | ≤ 10% change up to 1000 mg/dL of hemoglobin | ≤ 10% change up to 25.0 mg/dL of bilirubin | ≤ 10% change up to 1300 mg/dL of lipid | ≤ 10% change up to 3500 ng/mL of biotin | NA | NA | NA | NA | Internal Standard | 10995362_EN Rev. 03, 2020-11 | IFU Document | 0.341 | 6.969 | 2.48 | 2.913 | 0.155 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995607 |  | me02 | 00630414599304 | Global | N/AMetabolic | No | Atellica IM Homocysteine (HCY) | Atellica IM HCYHCY | HCY | 500 | 5 x 100 | Atellica IM HCY CAL SMN 10995498 | Atellica IM HCY DIL SMN 10995498 Atellica IM HCY MCM SMN 10995609 | 14 | 28 | Mouse monoclonal anti-SAH labeled with acridinium ester | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | 48 hrs at 2-8°C,; frozen at <-20 °C up to 13 weeks | 20 μL | No |  | Competitive immunoassay using direct chemiluminescent technology | 0.50–65.00 μmol/L | 0.50–65.00 μmol/L | 0.50– 260.00 μmol/L | 0.50– 260.00 μmol/L | 3.7 – 13.9 μmol/L | 3.7 – 13.9 μmol/L | NA | NA | NA | Atellica IM HCY CAL | 10995498 | 4 hours | Liquid | Ready to use Warm to RT and gently mix | 3 for lot calibration, 2 for pack recalibration | 2 | S-adenosylhomocysteine | 42 | 14 |  |  |  | Yes | Atellica IM HCY DIL | 10995608 |  | 0.50–130.00 μmol/L | 0.50–130.00 μmol/L | No | NA | NA | NA | NA | NA | NA | 0.67 μmol/L | < 0.01 μmol/L | NA | < 0.01 μmol/L | 0.67 μmol/L | NA | 5.53 μmol/L | 5.53 μmol/L | 2.9 | 3.3 | 11.55 μmol/L | 11.55 μmol/L | 1.6 | 2.7 | 26.38 μmol/L | 26.38 μmol/L | 1.4 | 2.5 | ADVIA Centaur HCY y = 1.06 + 0.22 μmol/L r=0.98 | ADVIA Centaur HCY y = 1.06 + 0.22 μmol/L r=0.98 | ≤ 10% change up to 1000 mg/dL of hemoglobin | ≤ 10% change up to 25.0 mg/dL of bilirubin | ≤ 10% change up to 1300 mg/dL of lipid | ≤ 10% change up to 3500 ng/mL of biotin | NA | NA | NA | NA | Internal Standard | 10995362_EN Rev. 03, 2020-11 | IFU Document | 1.346 | 10.748 | 2.48 | 6.063 | 0.612 | 27.3 | 6.3 | 15.4
Atellica® IM | 10995558 |  | th03 | 00630414598826 | Global | Therapeutic Drug Monitoring | No | Atellica IM Digoxin (Dig) | Atellica IM Dig | Dig | 50 |  | Atellica IM CAL B (10995503 or 10995504) | Atellica IM Dig MCM (10995559) | 10 | 28 | Mouse monoclonal anti-digoxin antibody | ;#Serum;# | Up to 8 hrs at RT | 50 μL | No |  | Competitive immunoassay using direct chemiluminescent technology. | 0.10–5.00 ng/mL | 0.13–6.40 nmol/L |  |  |  |  | 0.80–2.00 ng/mL (1.02–2.56 nmol/L) |  |  | Atellica IM CAL B | 10995503 (2‑pack) 10995504 (6‑pack) | 28 days recon at 2–8°C 4 hours at RT | Lyophilized |  | 3 for lot calibration, 2 for pack recalibration | 2 |  | 51 | 28 |  |  |  | No | NA | NA |  | NA | NA | No | NA | NA | < 0.01 ng/mL | < 0.01 nmol/L | NA | NA | 0.27 ng/mL | 0.10 ng/mL | NA | 0.13 nmol/L | 0.35 nmol/L | NA | 0.94 ng/mL | 1.20 nmol/L | 2.5 | 3.7 | 2.10 ng/mL | 2.69 nmol/L | 1.4 | 2.4 | 3.83 ng/mL | 4.9 nmol/L | 1.9 | 2.6 | ADVIA Centaur DIG y = 1.02x - 0.08 ng/mL r=1.00 | ADVIA Centaur DIG y = 1.02x - 0.10 nmol/L r=1.00 | insignificant effect up to 500 mg/dL of hemoglobin | insignificant effect up to 20 mg/dL of bilirubin | insignificant effect up to 1000 mg/dL of triglycerides | No Biotin in assay formulation | NA | NA |  |  | traceable to an internal standard manufactured using U.S.P. (United States Pharmacopeia) material | 10995341_EN Rev. 02, 2020-02 | IFU Document | 0.178 | 6.969 | 2.087 | 1.457 | 0.081 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995557 |  | th03 | 00630414598819 | Global | Therapeutic Drug Monitoring | No | Atellica IM Digoxin (Dig) | Atellica IM Dig | Dig | 250 |  | Atellica IM CAL B | Atellica IM Dig MCM (10995559) | 10 | 28 | Mouse monoclonal anti-digoxin antibody | ;#Serum;# | Up to 8 hrs at RT | 50 μL | No |  | Competitive immunoassay using direct chemiluminescent technology | 0.10–5.00 ng/mL | 0.13–6.40 nmol/L |  |  |  |  | 0.80–2.00 ng/mL (1.02–2.56 nmol/L) |  |  | Atellica IM CAL B | 10995503 (2‑pack) 10995504 (6‑pack) | 28 days recon at 2–8°C 4 hours at RT | Lyophilized |  | 3 for lot calibration, 2 for pack recalibration | 2 |  | 51 | 28 |  |  |  | No | NA | NA |  | NA | NA | No | NA | NA | < 0.01 ng/mL | < 0.01 nmol/L | NA | NA | 0.27 ng/mL | 0.10 ng/mL | NA | 0.13 nmol/L | 0.35 nmol/L | NA | 0.94 ng/mL | 1.20 nmol/L | 2.5 | 3.7 | 2.10 ng/mL | 2.69 nmol/L | 1.4 | 2.4 | 3.83 ng/mL | 4.9 nmol/L | 1.9 | 2.6 | ADVIA Centaur DIG y = 1.02x - 0.08 ng/mL r=1.00 | ADVIA Centaur DIG y = 1.02x - 0.10 nmol/L r=1.00 | insignificant effect up to 500 mg/dL of hemoglobin | insignificant effect up to 20 mg/dL of bilirubin | insignificant effect up to 1000 mg/dL of triglycerides | No Biotin in assay formulation | NA | NA |  |  | traceable to an internal standard manufactured using U.S.P. (United States Pharmacopeia) material | 10995341_EN Rev. 02, 2020-02 | IFU Document | 0.678 | 6.969 | 2.087 | 4.37 | 0.308 | 17.7 | 5.3 | 11.1
Atellica® IM | 10995554 |  | th04 | 00630414598789 | Global | Therapeutic Drug Monitoring | No | Atellica IM Digitoxin (Dgtn) | Atellica IM Dgtn | Dgtn | 50 |  | Atellica IM CAL Z | Atellica IM Multi‑Diluent 5 | 10 | 28 | Mouse monoclonal anti-digitoxin antibody | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | Up to 8 hrs at RT | 10 μL | No |  | competitive immunoassay using direct chemiluminescent technology | 1.5–80.0 ng/mL | 2.0–104.6 nmol/L |  |  |  |  | 10–30 ng/mL (13.1–39.2 nmol/L) |  |  | Atellica IM CAL Z | 10995521 (2‑pack) | 28 days recon at 2–8°C 8 hours at RT | Lyophilized |  | 3 for lot calibration, 2 for pack recalibration | 2 |  | 49 | 7 |  |  |  | Yes | Atellica IM Multi‑Diluent 5 | 10995647 |  | 1.5–160.0 ng/mL | 2.0–209.2 nmol/L | No | NA | NA | < 0.01 ng/mL | < 0.1 nmol/L | NA | NA | 2.3 ng/mL | 0.5 ng/mL | NA | 0.7 nmol/L | 3.0 nmol/L | NA | 11.6 ng/mL | 15.1 nmol/L | 2.4 | 4.3 | 18.9 ng/mL | 24.7 nmol/L | 1.7 | 3.7 | 44.4 ng/mL | 58.1 nmol/L | 1.9 | 3.5 | ADVIA Centaur DGTN y = 0.99x + 0.26 ng/mL r=1.00 | ADVIA Centaur DGTN y = 0.99x + 0.34 nmol/L r=1.00 | insignificant effect up to 500 mg/dL of hemoglobin | insignificant effect up to 20 mg/dL of bilirubin | insignificant effect up to 1000 mg/dL of triglycerides | No Biotin in assay formulation | NA | NA |  |  | traceable to an internal standard manufactured using USP (United States Pharmacopeia) material | 10995339_EN Rev. 02, 2019-07 | IFU Document | 0.198 | 6.969 | 2.087 | 1.457 | 0.09 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995553 |  | th04 | 00630414598772 | Global | Therapeutic Drug Monitoring | No | Atellica IM Digitoxin (Dgtn) | Atellica IM Dgtn | Dgtn | 250 |  | Atellica IM CAL Z | Atellica IM Multi‑Diluent 5 | 10 | 28 | Mouse monoclonal anti-digitoxin antibody | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | Up to 8 hrs at RT | 10 μL | No |  | competitive immunoassay using direct chemiluminescent | 1.5–80.0 ng/mL | 2.0–104.6 nmol/L |  |  |  |  | 10–30 ng/mL (13.1–39.2 nmol/L) |  |  | Atellica IM CAL Z | 10995521 (2‑pack) | 28 days recon at 2–8°C 8 hours at RT | Lyophilized |  | 3 for lot calibration, 2 for pack recalibration | 2 |  | 49 | 7 |  |  |  | Yes | Atellica IM Multi‑Diluent 5 | 10995647 |  | 1.5–160.0 ng/mL | 2.0–209.2 nmol/L | No | NA | NA | < 0.01 ng/mL | < 0.1 nmol/L | NA | NA | 2.3 ng/mL | 0.5 ng/mL | NA | 0.7 nmol/L | 3.0 nmol/L | NA | 11.6 ng/mL | 15.1 nmol/L | 2.4 | 4.3 | 18.9 ng/mL | 24.7 nmol/L | 1.7 | 3.7 | 44.4 ng/mL | 58.1 nmol/L | 1.9 | 3.5 | ADVIA Centaur DGTN y = 0.99x + 0.26 ng/mL r=1.00 | ADVIA Centaur DGTN y = 0.99x + 0.34 nmol/L r=1.00 | insignificant effect up to 500 mg/dL of hemoglobin | insignificant effect up to 20 mg/dL of bilirubin | insignificant effect up to 1000 mg/dL of triglycerides | No Biotin in assay formulation | NA | NA |  |  | traceable to an internal standard manufactured using USP (United States Pharmacopeia) material | 10995339_EN Rev. 02, 2019-07 | IFU Document | 0.689 | 6.969 | 2.087 | 4.37 | 0.313 | 17.7 | 5.3 | 11.1
Atellica® IM | 10995530 |  | ca03 | 00630414598574 | Global | Cardiac | No | Atellica IM Creatine Kinase MB (CKMB) | Atellica IM CKMB | CKMB | 100 |  | Atellica IM CKMB CAL (SMN 10995532) 2-pack | Atellica IM CKMB DIL (SMN 10995533) Atellica IM CKMB MCM (SMN 10995534) | 10 | 28 | mouse monoclonal | ;#Serum;#Heparinized Plasma;# | Up to 4 hrs at RT, 2-8C up to 48hr, 1 freeze thaw | 100 μL | No |  | direct chemiluminescent technology | 0.18–300.00 ng/mL | 0.002–3.75 nmol/L |  |  |  |  | CKMB results from 233 apparently healthy individuals gave a median result of 0.78 ng/mL (0.0098 nmol/L) | Serum CKMB results from 167 hospitalized patients with non-cardiac related disorders gave a median result of 1.49 ng/mL (0.0186 nmol/L) | Serum CKMB results from 42 patients with confirmed myocardial injury ranged up to 144 ng/mL (1.80 nmol/L), with a median result of 25.4 ng/mL (0.3175 nmol/L) | Atellica IM CKMB CAL | 10995532 (2‑pack) | 14 days recon at 2–8°C, 4 hrs at RT | Lyophilized | Reconstitued to 2.0 mL/vial | 3 for lot calibration, 2 for pack recalibration | 2 | Human source material | 66 | 28 |  |  |  | Yes | Atellica IM CKMB DIL | 10995533 |  | 0.18–3000.00 ng/mL | 0.0023–37.5 nmol/L | No | NA | NA | 0.08 ng/mL | 0.001 nmol/L | NA | NA | 0.31 ng/mL | 0.13 ng/mL | NA | 0.002 nmol/L | 0.004 nmol/L | NA | 2.18 | 0.027 | 3.8 | 7.3 | 12.48 | 0.156 | 1.1 | 2.4 | 63.31 | 0.791 | 1 | 2.2 | ADVIA Centaur CKMB y = 1.03x + 0.25 ng/mL r=1.00 | ADVIA Centaur CKMB y = 1.03x + 0.003 nmol/L r=1.00 | insignificant effect up to 150 mg/dL of hemoglobin | insignificant effect up to 40 mg/dL of bilirubin | insignificant effect up to 1000 mg/dL of triglycerides | No Biotin in assay formulation | 500 ng/mL | 6.25 nmol/L |  |  | traceable to an internal standard manufactured using highly purified material | 10995316_EN Rev. 02, 2019-07 | IFU Document | 0.205 | 6.969 | 2.087 | 1.457 | 0.093 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995531 |  | ca03 | 00630414598581 | Global | Cardiac | No | Atellica IM Creatine Kinase MB (CKMB) | Atellica IM CKMB | CKMB | 500 |  | Atellica IM CKMB CAL (SMN 10995532) 2-pack | Atellica IM CKMB DIL (SMN 10995533) Atellica IM CKMB MCM (SMN 10995534) | 10 | 28 | Mouse monoclonal anti-CK-MB antibody | ;#Serum;#Heparinized Plasma;# | Up to 4 hrs at RT, 2-8C up to 48hr, 1 freeze thaw | 100 μL | No |  | direct chemiluminescent technology | 0.18–300.00 ng/mL | 0.002–3.75 nmol/L |  |  |  |  | CKMB results from 233 apparently healthy individuals gave a median result of 0.78 ng/mL (0.0098 nmol/L) | Serum CKMB results from 167 hospitalized patients with non-cardiac related disorders gave a median result of 1.49 ng/mL (0.0186 nmol/L) | Serum CKMB results from 42 patients with confirmed myocardial injury ranged up to 144 ng/mL (1.80 nmol/L), with a median result of 25.4 ng/mL (0.3175) | Atellica IM CKMB CAL | 10995532 (2‑pack) | 14 days recon at 2–8°C 4 hrs at RT | Lyophilized | Reconstituted to 2.0 mL/vial | 3 for lot calibration, 2 for pack recalibration | 2 | Human source material | 66 | 28 |  |  |  | Yes | Atellica IM CKMB DIL | 10995533 |  | 0.18–3000.00 ng/mL | 0.0023–37.5 nmol/L | No | NA | NA | 0.08 ng/mL | 0.001 nmol/L | NA | NA | 0.31 ng/mL | 0.13 ng/mL | NA | 0.002 nmol/L | 0.004 nmol/L | NA | 2.18 | 0.027 | 3.8 | 7.3 | 12.48 | 0.156 | 1.1 | 2.4 | 63.31 | 0.791 | 1 | 2.2 | ADVIA Centaur CKMB y = 1.03x + 0.25 ng/mL r=1.00 | ADVIA Centaur CKMB y = 1.03x + 0.003 nmol/L r=1.00 | insignificant effect up to 150 mg/dL of hemoglobin | insignificant effect up to 40 mg/dL of bilirubin | insignificant effect up to 1000 mg/dL of triglycerides | No Biotin in assay formulation | 500 ng/mL | 6.25 nmol/L |  |  | traceable to an internal standard manufactured using highly purified material | 10995316_EN Rev. 02, 2019-07 | IFU Document | 0.818 | 6.969 | 2.087 | 4.37 | 0.372 | 17.7 | 5.3 | 11.1
Atellica® IM | 10995466 |  | ty07 | 00630414597935 | Global | N/AThyroid | No | Atellica IM Anti‑Thyroid Peroxidase (aTPO) | Atellica IM aTPOaTPO | aTPO | 100 | 1 x 100 | Atellica IM CAL O (SMN 10995516) | Atellica IM aTPO QC SMN 10995470 Atellica IM aTPO DIL SMN 10995468 Atellica IM aTPO MCM SMN 10995469 | 14 | 28 | acridinium-ester-labeled- mouse monoclonal anti-TPO antibody | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | 8hr at room temperature, separate cells from serum, 48 hours at 2-8C, >48 hrs freeze at -20C once | 30 | No |  | Competitive immunoassay using direct chemiluminescent technology | 28–1300 U/mL | NA | 1:5 28 - 6500 1:10 28 - 13,000 | 1:5 28 - 6500 1:10 28 - 13,000 | Cutoff 60 U/mL | Cutoff 60 U/mL | NA | NA | NA | Atellica IM CAL O | 10995516 | 8 hours | Liquid | Let stand 5 minutes at room temperature, Gently mix | 3 for lot calibration, 2 for pack recalibration | 2 | Human plasma with antiTPO (human) | 28 | 14 |  |  |  | Yes | Atellica IM aTPO DIL | 10995468 |  | 1:5 28 - 6500 1:10 28 - 13,000 | 1:5 28 - 6500 1:10 28 - 13,000 | Yes | Atellica IM aTPO QC | 10995470 | NA | NA | NA | NA | 28 U/mL | 13 U/mL | NA | 13 U/mL | 28 U/mL | NA | 70 U/mL | 70 U/mL | 4.1 | 7.3 | 490 U/mL | 490 U/mL | 1.3 | 3.1 | 898 U/mL | 898 U/mL | 1.5 | 3.4 | ADVIA Centaur aTPO y = 1.03x - 11 U/mL r = 0.99 | ADVIA Centaur aTPO y = 1.03x - 11 U/mL r = 0.99 | ≤ 5% change up to 500 mg/dL of hemoglobin | ≤ 5% change up to 40 mg/dL of bilirubin | ≤ 5% change up to 1000 mg/dL of lipid | No Biotin in assay formulation | NA | NA | NA | NA | WHO Reference Preparation 66/387 | 10995280_EN Rev. 02, 2019-07 | IFU Document | 0.211 | 6.969 | 2.087 | 1.457 | 0.096 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995467 |  | ty07 | 00630414597942 | Global | N/AThyroid | No | Atellica IM Anti‑Thyroid Peroxidase (aTPO) | Atellica IM aTPOaTPO | aTPO | 500 | 5 x 100 | Atellica IM CAL O (SMN 10995516) | Atellica IM aTPO QC SMN 10995470 Atellica IM aTPO DIL SMN 10995468 Atellica IM aTPO MCM SMN 10995469 | 14 | 28 | acridinium-ester-labeled mouse monoclonal anti-TPO antibody | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | 8hr at room temperature, separate cells from serum, 48 hours at 2-8C, >48 hrs freeze at -20C once | 30 | No |  | Competitive immunoassay using direct chemiluminescent technology | 28–1300 U/mL | NA | 1:5 28 - 6500 1:10 28 - 13,000 | 1:5 28 - 6500 1:10 28 - 13,000 | Cutoff 60 U/mL | Cutoff 60 U/mL | NA | NA | NA | Atellica IM CAL O | 10995516 | 8 hours | Liquid | Let stand 5 minutes at room temperature, Gently mix | 3 for lot calibration, 2 for pack recalibration | 2 | Human plasma with antiTPO (human) | 28 | 14 |  |  |  | Yes | Atellica IM aTPO DIL | 10995468 |  | 1:5 28 - 6500 1:10 28 - 13,000 | 1:5 28 - 6500 1:10 28 - 13,000 | Yes | Atellica IM aTPO QC | 10995470 | NA | NA | NA | NA | 28 U/mL | 13 U/mL | NA | 13 U/mL | 28 U/mL | NA | 70 U/mL | 70 U/mL | 4.1 | 7.3 | 490 U/mL | 490 U/mL | 1.3 | 3.1 | 898 U/mL | 898 U/mL | 1.5 | 3.4 | ADVIA Centaur aTPO y = 1.03x - 11 U/mL r = 0.99 | ADVIA Centaur aTPO y = 1.03x - 11 U/mL r = 0.99 | ≤ 5% change up to 500 mg/dL of hemoglobin | ≤ 5% change up to 40 mg/dL of bilirubin | ≤ 5% change up to 1000 mg/dL of lipid | No Biotin in assay formulation | NA | NA | NA | NA | WHO Reference Preparation 66/387 | 10995280_EN Rev. 02, 2019-07 | IFU Document | 0.849 | 6.969 | 2.087 | 4.37 | 0.386 | 17.7 | 5.3 | 11.1
Atellica® IM | 10995701 |  | id16 | 00630414600147 | Global | ToRCH | No | Atellica IM Toxoplasma IgM (Toxo M) | Atellica IM Toxo M | Toxo M | 50 |  | NA | Atellica IM Toxo M QC | 39 | 19 | Immunoglobulin class capture sandwich immunoassay using direct chemiluminescent technology | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | at 2–8°C for up to 7 days at ≤ -20°C for up to 180 days | 10 μL | No |  | Mouse monoclonal anti-human IgM μ antibody | 0.10–40.00 Index | NA |  |  |  |  | Nonreactive, Reactive, Equivocal |  |  | Atellica IM Toxo M CAL | Provided in Kit | At 2–8°C Until expiration date 8 hrs at RT | Liquid |  | 3 for lot calibration, 2 for pack recalibration | 2 |  | 21 | 14 |  |  |  | No | NA | NA |  | NA | NA | Yes | Atellica IM Toxo M QC | 10995702 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 0.95 Index | NA | 2 | 3.2 | 4.0 Index | NA | 2.2 | 3.1 | 14.98 Index | NA | 4.6 | 5.5 | NA | NA | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 40 mg/dL of conjugated bilirubin up to 20 mg/dL of unconjugated bilirubin | ≤ 10% change up to 1000 mg/dL of triglycerides | No Biotin in assay formulation | NA | NA |  |  | Standardization is based upon the relative clinical agreement with commercially available toxoplasma IgM assays | 10995432_EN Rev. 02, 2019-08 (OUS) 10995433_EN Rev. 02, 2019-06 (US) | IFU Document | 0.257 | 6.969 | 2.48 | 2.913 | 0.117 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995699 |  | id15 | 00630414600123 | Global | ToRCH | No | Atellica IM Toxoplasma IgG (Toxo G) | Atellica IM Toxo G | Toxo G | 100 |  | Atellica IM APW1 (SMN 10995458) | Atellica IM Toxo G QC | 28 | 28 | Mouse monoclonal anti-human IgGFc antibody | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | at 2–8°C for up to 7 days after 7 days at ≤ -20°C | 10 μL | No |  | Immunoglobulin class capture sandwich immunoassay using direct chemiluminescent technology | 0.5–700.0 IU/mL | NA |  |  |  |  | Nonreactive, Reactive, Equivocal |  |  | Atellica IM Toxo G CAL | Provided in Kit | At 2–8°C Until expiration date 8 hrs at RT | Liquid |  | 3 for lot calibration, 2 for pack recalibration | 2 |  | 56 | 14 |  |  |  | No | NA | NA |  | NA | NA | Yes | Atellica IM Toxo G QC | 10995700 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 8.1 IU/mL | NA | 1.8 | 2.8 | 441.5 IU/mL | NA | 0.9 | 1.4 | 690.5 IU/mL | NA | 1.2 | 1.7 | NA | NA | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 60 mg/dL of conjugated bilirubin up to 40 mg/dL of unconjugated bilirubin | ≤ 10% change up to 1000 mg/dL of triglycerides | No Biotin in assay formulation | NA | NA |  |  | Traceable to World Health Organization (WHO) 3rd International standard for human anti-toxoplasma serum | 10995430_EN Rev. 02, 2019-08 (OUS) 10995433_EN Rev. 02, 2019-06 (US) | IFU Document | 0.284 | 6.969 | 2.48 | 2.913 | 0.129 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995675 |  | si01 | 00630414599885 | OUS | Special ID | No | Atellica IM Syphilis (Syph) | Atellica IM Syph | Syph | 200 |  | Atellica IM APW1 (SMN 10995458) | Atellica IM Syph QC | 29 | 60 | T. pallidum recombinant antigens | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#EDTA Plasma;#Citrate Plasma;# | RT not recommended 2–8°C for up to 7 days >7 days store at ≤ -20°C | 100 μL | No |  | Fully automated, antigen sandwich assay, using direct chemiluminescent technology | 0.10-45.00 index | NA |  |  |  |  | Nonreactive, Reactive, Equivocal |  |  | Atellica IM Syph CAL | Provided in Kit | At 2–8°C Until expiration date 8 hrs at RT | Liquid |  | 3 for lot calibration, 2 for pack recalibration | 2 |  | 37 | 21 |  |  |  | No | NA | NA |  | NA | NA | Yes | Atellica IM Syph QC | 10995676 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 0.95 Index (Plasma) | NA | 1.4 | 2.5 | 20.68 Index (Plasma) | NA | 1.2 | 3.5 | 32.22 Index (Plasma) | NA | 1.1 | 3.1 | NA | NA | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 40 mg/dL of conjugated bilirubin up to 40 mg/dL of unconjugated bilirubin | 500 ng/mL 1000 mg/dL of triglycerides | Insignificant effect on the assay up to 1500 ng/mL | NA | NA |  |  | Standardization is based upon relative clinical agreement with commercially available syphilis assays | 10995422_EN Rev. 04, 2019-08 | IFU Document | 0.341 | 6.969 | 2.48 | 2.913 | 0.155 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995668 |  | id12 | 00630414599830 | Global | ToRCH | No | Atellica IM Rubella IgM (Rub M) | Atellica IM Rub M | Rub M | 50 |  | Atellica IM APW1 (SMN 10995458) Atellica IM Multi‑Diluent 2 | Atellica IM Rub M QC | 40 | 28 | Mouse monoclonal anti‑human IgMμ antibody | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | at 2–8°C up to 7 days after 7 days at ≤ -20°C | 20 μL | No |  | Sandwich immunoassay using direct chemiluminescent technology | 0.00–15.00 Index | NA |  |  |  |  | Nonreactive, Reactive, Equivocal |  |  | Atellica IM Rub M CAL | Provided in Kit | At 2–8°C Until expiration date 8 hrs at RT | Liquid |  | 3 for lot calibration, 2 for pack recalibration | 2 |  | 91 | 14 |  |  |  | Yes | Atellica IM Multi‑Diluent 2 | 10995644 |  | ? | ? | Yes | Atellica IM Rub M QC | 10995669 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 0.92 Index | NA | 3.7 | 9.9 | 2.39 Index | NA | 2.9 | 4 | NA | NA |  |  | NA | NA | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 60 mg/dL of conjugated bilirubin up to 40 mg/dL of unconjugated bilirubin | ≤ 10% change up to 1000 mg/dL of triglycerides | No Biotin in assay formulation | NA | NA |  |  | Standardization is based upon relative clinical agreement with commercially available Rubella M assays | 10995419_EN Rev. 02, 2019-08 (OUS) 10995420_EN Rev. 02, 2019-06 (US) | IFU Document | 0.249 | 6.969 | 2.48 | 2.913 | 0.113 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995670 |  | id14 | 00630414600529 | Global | ToRCH | No | Atellica IM Rubella IgG (Rub G) | Atellica IM Rub G | Rub G | 100 |  | Atellica IM APW1 (10995458) | Atellica IM Rub G QC | 19 | 23 | Mouse monoclonal anti-human IgGFc antibody | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | at 2–8°C up to 7 days after 7 days at ≤ -20°C | 10 μL | No |  | Sandwich immunoassay using direct chemiluminescent technology | 0.2–500.0 IU/mL | NA |  |  |  |  | Nonreactive, Reactive, Equivocal |  |  | Atellica IM Rub G CAL | Provided in kit | At 2–8°C Until expiration date 8 hrs at RT | Liquid |  | 3 for lot calibration, 2 for pack recalibration | 2 |  | 53 | 14 |  |  |  | No | NA | NA |  | NA | NA | Yes | Atellica IM Rub G QC | 10995671 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 9.1 IU/mL | NA | 5.4 | 8.3 | 82.2 IU/mL (Control) | NA | 2 | 3.6 | 163.8 IU/mL | NA | 2.7 | 4.8 | NA | NA | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 60 mg/dL of conjugated bilirubin up to 40 mg/dL of unconjugated bilirubin | ≤ 10% change up to 1000 mg/dL of triglycerides | No Biotin in assay formulation | NA | NA |  |  | Traceable to the World Health Organization (WHO) 1st International Standard for human anti-rubella immunoglobulin (RUBI‑1‑94) | 10995417_EN Rev. 02, 2019-08 (OUS) 10995418_EN Rev. 02, 2019-06 (US) | IFU Document | 0.271 | 6.969 | 2.48 | 2.913 | 0.123 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995453 |  | id12 | 00630414597836 | OUS | Hepatitis | No | Atellica IM Hepatitis B surface antibody 2 (aHBs2) | Atellica IM aHBs2aHBs2 | aHBs2 | 200 | 1 x 200 | NA | Atellica IM aHBs2 QC Atellica IM Multi‑Diluent 11 | 14 | 90 | Inactivated human hepatitis B surface antigen | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | 2–8°C for up to 7 days ≤ -20°C for longer storage | 100 μL | No |  | direct sandwich | 3.1–1000.0 mIU/mL | 3.1–1000.0 mIU/mL | NA | NA | NA | NA | nonreactive <8.0 mIU/mL | reactive ≥ 12.0 mIU/mL | retest zone ≥8.0 mIU/mL and <12.0 mIU/mL | Atellica IM aHBs2 CAL | Provided in kit | At 2–8°C Until expiration date, 8 hrs at RT | Liquid | N/A | 3 for lot calibration, 2 for pack recalibration | 2 | processed human plasma | 55 | 42 |  |  |  | Yes | Atellica IM Multi‑Diluent 11 | 10995642 (2‑pack) 10995641 (vial) |  | 3.1 - 10,000 mIU/mL | N/A | Yes | Atellica IM aHBs2 QC | 10995454 | N/A | N/A | N/A | N/A | 3.1 mIU/mL | 1.7 mIU/mL | 9.0 mIU/mL | N/A | N/A | N/A | 10.3 mIU/mL | 10.3 mIU/mL | 5.1 | 5.3 | 131.1 mIU/mL | 131.1 mIU/mL | 1.7 | 3 | 797.3 mIU/mL | 797.3 mIU/mL | 2.3 | 3.7 | N/A | N/A | ≤ 15% change up to 500 mg/dL of hemoglobin | ≤ 15% change up to 40 mg/dL of conjugated and unconjugated bilirubin | ≤ 15% change up to 1000 mg/dL triglycerides | ≤15% change up to 1500 ng/mL of biotin | 200,000 mIU/mL | N/A |  |  | Standardized against the World Health Organization (WHO) 1st International Reference Preparation (1977) | 10995277_EN Rev. 04, 2019-07 | IFU Document | 0.299 | 6.969 | 2.48 | 2.913 | 0.136 | 17.7 | 6.3 | 7.4
Atellica® IM | 10995538 |  | me01 | 00630414598659 | Global | N/AMetabolic | No | Atellica IM Cortisol (Cor) | Atellica IM CorCor | Cor | 50 | 1 x 50 | Atellica IM CAL E | Atellica IM Multi‑Diluent 3 Atellica IM Cor MCM (SMN 10995539) | 10 | 10 | Rabbit anti-cortisol antibody bound to mouse monoclonal anti-rabbit IgG antibody | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | Serum: 8hrs at RT; 48hrs at 2-8C; frozen at -20C | 20 μL | No |  | Competitive immunoassay using direct chemiluminescent | 0.50-75.00 μg/dL | 13.80–2069.25 nmol/L | 0.50–150.00 μg/dL | 13.80–4138.50 nmol/L | 7 - 9; (a.m.): 5.27–22.45 μg/dL 3 - 5 (p.m.) 3.44 - 16.76 μg/dL | 7–9 (a.m): 145.4–619.4 nmol/L 3 - 5 (p.m.): 94.9 - 462.4 nmol/L | NA | NA | NA | Atellica IM CAL E | 10995512 (2‑pack) 10995513 (6‑pack) | 14 days recon at 2–8°C 4 hours recons at RT | Lyophilized | Add 2.0 mL of special reagent water into each vial. Let vials stand for 15-20 minutes at room temperature. | 3 for lot calibration, 2 for pack recalibration | 2 | defibrinated human plasma | 19 | 10 |  |  |  | Yes | Atellica IM Multi‑Diluent 3 | 10995645 |  | 0.50–150.00 μg/dL | 13.80–4138.50 nmol/L | No | Use commercially available controls | NA | NA | NA | NA | NA | ≤ 0.14 μg/dL | ≤ 0.06 μg/dL | ≤ 0.31 μg/dL | 1.66 nmol/L | 3.86 nmol/L | 8.55 nmol/L | 3.11 μg/dL | 85.80 nmol/L | 2.1 | 8.6 | 18.82 μg/dL | 519.24 nmol/L | 1.8 | 7.8 | 31.57 μg/dL | 871.02 nmol/L | 1.5 | 7.6 | ADVIA Centaur Cortisol y = 1.02x + 0.14 μg/dL r=1.00 | ADVIA Centaur Cortisol y = 1.02x + 3.86 nmol/L r=1.00 | insignificant effect up to 500 mg/dL of hemoglobin | insignificant effect up to 20 mg/dL of conjugated bilirubin; insignificant effect up to 20 mg/dL of unconjugated bilirubin | insignificant effect up to 1170 mg/dL of triglycerides | No Biotin in assay formulation | NA | NA | NA | NA | Internal standards manufactured analytically which are traceable to gas chromatography-mass spectrometry (GCMS) | 10995336_EN Rev. 03, 2020-03 | IFU Document | 0.196 | 6.969 | 2.087 | 1.457 | 0.089 | 17.7 | 5.3 | 3.7
Atellica® IM | 10995537 |  | me01 | 00630414600413 | Global | N/AMetabolic | No | Atellica IM Cortisol (Cor) | Atellica IM CorCor | Cor | 250 | 5 x 50 | Atellica IM CAL E | Atellica IM Multi‑Diluent 3 Atellica IM Cor MCM (SMN 10995539) | 10 | 10 | Rabbit anti-cortisol antibody bound to mouse monoclonal anti-rabbit IgG antibody | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | RT for up to 8 hrs >8hrs at 2–8°C >48hrs at ≤ -20°C | 20 μL | No |  | Competitive immunoassay using direct chemiluminescent | 0.50-75.00 μg/dL | 13.80–2069.25 nmol/L | 0.50–150.00 μg/dL | 13.80–4138.50 nmol/L | 7 - 9; (a.m.): 5.27–22.45 μg/dL 3 - 5 (p.m.) 3.44 - 16.76 μg/dL | 7–9 (a.m): 145.4–619.4 nmol/L 3 - 5 (p.m.): 94.9 - 462.4 nmol/L | NA | NA | NA | Atellica IM CAL E | 10995512 (2‑pack) 10995513 (6‑pack) | 14 days recon at 2–8°C, 4 hours recons at RT | Lyophilized | Add 2.0 mL of special reagent water into each vial. Let vials stand for 15-20 minutes at room temperature. | 3 for lot calibration, 2 for pack recalibration | 2 | defibrinated human plasma | 19 | 10 |  |  |  | Yes | Atellica IM Multi‑Diluent 3 | 10995645 |  | 0.50–150.00 μg/dL | 13.80–4138.50 nmol/L | No | Use commercially available controls | NA | NA | NA | NA | NA | ≤ 0.14 μg/dL | ≤ 0.06 μg/dL | ≤ 0.31 μg/dL | ≤1.66 nmol/L | ≤3.86 nmol/L | ≤8.55 nmol/L | 3.11 μg/dL | 85.80 nmol/L | 2.1 | 8.6 | 18.82 μg/dL | 519.24 nmol/L | 1.8 | 7.8 | 31.57 μg/dL | 871.02 nmol/L | 1.5 | 7.6 | ADVIA Centaur Cortisol y = 1.02x + 0.14 μg/dL r=1.00 | ADVIA Centaur Cortisol y = 1.02x + 3.86 nmol/L r=1.00 | insignificant effect up to 500 mg/dL of hemoglobin | insignificant effect up to 20 mg/dL of bilirubin; insignificant effect up to 20 mg/dL of unconjugated bilirubin | insignificant effect up to 1170 mg/dL of triglycerides | No Biotin in assay formulation | NA | NA | NA | NA | Internal standards manufactured analytically which are traceable to gas chromatography-mass spectrometry (GCMS) | 10995336_EN Rev. 03, 2020-03 | IFU Document | 0.682 | 6.969 | 2.087 | 4.37 | 0.31 | 17.7 | 5.3 | 11.1
Atellica® IM | 11200588 |  | ca05 | 00630414228945 | OUS | Cardiac | No | Atellica IM NT‑proBNP (PBNP) | Atellica IM PBNP | PBNP | 100 | 1 X 100 | NA | Atellica IM Multi‑Diluent 1 (SMN 10995637) Atellica IM PBNP MCM (SMN 11200590) | 14 | 28 | sheep monoclonal anti-human antibody | ;#Serum;#Lithium Heparin Plasma;#Potassium EDTA Plasma;# | Serum: 3 days at RT 8 days at 2–8°C Lithium heparin: 2 days at RT 8 days at 2–8°C Dipotassium EDTA: 4 days at RT 8 days at 2–8°C All samples: 12 months at ≤ -20°C | 20 μL | No |  | direct chemiluminescent technology | 35–35,000 pg/mL | 4.13–4130 pmol/L |  |  |  |  | 125 pg/mL for patients aged <75 years | 450 pg/mL for patients aged >75 years |  | IM PBNP CAL | NA | Lyophilized at 2–8°C until expiration date Recon at 2–8°C 24 hrs Recon at ≤ -20°C 30 days | Lyophilized | Reconstituted to 2.0 mL/vial | 3 for lot calibration, 2 for pack recalibration | 2 |  | 41 | 28 |  |  |  | Yes | Atellica IM Multi‑Diluent 1 | 10995637 (2‑pack) 10995638 (6‑pack) 10995639 (vial) |  | 35-350,000 pg/mL | 4.13-41,300 pmol/L | No | N | NA | NA | NA | NA | NA | 20 pg/mL | 13 pg/mL | 18 pg/mL | 1.53 pmol/L | 2.36 pmol/L | 2.12 pmol/L | 131 pg/mL (serum) | 16 pmol/L | 2.3 | 3.9 | 3202 pg/mL (serum) | 378 pmol/L | 1.2 | 3.2 | 31,794 pg/mL (serum) | 3,752 pmol/L | 1.8 | 3.7 | ADVIA Centaur PBNP y = 1.07x + 10 pg/mL r=1.00 | ADVIA Centaur PBNP y = 1.07x + 1.18 pmol/L r=1.00 | ≤ 10% change up to 1000 mg/dL of hemoglobin | ≤ 10% change up to 60 mg/dL of conjugated bilirubin up to 60 mg/dL of unconjugated bilirubin | ≤ 10% change up to 3000 mg/dL of Lipemia | ≤ 10% change in results up to 75 ng/mL. | 300,000 pg/mL | 35,400 pmol/L |  |  | Traceable to an internal standard manufactured using synthetic human NT‑proBNP | 11200584_EN Rev. 04, 2019-05 | IFU Document | 0.286 | 6.875 | 2.125 | 2.438 | 0.13 | 17.463 | 5.398 | 6.193
Atellica® IM | 11200589 |  | ca05 | 00630414228952 | OUS | Cardiac | No | Atellica IM NT‑proBNP (PBNP) | Atellica IM PBNP | PBNP | 500 | 5 X 100 | NA | Atellica IM Multi‑Diluent 1 (SMN 10995637) Atellica IM PBNP MCM (SMN 11200590) | 14 | 28 | sheep monoclonal anti-human antibody | ;#Serum;#Lithium Heparin Plasma;#Potassium EDTA Plasma;# | Serum: 3 days at RT 8 days at 2–8°C Lithium heparin: 2 days at RT 8 days at 2–8°C Dipotassium EDTA: 4 days at RT 8 days at 2–8°C All samples: 12 months at ≤ -20°C | 20 μL | No |  | direct chemiluminescent technology | 35–35,000 pg/mL | 4.13–4130 pmol/L |  |  |  |  | 125 pg/mL for patients aged <75 years | 450 pg/mL for patients aged >75 years |  | IM PBNP CAL | NA | Lyophilized at 2–8°C until expiration date Recon at 2–8°C 24 hrs Recon at ≤ -20°C 30 days | Lyophilized | Reconstitute 2.0 mL/vial | 3 for lot calibration, 2 for pack recalibration | 2 |  | 41 | 28 |  |  |  | Yes | Atellica IM Multi‑Diluent 1 | 10995637 (2‑pack) 10995638 (6‑pack) 10995639 (vial) |  | 35-350,000 pg/mL | 4.13 - 41,300 pmol/L | No | N | NA | NA | NA | NA | NA | 20 pg/mL | 13 pg/mL | 18 pg/mL | 1.53 pmol/L | 2.36 pmol/L | 2.12 pmol/L | 131 pg/mL (serum) | 16 pmol/L | 2.3 | 3.9 | 3,202 pg/mL (serum) | 378 pmol/L | 1.2 | 3.2 | 31,794 pg/mL (serum) | 3,752 pmol/L | 1.8 | 3.7 | ADVIA Centaur PBNP y = 1.07x + 10 pg/mL r=1.00 | ADVIA Centaur PBNP y = 1.07x + 1.18 pmol/L r=1.00 | ≤ 10% change up to 1000 mg/dL of hemoglobin | ≤ 10% change up to 60 mg/dL of conjugated bilirubin up to 60 mg/dL of unconjugated bilirubin | ≤ 10% change up to 3000 mg/dL of Lipemia |  | 300,000 pg/mL | 35,400 pmol/L |  |  | Traceable to an internal standard manufactured using synthetic human NT‑proBNP | 11200584_EN Rev. 04, 2019-05 | IFU Document | 0.953 | 6.875 | 2.125 | 6.75 | 0.433 | 17.463 | 5.398 | 17.145
Atellica® CH | 11097643 |  | pl19 | 00630414210827 | Global | Specific Proteins | No | Atellica CH Haptoglobin | HaptHapt | 4038 | 300 | 2 x 150 | N/A | N/A | 10 | 30 days per pack | Anti-human haptoglobin antibody (goat) | ;#Serum;# | Frozen at -20C up to 14 days | 3.4 μL | No |  | PEG-enhanced immunoturbidimetric | 1–340 mg/dL | 0.01–3.40 g/L | N/A | N/A | 40–280 mg/dL for adults | 0.40–2.80 g/L for adults | N/A | N/A | N/A | Atellica CH LSP CAL | 11099434 | 28 | Liquid | N/A | 3 | 6 | Proteins derived from human serum | 10 | 30 |  | No |  | No | Atellica CH Diluent | 11099300 |  | N/A | N/A | No | N/A | N/A | 1 mg/dL | 0 mg/dL | N/A | N/A | 1 mg/dL | 0 mg/dL | N/A | 0.00 g/L | 0.01 g/L | N/A | 80 (Control) | 0.80 (Control) | 0.7 | 2.2 | 125 (Control) | 1.25 (Control) | 0.7 | 2.2 | 214 (Control) | 2.14 (Control) | 1.2 | 2.9 | ADVIA Chemistry 1800 y = 1.06x + 0 mg/dL r=0.997 | ADVIA Chemistry 1800 y = 1.06x + 0.00 g/L r=0.997 | N/A | N/A | N/A | No Biotin in assay formulation | 600 mg/dL | 6.00 g/L | N/A | N/A | (IRMM) CRM-470 from IFCC | 11111580_EN Rev. 03, 2020-03 | IFU Document | 0.381 | 5.625 | 3.5 | 2.75 | 0.173 | 14.287 | 8.89 | 6.985
Atellica® CH | 10470440 |  | dr12 | 00842768034114 | Global | Drugs of Abuse/Toxicology | Yes | Emit II Plus 6-Acetylmorphine Assay (Alliance Application) | 6AM | 8053 | 200 | 2 x 100 |  |  | 6 | 28 | Mouse Mono-clonal Antibody to 6-AM | ;#Urine;# | ​7 days at 25°C, samples should be frozen after 7 days | 15 μL |  |  | EMIT | 10 ng/mL; Semi Quant 2.5 - 20.0 ng/mL | n/a | n/a | n/a | qualitative/semi-quantitative assay | n/a |  |  |  | Emit® Calibrator/Control Level 0 Emit® II Plus 6-AM/Ecstasy Calibrator/Control Level 1, 2, 3, 4 | 10445406 10470441 10470442 10470443 10470444 | expiration date printed on the vial label | Liquid |  | 3 | 4 |  | 60 | 6 |  | No |  | No | [12 empty] | n/a | n/a | n/a | n/a | n/a | n/a | 11.9 | n/a | 1.5 | 2.8 | 10.6 | n/a | 1.7 | 3.4 | 7.9 | n/a | 1.9 | 4.3 | Viva-E Analyzer ± 50% of 10 ng/mL cutoff |  | n/a | n/a | n/a | No Biotin in assay formulation |  |  | n/a | n/a | Traceable to Syva Emit Cal/Cont and Emit II Plus 6-AM Ecstasy which are gravimetrically prepared standards- qualified by GC/MS within ± 10% of nominal. | 9R022UL_US_E | IFU Document
Atellica® CH | 10445371 |  | th16 | 00842768001413 | Global | Therapeutic Drug Monitoring (TDM) | Yes | Emit Caffeine Assay (Alliance Application) | Caff | 8050 | 225 | 3 x 75 |  |  | 7 | 30 | Sheep antibodies reactive to caffeine | ;#Serum;# | ​Store at 2–8°C | 7 μL |  |  | EMIT | 1–30 µg/mL | 5.2–155 µmol/L | n/a | n/a | n/a | n/a |  |  |  | Emit Caffeine Calibrators (included in reagent kit) 0, 1, 3, 7, 15, 30 | Included | 84 | Lyophilized | Reconstitute 1.0 mL, 1 Hour | 3 | 6 | Human Serum | 60 | 8 |  | No |  | No | [12 empty] | n/a | n/a | n/a | n/a | n/a | n/a | 2.17 | n/a | 2.5 | 9.4 | 8.62 | n/a | 1.4 | 6 | 15.72 | n/a | 1.7 | 5.4 | Viva-E Analzyer, y = 0.96x + 0.9 μg/mL n = 102 r = 0.982 |  | n/a | n/a | n/a | No Biotin in assay formulation |  |  | n/a | n/a |  | 6P594UL.12DS_D | IFU Document
Atellica® CH | 10445389 |  | dr18 | 00842768014314 | Global | Drugs of Abuse/Toxicology | Yes | EMIT Tox Serum Tricyclic Antidepressants Assay (Alliance Application) | TCA | 8054 | 225 | 3 x 75 |  |  | 7 | 25 | Sheet Antibodies | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | ​Store at 2–8°C | 11 μL |  |  | EMIT | 300 ng/mL | n/a | [7 empty] | Emit tox Serum Tricyclic Antidepressants Cal/Control | 10445390 | 12 weeks | Lyophilized | Add 3 mL of distilled or deionized water to vial |  |  | Human serum | 25 | 8 |  | No |  | No | [18 empty] | 196.2 mAu/min |  | 0.3 | 0.9 | 229.6 mAu/min |  | 0.5 | 1.4 | 307.2 mAu/min |  | 0.4 | 1.2 | Viva-E Analyzer ± 50% of 300 ng/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation | [5 empty] | 7C244UL.2DS_F | IFU Document
Atellica® CH | 10445391 |  | dr17 | 00842768001604 | Global | Drugs of Abuse/Toxicology | Yes | EMIT Tox Serum Barbiturate Assay (Alliance Application) | sBrb | 8051 | 225 | 3 x 75 |  |  | 7 | 30 | Sheep Antibodies | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | ​Store at 2–8°C | 11 μL |  |  | EMIT | 3 µg/mL, Semi-quantitive 0.45-6 µg/mL | [8 empty] | Emit tox Serum Calibrators | 10445387 | 12 weeks | Lyophilized | Add 3 mL of distilled or deionized water to the vial. |  | 3 | Human Serum | 30 | 8 |  | No |  | No | [18 empty] | 0.92 µg/mL | 163.3 mAu/min | 0.4 | 1.2 | 3.19 µg/mL | 205.1 mAu/min | 0.5 | 1.1 | 5.83 µg/mL | 253.8 mAu/min | 0.6 | 1.2 | Viva-E Analyzer ± 50% of 3 µg/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation | [5 empty] | 7D164UL.12DS_E | IFU Document
Atellica® CH | 10445388 |  | dr16 | 00842768001581 | Global | Drugs of Abuse/Toxicology | Yes | EMIT Tox Serum Benzodiazepine Assay (Alliance Application) | sBnz | 8052 | 225 | 3 x 75 |  |  | 7 | 30 | Sheep Antibodies | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | ​​Store at 2–8°C | 11 μL |  |  | EMIT | 0.3 µg/mL, Semi-quantitive 0.07-2 µg/mL | [8 empty] | Emit tox Serum Calibrators | 10445387 | 12 weeks | Lyophilized | Add 3 mL of distilled or deionized water to the vial. |  |  | Human Serum | 30 | 8 |  | No |  | No | [18 empty] | 0.04 µg/mL | 155.8 mAu/min | 0.6 | 0.9 | 0.29 µg/mL | 176.3 mAu/min | 0.5 | 1.1 | 1.89 µg/mL | 322.3 mAu/min | 0.7 | 1.2 | Viva-E Analyzer ± 50% of 0.3 µg/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation | [5 empty] | 7B164UL.13DS_F | IFU Document
Atellica® CH | 10720048 |  | dr13 | 00630414601120 | Global | Drugs of Abuse/Toxicology | Yes | Emit II Plus Buprenorphine Assay (Alliance Application) | Bupn | 8049 | 200 | 2 x 100 |  |  | 6 | 28 | Mouse Monoclonal Antibodies | ;#Urine;# | 5 days at 25°C, samples should be frozen after 5 days | 15 μL |  |  | EMIT | 5 ng/mL; Semi Quant 2.5-25 ng/mL | [8 empty] | Emit® Calibrator/Control Level 0 Emit® II Plus Specialty Drug Calibrator/Control Level 1, 2, 3, 4 | 10445406 10720049 10720050 10720051 10720052 | expiration date printed on the vial | Liquid |  |  | 5 | Human Urine | 28 | 8 |  | No |  | No | [18 empty] | 2.93 ng/mL |  | 2.6 | 3.5 | 4.55 ng/mL |  | 1.4 | 2.8 | 6.83 ng/mL |  | 1.2 | 1.7 | Viva-E Analyzer ± 50% of 5 ng/mL cutoff |  | N/A | N/A | N/A | No Biotin in assay formulation | [5 empty] | 10872256_B | IFU Document
Atellica® CH | 10709280 |  | th17 | 00858724003056 | Global | Therapeutic Drug Monitoring (TDM) | Yes | ARK Methotrexate Assay (Alliance Application) | Mtx | 8048 | 110 | 1 x 110 |  |  | 7 | 62 | Rabbit Polyclonal Antibody to Methotrexate | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#Potassium EDTA Plasma;# | ​Clarified samples at 2-8°C up to 2 weeks, is testing delayed then store at less than -10°C | 20 μL |  |  | Homogenous Enzyme Immunoassay | 0.04 - 1.20 μmol/L | 0.02 - 0.55 μg/mL |  |  | ​See IFU |  |  |  |  | ARK Methotrexate Calibrators | 10709281 |  | Liquid |  | 2 | 6 |  | 62 | 15 |  | No |  | No |  | 10709285 | [10 empty] | 0.02 μmol/L | 0.00 μmol/L | 0.04 μmol/L |  |  |  | 0.07 |  | 6.5 | 8.8 | 0.42 |  | 1.7 | 3.5 | 0.81 |  | 2.6 | 3.8 | Roche/Hitachi 917 y = 0.95x + 0 μmol/L r = 0.97 |  | N/A | N/A | N/A | No Biotin in assay formulation | [5 empty] | 1600-0213-00 Rev. 07 | IFU Document
Atellica® IM | 10995541 |  | di04 | 630414600413 | Global | DiabetesDiabetes | No | Atellica IM C-peptide (CpS) | Atellica IM CpSCpS | CpS | 100 | 1 x 100 | Atellica IM CpS CAL SMN 10995542 | Atellica IM Multi-Diluent 10 SMN 10995640 Atellica IM CpS MCM SMN 10995543 | 14 | 28 | Monoclonal mouse anti-C-peptide antibody | ;#Serum;# | Serum : Up to 8 hrs at RT, Up to 24 hrs at 2-8°C, Frozen at <-20 °C | 50 μL | No |  | 2-site sandwich immunoassay using direct chemiluminescent technology | 0.05 - 30.00 ng/mL | 0.05 - 30.00 ng/mL | 1:2 0.05 - 60.00ng/mL 1:5 0.05 - 150 .0 ng/mL | 1:2 0.05 - 60.00ng/mL 1:5 0.05 - 150 .0 ng/mL | 0.81 - 3.85 ng/mL | 0.81 - 3.85 ng/mL | NA | NA | NA | Atellica IM CpS CAL | 10995542 | 4 hours at RT | Liquid | Let the calibrators stand for 5 minutes at RT, Gently mix | 3 | 2 | citric acid buffer with casein and preservatives | 50 | 14 |  |  |  | Yes | Multi-Diluent 10 | 10995640 |  | 1:2 0.05 - 60.00ng/mL 1:5 0.05 - 150 .0 ng/mL | 1:2 0.05 - 60.00ng/mL 1:5 0.05 - 150 .0 ng/mL | No | Use comercially available control | NA | 0.03 ng/mL | 0.03 ng/mL | NA | NA | 0.11 ng/mL | 0.09 ng/mL | NA | 0.09 ng/mL | 0.11 ng/mL | NA | 1.08 ng/mL | 1.08 ng/mL | 2 | 2.9 | 13.81 ng/mL | 13.81 ng/mL | 1.5 | 3.3 | 24.12 ng/mL | 24.12 ng/mL | 3.2 | 4.9 | ADVIA Centaur CpS; y = 0.98x - 0.03 ng/mL r = 1.00 | ADVIA Centaur CpS; y = 0.98x - 0.03 ng/mL r = 1.00 | ≤ 10% change up to 250 mg/dL of hemoglobin | ≤ 10% change up to 20 mg/dL of bilirubin | ≤ 10% change up to 1000 mg/dL of lipid | No Biotin in assay formulation | 125 ng/mL | 125 ng/mL | 125 ng/mL | 125 ng/mL | WHO International Standard 84/510 | 10997742_EN Rev. 01, 2017-10 | IFU Document | 0.227 | 0 | 0 | 0 | 0.103 | 17.7 | 3.7 | 5.3
Atellica® IM | 10995628 |  | di05 | 630414600444 | Global | DiabetesDiabetes | No | Atellica IM Insulin (IRI) | Atellica IM IRIIRI | IRI | 100 | 1 x 100 | Atellica IM IRI CAL SMN 10995629 | Atellica IM IRI MCM SMN 10995632 | 14 | 21 | mouse monoclonal anti-insulin antibodies | ;#Serum;# | Up to 8 hrs at RT, Up to 24 hrs at 2-8°C, Freeze at -20 °C after 24 hrs | 25 μL | No |  | 2-site sandwich chemiluminescent with acridinium-ester technology | 0.5 - 300.0 mU/L | 0.5 - 300.0 mU/L | 1:2 05 - 600 mU/L 1:5 0.5 - 1500 mU/L | 1:2; 05 - 600 mU/L; 1:5 0.5 - 1500 mU/L | 3.0 - 25 mU/L | 3.0 - 25 mU/L | NA | NA | NA | Atellica IM IRI CAL | 10995629 | 8 hrs at RT | Liquid | Let the calibrators stand for 5 minutes at RT, Gently mix | 3 | 2 | Buffered saline with casein, potassium thiocyanate, sodium azide, and preservatives | 91 | 14 |  |  |  | Yes | Atellica IM IRI DIL | 10995630 (2 x 10 mL) 10995631 (1 x 20 mL) |  | 1:2 0.5 - 600 mU/L 1:5 0.5 - 1500 mU/L | 1:2 0.5 - 600 mU/L 1:5 0.5 - 1500 mU/L | No | Use Commercially Available Controls | NA | 0.3 mU/L | 0.3 mU/L | NA | NA | 0.8 mU/L | 0.6 mU/L | NA | 0.6 mU/L | 0.8 mU/L | NA | 3.7 mU/L | 3.7 mU/L | 1.8 | 3.5 | 52.0 mU/L | 52.0 mU/L | 1.6 | 1.6 | 255.7 mU/L | 255.7 mU/L | 1.4 | 2.8 | ADVIA Centaur Insulin y = 0.99x + 0.06 mU/L r = 0.99 | ADVIA Centaur Insulin y = 0.99x + 0.06 mU/L r = 0.99 | ≤ 6% change up to 125 mg/dL of hemoglobin | ≤ 6% change up to 20 mg/dL of bilirubin | ≤ 6% change up to 1000 mg/dL of lipid | No Biotin in assay formulation | 3000 mU/L | 3000 mU/L | 3000 mU/L | 3000 mU/L | WHO 1st IRP 66/3024 | 10997752_EN Rev. 04, 2020-08 | IFU Document | 0.213 | 6.969 | 1.451 | 2.087 | 0.097 | 17.7 | 3.7 | 5.3
Atellica® CH | 10317369 |  | pl16 | 08058056680730 | Global | Specific Proteins | Yes | Ceruloplasmin (Sentinel Alliance Application) | CERULCERUL | 7054 | 230 | 2 x 115 | Atellica CH Empty Reagent Packs (SMN 11097534) | N/A | 9.7 | 35 | anti-ceruloplasmin polyclonal antiserum (goat) | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | 3 days at 2-8 °C | 4 µL | No |  | Turbidimetric | 5 - 90 mg/dL | 5 - 90 mg/dL | N/A | N/A | 20 - 60 mg/dL | 20 - 60 mg/dL | N/A | N/A | N/A | Plasma proteins Cal 3X | 10330537 | 30 days at 2-8 °C | Liquid | No Preparation required | 3 | 5 | Plasmatic plasma proteins in human based matrix | 35 | 35 |  | No |  | No | N/A | 11099300 |  | N/A | N/A | No | N/A | N/A | N/A | N/A | N/A | N/A | 2 mg/dL | N/A | N/A | N/A | 2 mg/dL | N/A | 8.28 mg/dL | 8.28 mg/dL | 1.5 | 3.1 | 33.80 mg/dL | 33.80 mg/dL | 1 | 1.8 | 73.01 mg/dL | 73.01 mg/dL | 0.5 | 1.7 | ADVIA Chemistry 2400 y = 0.960 x - 0.690 mg/dL r = 0.977 n = 143 | ADVIA Chemistry 2400 y = 0.960 x - 0.690 mg/dL r = 0.977 n = 143 | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 30 mg/dL of bilirubin | ≤ 10% change up to 1000 mg/dL of Trig. | N/A | no prozone effect was observed up to a concentration of 1500 mg/dL | no prozone effect was observed up to a concentration of 1500 mg/dL | no prozone effect was observed up to a concentration of 1500 mg/dL | no prozone effect was observed up to a concentration of 1500 mg/dL | CAP RPSP-3 (Reference Preparation for Serum Proteins Lot-3) / RPPHS (Reference Preparation for Protein in Human Serum) | IFU: 11065N-2.0/ 02 2015/09/18 and App. Sheet: F57 2019/12/17 | IFU Document | 0.463 | 4.37 | 2.42 | 4.69 | 0.21 | 11.1 | 6.15 | 11.9
Atellica® CH | 11097647 |  | pl18 | 00630414236919 | Global | Specific Proteins | No | Atellica CH Cystatin C_2 (CYSC_2) | CYSC_2CYSC_2 | 4051 | 800 | 4 x 200 | N/A | N/A | 10 | 40 | anti-cystatin C antibody (rabbit) | ;#Serum;#Lithium Heparin Plasma;#Potassium EDTA Plasma;# | 7 days at 2-8C; Frozen at -20C or below at least 3 months if they are frozen within 24 hours after collection and if repeated freeze-thaw cycles are avoided. | 5.2 μL | No |  | Latex-enhanced immunoturbidimetric | 0.25–8.93 mg/L | 0.25–8.93 mg/L | 0.25–26.79 mg/L | 0.25–26.79 mg/L | 0.64–1.23 mg/L for adults | 0.64–1.23 mg/L for adults | N/A | N/A | N/A | Atellica® CH Cystatin C_2 Calibrators (CYSC_2 CAL) | 11097648 | 60 | Liquid | N/A | 3 | 6 | Human serum; recombinant human cystatin C | 61 | 40 |  | No |  | Yes | Atellica CH Diluent | 11099300 |  | 0.25–26.79 mg/L | 0.25–26.79 mg/L | No | N/A | N/A | 0.12 mg/L | 0.12 mg/L | N/A | N/A | 0.12 mg/L | 0.08 mg/L | 0.25 mg/L | 0.08 mg/L | 0.12 mg/L | 0.25 mg/L | 0.55 mg/L | 0.55 mg/L | 3.6 | 3.7 | 0.86 mg/L | 0.86 mg/L | 2.1 | 2.2 | 3.86 mg/L | 3.86 mg/L | 0.6 | 0.8 | ADVIA Chemistry 1800 CYSC_2 y = 1.01x + 0.04 mg/L r = 0.999 n=100 | ADVIA Chemistry 1800 CYSC_2 y = 1.01x + 0.04 mg/L r = 0.999 n=100 | ≤ 10% change up to 1000 mg/dL of hemoglobin. | ≤ 10% change up to 60 mg/dL of conjugated bilirubin and 60 mg/dL of unconjugated bilirubin. | ≤ 10% change up to 1000 mg/dL of Intralipid. | No Biotin in assay formulation | 56.00 mg/L | 56.00 mg/L | 56.00 mg/L | 56.00 mg/L | IFCC Standard ERM–DA471 | 11111654_EN Rev. 04, 2020-08 | IFU Document | 0.679 | 5.625 | 3.5 | 2.75 | 0.308 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097645 |  | pl17 | 00630414220703 | OUS | Specific Proteins | No | Atellica CH Wide Range C-Reactive Protein (wrCRP) | wrCRPwrCRP | 4042 | 2000 | 4 x 500 | N/A | N/A | 8 | 51 | CRP antibody synthetic latex (rabbit) | ;#Serum;#Lithium Heparin Plasma;# | 13 days at 2-8C; Frozen at -20C up to 6 months | 10 μL | No |  | Latex-enhanced immunoturbidimetric assay | 0.050–15.600 mg/dL | 0.50–156.00 mg/L | 0.050– 62.400 mg/dL | 0.50–624.00 mg/L | 0.000–0.500 mg/dL for adults | 0.00–5.00 mg/L for adults | N/A | N/A | N/A | Atellica CH Wide Range C-Reactive Protein Calibrators (wrCRP CAL) | 11097646 | 28 | Liquid | Ready to Use | 3 | 6 | Recombinant human C-reactive protein (CRP) in a stabilized protein | 61 | 51 |  | No |  | Yes | NA | 11099300 |  | 0.050– 62.400 mg/dL | 0.50–624.00 mg/L | No | NA | NA | 0.030 mg/dL | 0.30 mg/L | N/A | N/A | 0.030 mg/dL | 0.000 mg/dL | 0.050 mg/dL | 0.00 mg/L | 0.30 mg/L | 0.50 mg/L | 0.562 mg/dL | 5.62 mg/L | 6 | 8 | 5.318 mg/dL | 53.18 mg/L | 4 | 6 | 12.762 mg/dL | 127.62 mg/L | 3 | 5 | BN ProSpec CardioPhase hsCRP y = 1.08x - 0.030 mg/dL r=0.997 n=127 | BN ProSpec CardioPhase hsCRP y = 1.08x - 0.30 mg/L r=0.997 n=127 | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 30 mg/dL of conjugated bilirubin and 30 mg/dL unconjugated bilirubin. | ≤ 10% change up to 500 mg/dL of Intralipid | No Biotin in assay formulation | 95.000 mg/dL | 950.00 mg/L | 95.000 mg/dL | 950.00 mg/L | IRMM reference material ERM-DA474/IFCC | 11111647_EN Rev. 02, 2019-10 | IFU Document | 0.974 | 5.625 | 3.5 | 2.75 | 0.442 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097644 |  | ge05 | 00630414220697 | Global | General Chemistry | No | Atellica CH Calcium (Arsenazo) Serum/plasma | CA_2CA_2 | 4041 | 8200 | 4 x 2050 | N/A | N/A | 3 | 63 | Sodium Acetate, Arsenazo III, Non-reactive stabilizers | ;#Serum;#Lithium Heparin Plasma;# | 8hr at RT; 2 days at 4C; frozen at -20C at least 6 months | 4 μL |  |  | Arsenazo - Colorimetric | 1.0–16.0 mg/dL | 0.25–4.00 mmol/L | 1.0–32.0 mg/dL | 0.25–8.00 mmol/L | 8.7–10.4 mg/dL | 2.18–2.60 mmol/L | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 48 hours | Lyophilized | Add 3.0 mL of deionized water into the vial | 3 | 2 | Bovine Serum | 63 | 63 |  | No |  | Yes | N/A | 11099300 |  | 1.0–32.0 mg/dL | 0.25–8.00 mmol/L | No | N/A | N/A | 0.1 mg/dL | 0.03 mmol/L | N/A | N/A | 0.1 mg/dL | 0.0 mg/dL | N/A | 0.00 mmol/L | 0.03 mmol/L | N/A | 6.5 mg/dL | 1.63 mmol/L | 1.1 | 1.5 | 10.2 mg/dL | 2.55 mmol/L | 0.7 | 0.7 | 14.1 mg/dL | 3.53 mmol/L | 0.6 | 0.7 | ADVIA Chemistry 1800 CA_2 y = 1.02x - 0.1 mg/dL r=1.000 n=100 | ADVIA Chemistry 1800 CA_2 y = 1.02x - 0.03 mmol/L r=1.000 n=100 | ≤ 10% change up to 1000 mg/dL of hemoglobin. | ≤ 10% change up to 50 mg/dL of conjugated bilirubin and 50 mg/dL of unconjugated bilirubin. | ≤ 10% change up to 1000 mg/dL of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | Inductively Coupled Plasma Atomic Emission/NIST | 11111651_EN Rev. 03, 2019-03 | IFU Document | 0.604 | 2.875 | 3.5 | 2.75 | 0.274 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097649 |  | ge19 | 00630414238357 | Global | General Chemistry | No | Atellica CH Amylase_2 | AMY_2AMY_2 | 4052 | 1050 | 3 x 350 | N/A | N/A | 9 | 31 | α-Glucosidase (microorganism); Ethylidene-4-NP-G7 | ;#Serum;#Lithium Heparin Plasma;#Urine;# | Serum and Plasma: 8 days at RT; 31 days at 2–8°C; at least 1 year at -20°C Urine: 10 days at RT; 31 days at 2–8°C | 16 µL | No |  | Ethylidene Blocked-pNPG7 | 20 – 1500 U/L | 20 – 1500 U/L | 20 – 4500 U/L | 20 – 4500 U/L | Serum and plasma: 30–118 U/L Urine: ≤ 650 U/L | Serum and plasma: 30–118 U/L Urine: ≤ 650 U/L | N/A | N/A | N/A | Atellica CH SPCL CHEM CAL | 11099438 | 7 | Lyophilized | 5.0 mL of distilled water; Stand for 30 minutes at room temperature; Invert to mix Prior to use mix the contents by gently inverting the vials | 3 | 2 | Human serum | 62 | 31 |  | No |  | Yes | N/A | 11099300 |  | 20 – 4500 U/L | 20 – 4500 U/L | No | N/A | N/A | 7 U/L | 7 U/L | N/A | N/A | 7 U/L | 1 U/L | 20 U/L | 1 U/L | 7 U/L | 20 U/L | 52 U/L | 52 U/L | 1.1 | 1.4 | 187 U/L | 187 U/L | 0.4 | 0.6 | 1128 U/L | 1128 U/L | 0.2 | 0.4 | Roche cobas AMYL2 (serum) y = 1.09x + 0 U/L r = 0.996 n= 118 Roche cobas AMYL2 (urine) y = 1.11x - 1 U/L r = 0.998 n= 114 | Roche cobas AMYL2 (serum) y = 1.09x + 0 U/L r = 0.996 n= 118 Roche cobas AMYL2 (urine) y = 1.11x - 1 U/L r = 0.998 n= 114 | < 10% change up to 500 mg/dL of hemoglobin. | < 10% change up to 30 mg/dL of bilirubin. | < 10% change up to 650 mg/dL of Intralipid® | No Biotin in assay formulation | N/A | N/A | N/A | N/A | Traceable to the IRMM/IFCC-456 reference material and commutable to the IFCC Alpha-Amylase Primary Reference Procedure | 11111707_EN Rev. 02, 2019-08 | IFU Document | 0.6 | 4.187 | 3.5 | 2.75 | 0.272 | 10.636 | 8.89 | 6.985
Atellica® CH | 11097579 |  | ge19 | 00630414238364 | OUS | General Chemistry | No | Atellica CH Pancreatic Amylase_2 (serum and plasma) | PAMY_2PAMY_2 | 4053 | 1720 | 4 x 430 | N/A | N/A | 9 | 31 | monoclonal mouse antibodies | ;#Serum;#Lithium Heparin Plasma;# | 8 days at RT; 31 days at 2-8C; at least 1 year frozen at -20C | 16 | No |  | Enzymatic reaction | 20–1500 U/L | 20–1500 U/L | 20–7500 U/L | 20–7500 U/L | Serum and Plasma: 13–53 U/L | Serum and Plasma: 13-53 U/L | N/A | N/A | N/A | Atellica CH SPCL CHEM CAL | 11097579 | 7 | Lyophilized | 5.0 mL of distilled water; Stand for 30 minutes at room temperature; Invert to mix Prior to use mix the contents by gently inverting the vials | 3 | 2 | Human serum | 62 | 31 |  | No |  | Yes | N/A | 11099300 |  | 20–7500 U/L | 20–7500 U/L | No | N/A | N/A | 9 U/L | 9 U/L | N/A | N/A | 9 U/L | 1 U/L | 20 U/L | 1 U/L | 9 U/L | 20 U/L | 54 U/L | 54 U/L | 0.9 | 1.2 | 186 U/L | 186 U/L | 0.4 | 0.5 | 1129 U/L | 1129 U/L | 0.2 | 0.3 | Roche cobas AMY-P (serum) y = 1.07x + 1.2 U/L r=1.000 n=106 | Roche cobas AMY-P (serum) y = 1.07x + 1.2 U/L r=1.000 n=106 | ≤ 10% change up to 500 mg/dL of hemoglobin. | 10% change up to 30 mg/dL of bilirubin. | ≤ 10% change up to 1000 mg/dL of Intralipid |  | N/A | N/A | N/A | N/A | IRMM/IFCC-456 Pancreatic Alpha- Amylase Reference Material and traceable to JCCLS-SOP through the CRM-001c reference material. | 11111708_EN Rev. 03, 2020-01 | IFU Document | 0.886 | 5.625 | 3.5 | 2.75 | 0.402 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097644 |  | ge05 | 00630414220697 | Global | General Chemistry | No | Atellica CH Calcium (Arsenazo) urine | CA_2CA_2 | 4041 | 8200 | 4 x 2050 | N/A | N/A | 3 | 63 | Sodium Acetate, Arsenazo III, Non-reactive stabilizers | ;#Urine;# | Urine specimens should be collected in a bottle containing 10 mL of 6 M HCl per 24–hour specimen to prevent calcium salt precipitation. | 4 μL | No |  | Arsenazo - Colorimetric | 1.0–32.0 mg/dL | 0.25–8.00 mmol/L | 1.0–160.0 mg/dL | 0.25–40.00 mmol/L | Urine: 100–300 mg/day | Urine: 2.50–7.50 mmol/day | N/A | N/A | N/A | Atellica CH CHEM CAL | 11099411 | 48 hours | Lyophilized | Add 3.0 mL of deionized water into the vial | 3 | 2 | Bovine serum | 63 | 63 |  | No |  | Yes | N/A | 11099300 |  | 1.0–160.0 mg/dL | 0.25–40.00 mmol/L | No | N/A | N/A | 0.2 mg/dL | 0.50 mmol/L | N/A | N/A | 0.2 mg/dL | 0.1 mg/dL | N/A | 0.03 mmol/L | 0.50 mmol/L | N/A | 6.2 mg/dL | 1.55 mmol/L | 2 | 2.3 | 12.1 mg/dL | 3.03 mmol/L | 0.6 | 1 | 24.0 mg/dL | 6.00 mmol/L | 0.4 | 0.5 | ADVIA Chemistry 1800 CA_2 (urine) y = 1.03x - 0.2 mg/dL r=0.997 n=107 | ADVIA Chemistry 1800 CA_2 (urine) y = 1.03x - 0.05 mmol/L r=0.997 n=107 | N/A | N/A | N/A | No Biotin in assay formulation | N/A | N/A | N/A | N/A | Inductively Coupled Plasma Atomic Emission/NIST | 11111651_EN Rev. 03, 2019-03 | IFU Document | 0.604 | 2.875 | 3.5 | 2.75 | 0.274 | 7.302 | 8.89 | 6.985
Atellica® CH | 10482437 |  | pl28 | 00842768031410 | OUS | Specific Proteins | Yes | N LATEX FLC KAPPA (BN Alliance Application) | FLCKFLCK | 7040 | 72 | 3 x 1.7 mL | N FLC SUPPLEMENTARY REAGENT (10482439) | N/A | 9.26 | 14 | Suspension of polystyrene particles coated with monoclonal antibodies (mouse) to human FLC kappa | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | S,P: 5 days at 2-8C; frozen at or below -20C up to 6 months | 2 μL | No |  | Particle-enhanced turbidimetry | 3.91 to 60.00 mg/L | 3.91 to 60.00 mg/L | 1.17 to 600.00 mg/L | 1.17 to 600.00 mg/L | S,P: Kappa: 6.70–22.4 mg/L FLC κ/λ ratio: 0.86 (0.31 to 1.56) - median, (minimum – maximum) | S,P: Kappa: 6.70–22.4 mg/L FLC κ/λ ratio: 0.86 (0.31 to 1.56) - median, (minimum – maximum) | N/A | N/A | N/A | N FLC STANDARD SL | 10482440 | 42 | Liquid | N/A | 3 | 6 | Human free light chain proteins, human serum albumin and protease inhibitors | 42 | 14 |  | No |  | Yes | N/A | 11099300 |  | 1.17 to 600.00 mg/L | 1.17 to 600.00 mg/L | Yes | N FLC CONTROL SL1 and SL2 | 10482441 and 10482442 | N/A | N/A | N/A | N/A | N/A | N/A | 1.037 mg/L | N/A | N/A | 1.037 mg/L | 7.82 mg/L | 7.82 mg/L | 1.02 | 1.72 | 26.61 mg/L | 26.61 mg/L | 0.68 | 1.33 | 50.08 mg/L | 50.08 mg/L | 0.72 | 1.41 | BN ProSpec® System N Latex FLC kappa y=0.942 x –0.075 mg/L r=0.982 n=167 BN ProSpec® System FLC κ/λ ratio y=1.03 x –0.013 Ratio r=0.971 n=163 | BN ProSpec® System N Latex FLC kappa y=0.942 x –0.075 mg/L r=0.982 n=167 BN ProSpec® System FLC κ/λ ratio y=1.03 x –0.013 Ratio r=0.971 n=163 | ≤ 10% change up to 1100 mg/dL of hemoglobin. | ≤ 10% change up to 65 mg/dL of bilirubin. | ≤ 10% change up to 1115 mg/dL of Trig. | N/A | No high-dose hook effect up to an analyte concentration of 30 000 mg/L | No high-dose hook effect up to an analyte concentration of 30 000 mg/L | No high-dose hook effect up to an analyte concentration of 30 000 mg/L | No high-dose hook effect up to an analyte concentration of 30 000 mg/L | Internal Reference Plasma-Pool | OPJAG03C11 Rev. 08 – en | IFU Document | 0.22 | 3.15 | 1.38 | 2.36 | 0.01 | 8 | 3.5 | 6
Atellica® CH | 10482438 |  | pl29 | 00842768031427 | OUS | Specific Proteins | Yes | N LATEX FLC LAMBDA (BN Alliance Application) | FLCLFLCL | 7041 | 83 | 3 x 2.1 mL | N FLC SUPPLEMENTARY REAGENT (10482439) | N/A | 9.26 | 14 | Suspension of polystyrene particles coated with monoclonal antibodies (mouse) to human FLC lambda | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | S,P: 5 days at 2-8C; frozen at or below -20C up to 6 months | 5 μL | No |  | Particleenhanced turbidimetry | 5.47 to 70.00 mg/L | 5.47 to 70.00 mg/L | 1.64 to 700.00 mg/L | 1.64 to 700.00 mg/L | S,P: Lambda: 8.30–27.0 mg/L FLC κ/λ ratio: 0.86 (0.31 to 1.56) - median, (minimum – maximum) | S,P: Lambda: 8.30–27.0 mg/L FLC κ/λ ratio: 0.86 (0.31 to 1.56) - median, (minimum – maximum) | N/A | N/A | N/A | N FLC STANDARD SL | 10482440 | 42 | Liquid | N/A | 3 | 6 | Human free light chain proteins, human serum albumin and protease inhibitors | 42 | 14 |  | No |  | Yes | N/A | 11099300 |  | 1.64 to 700.00 mg/L | 1.64 to 700.00 mg/L | Yes | N FLC CONTROL SL1 and SL2 | 10482441 and 10482442 | N/A | N/A | N/A | N/A | N/A | N/A | 1.661 mg/L | N/A | N/A | 1.661 mg/L | 8.04 mg/L | 8.04 mg/L | 2.27 | 3.63 | 32.89 mg/L | 32.89 mg/L | 1.12 | 3.22 | 60.64 mg/L | 60.64 mg/L | 1.72 | 4.1 | BN ProSpec® System N Latex FLC lambda 0.927 x –0.067 mg/L r=0.977 n=170 BN ProSpec® System FLC κ/λ ratio y=1.03 x –0.013 Ratio r=0.971 n=163 | BN ProSpec® System N Latex FLC lambda 0.927 x –0.067 mg/L r=0.977 n=170 BN ProSpec® System FLC κ/λ ratio y=1.03 x –0.013 Ratio r=0.971 n=163 | ≤ 10% change up to 820 mg/dL of hemoglobin. | ≤ 10% change up to 65 mg/dL of bilirubin. | ≤ 10% change up to 734 mg/dL of Trig. | N/A | No high-dose hook effect up to an analyte concentration of 30 000 mg/L | No high-dose hook effect up to an analyte concentration of 30 000 mg/L | No high-dose hook effect up to an analyte concentration of 30 000 mg/L | No high-dose hook effect up to an analyte concentration of 30 000 mg/L | Internal Reference Plasma-Pool | OPJAG03C11 Rev. 08 – en | IFU Document | 0.37 | 3.15 | 1.38 | 2.36 | 0.168 | 8 | 3.5 | 6
Atellica® CH | 11097542 |  | is04 | 00630414238340 | Global | Therapeutic Drug Monitoring (TDM) | No | Atellica CH Tacrolimus (TACR) | TACR | 4048 | 400 | 2 x 200 |  |  | 6 | 31 | Mouse monoclonal antibody reactive to tacrolimus | ;#Whole blood;# | 8 hours at 18–25°C or for at least 1 week at 2–8°C or stored frozen for at least 6 months at -20°C | 25 μL |  |  | EMIT | 2.0 - 30.0 ng/mL | n/a | 2.0 - 80.0 ng/mL | n/a | No firm therapeutic range exists for tacrolimus in whole blood. | n/a |  |  |  | Emit 2000 Tacrolimus Calibrators (0, 2.5, 5, 10, 20, 30 ng/mL) | 10445399 | 56 | Liquid | Thaw at a room temperature of 18–25°C for approximately 1 hour or at a refrigerated temperature of 2–8°C overnight. Do not vortex | 3 | 6 | human whole blood | 64 | 10 |  | No |  | No |  | 10445399 | [10 empty] | 1.1 ng/mL | 0.6 ng/mL | 2.8 ng/mL | n/a | n/a | n/a | 4.7 ng/mL | n/a | 8.4 | 13.7 | 10.5 ng/mL | n/a | 3.9 | 6.7 | 15.1 ng/mL | n/a | 3 | 6.2 | Emit 2000 Tacrolimus on the Viva-ProE System (Liver and Kidney) y = 1.00x + 0.1 ng/mL r = 0.989 n = 125 |  | N/A | ≤ 10% interference up to 45 mg/dL bilirubin, conjugated (Tacrolimus concentration 9.1 and 15.2 ng/mL); ≤ 10% interference up to 60 mg/dL bilirubin, uncconjugated (Tacrolimus concentration 8.7 and 13.6 ng/mL); | ≤ 10% interference up to 1125 mg/dL Lipemia (Intralipid®) (Tacrolimus concentration 8.9 and 16.6 ng/mL) |  |  |  | N/A |  | Purified tacrolimus | 11111631_EN Rev. 01, 2018-09 | IFU Document |  |  |  |  | 0.135 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097540 |  | is02 | 00630414238081 | OUS | Therapeutic Drug Monitoring (TDM) | No | Atellica CH Mycophenolic Acid (MPA) | MPA | 4043 | 600 | 2 x 300 |  |  | 9 | 33 | Mouse monoclonal antibody to mycophenolic acid | ;#Lithium Heparin Plasma;#Potassium EDTA Plasma;# | 7 days at 2–8°C or stored frozen for at least 11 months at -20°C or colder. Samples should not be exposed to more than 3 freeze-thaw cycles | 5.7 μL |  |  | EMIT | 0.10 - 15.00 μg/mL | n/a | 0.10 - 30.00 μg/mL | n/a | No established therapeutic range exists for mycophenolic acid in plasma. | n/a |  |  |  | Emit 2000 Mycophenolic Acid Calibrators (0, 0.5, 2, 5, 10, 15 μg/mL) | 10445380 | expiration on bottle | Liquid | n/a | 3 | 6 | Bovine Serum Albumin | 63 | 6 |  | No |  | No |  | 10445380 | [10 empty] | 0.06 μg/mL | 0.03 μg/mL | n/a | n/a | n/a | n/a | 1.09 μg/mL | n/a | 2 | 5.2 | 7.33 μg/mL | n/a | 2.3 | 5.7 | 11.44 μg/mL | n/a | 3.3 | 7.2 | Syva Emit 2000 Mycophenolic Acid on Viva-ProE y = 0.95x - 0.04 μg/mL r = 0.991 n = 102 |  | ≤ 10% change up to 500 mg/dL of hemoglobin at analyte concentration of 1.11 and 4.20 μg/mL | ≤ 10% change up to 20 mg/dL of Bilirubin, conjugated at analyte concentration of 1.00 and 3.89 μg/mL; ≤ 10% change up to 20 mg/dL of Bilirubin, unconjugated at analyte concentration of 0.97 and 3.84 μg/mL | ≤ 10% change up to 937.5 mg/dL of Lipemia (Intralipid®) at analyte concentration of 1.07 and 3.85 μg/mL; ≤ 10% change up to 1250 mg/dL of Lipemia (Triglycerides concentrate) at analyte concentration of 0.99 and 3.82 μg/mL |  |  |  | N/A |  | Purified mycophenolic acid powder | 11111627_EN Rev. 01, 2018-09 | IFU Document |  |  |  |  | 0.172 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097541 |  | dr19 | 00630414238098 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Oxycodone (OXY) | OXY100 | 7046 | 1520 | 4 x 380 |  |  | 7 | 40 | Mouse monoclonal antibody to oxycodone | ;#Urine;# | Refrigerated or unrefrigerated for up to 7 days or stored frozen after 7 days at -20°C for up to 3 years. | 12 μL |  |  | EMIT | 100 ng/mL, Semi-quantitative 50 - 400 ng/mL | n/a | n/a | n/a | n/a | n/a |  |  |  | Emit® Calibrator/Control Level 0 Emit® II Plus Specialty Multi Drug Calibrators/Controls Level 1, 2, 3 | 10445406 10720807 10720813 10720814 | Once opened, the Emit® Calibrator/Control level 0 is stable for 5 weeks when recapped and stored at 2–8°C when not in use. For Emit® II Plus Specialty Multi Drug Level 1, 2 and 3 Expiration date on bottle | Liquid | n/a | 3 | 4 | human urine | 62 | 28 |  | No |  | No | [12 empty] | n/a | n/a | n/a |  |  |  | n/a | n/a |  |  | n/a | n/a |  |  | n/a | n/a |  |  | Atellica CH compared to LC-MS/MS n=100 100% agreement to positives (51/51) - Semiquantitative and Qualitative 84% agreement to negatives (41/49) - Semiquantitative and Qualitative |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  |  |  | >95% pure oxycodone solution confirmed by LC/MS/MS or GC/MS. | 11111624_EN Rev. 01, 2018-09 | IFU Document |  |  |  |  | 0.374 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097541 |  | dr19 | 00630414238098 | Global | Drugs of Abuse/Toxicology | No | Atellica CH Oxycodone (OXY) | OXY300 | 7055 | 1520 | 4 x 380 |  |  | 7 | 40 | Mouse monoclonal antibody to oxycodone | ;#Urine;# | Refrigerated or unrefrigerated for up to 7 days or stored frozen after 7 days at -20°C for up to 3 years. | 4 μL |  |  | EMIT | 300 ng/mL, Semi-quantitative 75 - 1000 ng/mL | n/a | n/a | n/a | n//a | n/a |  |  |  | Emit® Calibrator/Control Level 0 Emit® II Plus Specialty Multi Drug Calibrators/Controls Level 1, 2, 3, 4 | 10445406 10720807 10720813 10720814 10720815 | Once opened, the Emit® Calibrator/Control level 0 is stable for 5 weeks when recapped and stored at 2–8°C when not in use. For Emit® II Plus Specialty Multi Drug Level 1, 2, 3 and 4 Expiration date on the bottle | Liquid | N/A | 3 | 5 | human urine | 62 | 28 |  | No |  | No | [12 empty] | n/a | n/a | n/a |  |  |  | n/a | n/a |  |  | n/a | n/a |  |  | n/a | n/a |  |  | Atellica CH compared to LC-MS/MS n=100 100% agreement to positives (37/37) - Semiqualitative and Qualitative 89% agreement to negatives (56/63) - Semiqualitative and Qualitative |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  |  |  | >95% pure oxycodone solution confirmed by LC/MS/MS or GC/MS. | 11111624_EN Rev. 01, 2018-09 | IFU Document |  |  |  |  | 0.374 | 14.287 | 8.89 | 6.985
Atellica® CH | 11097539 |  | th07 | 00630414220680 | Global | Therapeutic Drug Monitoring (TDM) | No | Atellica CH Lidocaine (LIDO) | LIDO | 4044 | 200 | 2 x 100 |  |  | 8 | 31 | Mouse monoclonal antibody | ;#Serum;#Sodium Heparin Plasma;#Potassium EDTA Plasma;# | Serum specimens may be stable for at least 6 hours at room temperature.6 Serum and plasma specimens may be stable for at least 7 days at 2–8°C.8 pH neutral samples are stable for at least 1–2 years if stored at -20°C. | 7.5 μL |  |  | PETINIA | 0.5 - 12.0 μg/mL | 2.1–51.2 μmol/L | 0.5 - 24.0 μg/mL | 2.1–102.5 μmol/L | 1.5–5.0 μg/mL | 6.4–21.3 μmol/L |  |  |  | Atellica CH DRUG CAL II | 11099405 | 30 | Liquid | Allow to equilibrate to room temperature and mix thoroughly before use. | 3 | 5 | Bovine serum | 61 | 18 |  | No |  | Yes | [12 empty] | 0.3 μg/mL | 0.2 μg/mL | n/a | 0.9 μmol/L | 1.3 μmol/L | n/a | 1.5 μg/mL | 6.4 μmol/L | 2.5 | 3.1 | 7.5 μg/mL | 32.0 μmol/L | 1.2 | 1.6 | 10.4 μg/mL | 44.4 μmol/L | 1 | 1.3 | LIDO on Dimension EXL with LOCI Module y = 0.98x + 0.1 μg/mL (y = 0.98x + 0.4 μmol/L) n =110 r = 0.998 |  | ≤ 10% interference up to 375 mg/dL of hemoglobin at analyte concentration 5.1 and 8.3 μg/mL | ≤ 10% interference up to 60 mg/dL of Bilirubin, conjugated at analyte concentration 5.0 and 8.0 μg/mL; ≤ 10% interference up to 60 mg/dL of Bilirubin, unconjugated at analyte concentration 5.0 and 8.1 μg/mL; | ≤ 10% interference up to 1000 mg/dL of Lipemia (Intralipid®) at analyte concentration 5.0 and 8.0 μg/mL |  |  |  | N/A |  | USP Lidocaine standard | 11111635_EN Rev. 01, 2018-09 | IFU Document |  |  |  |  | 0.114 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097537 |  | th08 | 00630414220659 | Global | Therapeutic Drug Monitoring (TDM) | No | Atellica CH N-Acetylprocainamide (NAPA) | NAPA | 4045 | 200 | 2 x 100 |  |  | 7 | 31 | Mouse monoclonal antibody | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#Potassium EDTA Plasma;# | Separated serum specimens may be stored for at least 8 hours at room temperature. Serum and plasma specimens may be stored for at least 2 weeks at 2–8°C8 or stored frozen for at least 6 months at -20°C | 2.5 μL |  |  | PETINIA | 0.5–30.0 μg/mL | 1.8–108.3 μmol/L | 0.5–60.0 μg/mL | 1.8–216.6 μmol/L | The combined procainamide and N‑acetylprocainamide therapeutic range is 10–30 μg/mL. |  |  |  |  | Atellica CH DRUG CAL II | 11099405 | 30 | Liquid | Allow to equilibrate to room temperature and mix thoroughly before use. | 3 | 5 | Bovine Serum | 61 | 12 |  | No |  | Yes | [12 empty] | 0.3 μg/mL | 0.2 μg/mL | n/a | 0.7 μmol/L | 1.1 μmol/L | n/a | 2.2 μg/mL | 7.9 μmol/L | 2.8 | 4.5 | 13.4 μg/mL | 48.4 μmol/L | 0.7 | 1.4 | 23.6 μg/mL | 85.2 μmol/L | 0.5 | 1.1 | NAPA on Dimension EXL with LOCI Module y = 1.04x - 0.1 μg/mL (y = 1.04x - 0.4 μmol/L) n = 109 r = 0.997 |  | ≤ 10% interference from up to 500 mg/dL hemoglobin at analyte concentration of 9.8 and 14.4 μg/mL | ≤ 10% interference from up to 60 mg/dL Bilirubin, conjugated at analyte concentration of 9.4 and 14.3 μg/mL; ≤ 10% interference from up to 60 mg/dL Bilirubin, unconjugated at analyte concentration of 9.8 and 14.5 μg/mL; | ≤ 10% interference from up to 1000 mg/dL Lipemia (Intralipid®) at analyte concentration of 9.7 and 14.4 μg/mL |  |  |  | N/A |  | Traceable to N-acetylprocainamide of high purity. | 11111639_EN Rev. 01, 2018-09 | IFU Document |  |  |  |  | 0.114 | 7.302 | 8.89 | 6.985
Atellica® CH | 11097538 |  | th10 | 00630414220673 | Global | Therapeutic Drug Monitoring (TDM) | No | Atellica CH Procainamide (PROC) | PROC | 4046 | 200 | 2 x 100 |  |  | 7 | 31 | Mouse monoclonal antibody | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#Potassium EDTA Plasma;# | Separated serum specimens may be stored for at least 8 hours at room temperature. Serum and plasma specimens may be stored for at least 2 weeks at 2–8°C or stored frozen for at least 6 months at -20°C. | 2.5 μL |  |  | PETINIA | 0.5–20.0 μg/mL | 2.1–85.0 μmol/L | 0.5–40.0 μg/mL | 2.1–170.0 μmol/L | 4.0–12.0 μg/mL | 17.0–51.0 μmol/L |  |  |  | Atellica CH DRUG CAL II | 11099405 | 30 | Liquid | Allow to equilibrate to room temperature and mix thoroughly before use. | 3 | 5 | Bovine serum | 61 | 18 |  | No |  | Yes | [12 empty] | 0.4 μg/mL | 0.2 μg/mL | n/a | 0.9 μmol/L | 1.7 μmol/L | n/a | 3.5 μg/mL | 14.9 μmol/L | 1.6 | 2.3 | 7.2 μg/mL | 30.6 μmol/L | 1.3 | 1.6 | 17.7 μg/mL | 75.2 μmol/L | 0.7 | 1 | PROC on Dimension EXL with LOCI Module y = 1.01x + 0.1 μg/mL (y = 1.01x + 0.4 μmol/L) n = 101 r = 0.998 |  | ≤ 10% interference from up to 500 mg/dL hemoglobin at analyte concentration 5.8 and 9.7 μg/mL | ≤ 10% interference from up to 60 mg/dL Bilirubin, conjugated at analyte concentration 5.6 and 9.5 μg/mL; ≤ 10% interference from up to 60 mg/dL Bilirubin, unconjugated at analyte concentration 5.8 and 9.5 μg/mL | ≤ 10% interference from up to 1000 mg/dL Lipemia (Intralipid®) at analyte concentration 5.7 and 9.5 μg/mL |  |  |  | N/A |  | Traceable to USP Procainamide standard | 11111643_EN Rev. 01, 2018-09 | IFU Document |  |  |  |  | 0.114 | 7.302 | 8.89 | 6.985
Atellica® CH | 11354454 |  | dr19 | 00858724003469 | OUS | Drugs of Abuse/Toxicology | Yes | ARK Ethyl Glucuronide (EtG) Assay (Alliance Application) | ETG500 | 7044 | 154 | 1 x 154 |  |  | 7.1 | 60 | Sheep monoclonal antibody to ethyl glucuronide | ;#Urine;# | Store at 2-8°C (36–46°F) and assay within 7 days after collection. After 7 days store the urine sample frozen. | 6 μL |  |  | Homogeneous enzyme immunoassay technique | 100 - 2000 ng/mL | n/a | n/a | n/a | n/a | n/a |  |  |  | ARK Ethyl Glucuronide Calibrators | 11354457 | Use prior to the expiration date. | Liquid | Calibrators are ready to use. Mix each level by gentle inversion before dispensing | 3 | 5 | Human Urine | 60 | 8 |  | No |  | No |  | 11354458 | [10 empty] | n/a | n/a | 50.0 ng/mL |  |  |  | 384.1 ng/mL | n/a | 2 | 2.7 | 505.6 ng/mL | n/a | 2.3 | 3 | 630.5 ng/mL | n/a | 1.4 | 2.3 | Atellica CH compared to Beckman AU680 100% agreement to positives (60/60) 100% agreement to negatives (40/40) n = 100 |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | There is no internationally recognized standard for ethyl glucuronide. A certified solution of ethyl glucuronide is traceable to LC/MS. | 1600-0609-00 Rev. 02 | IFU Document |  |  |  |  | 0.109 | 12.7 | 7.9 | 7.6
Atellica® CH | 11354454 |  | dr19 | 00858724003469 | OUS | Drugs of Abuse/Toxicology | Yes | ARK Ethyl Glucuronide (EtG) Assay (Alliance Application) | ETG1K | 7056 | 154 | 1 x 154 |  |  | 7.1 | 60 | Sheep monoclonal antibody to ethyl glucuronide | ;#Urine;# | Store at 2-8°C (36–46°F) and assay within 7 days after collection. After 7 days store the urine sample frozen. | 6 μL |  |  | Homogeneous enzyme immunoassay technique | 100 - 2000 ng/mL | n/a | n/a | n/a | n/a | n/a |  |  |  | ARK Ethyl Glucuronide Calibrators | 11354457 | Use prior to the expiration date. | Liquid | Calibrators are ready to use. Mix each level by gentle inversion before dispensing | 3 | 5 | Human Urine | 60 | 8 |  | No |  | No |  | 11354458 | [10 empty] | n/a | n/a | 50.0 ng/mL | n/a | n/a | n/a | 751.2 ng/mL | n/a | 1.8 | 2.3 | 1000.6 ng/mL | n/a | 1.3 | 2.4 | 1257.4 ng/mL | n/a | 1.8 | 2.5 | Atellica CH compared to Beckman AU680 98% agreement to positives (41/42) 100% agreement to negatives (58/58) n = 100 |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | There is no internationally recognized standard for ethyl glucuronide. A certified solution of ethyl glucuronide is traceable to LC/MS. | 1600-0609-00 Rev. 02 | IFU Document |  |  |  |  | 0.109 | 12.7 | 7.9 | 7.6
Atellica® CH | 10445383 |  | th14 | 00842768001543 | Global | Therapeutic Drug Monitoring (TDM) | No | Emit® Amikacin Assay (Alliance Application) | Amik | 7052 | 450 | 6 x 75 |  |  | 8.5 | 30 | Sheep antibodies reactive to amikacin | ;#Serum;#EDTA Plasma;#Citrate Plasma;# | Store the serum or plasma refrigerated at 2–8°C. | 6 µL |  |  | EMIT | 2.5-50 μg/mL | 4.3–86 μmol/L | 2.5 - 150 μg/mL | 4.3–258 μmol/L | 20–25 μg/mL for moderately severe infections and 25–30 μg/mL for life threatening infections or less sensitive pathogens. Trough amikacin oncentrations of 1–4 μg/mL for less severe infections and 5–8 μg/mL for life-threatening infections usually ensure that the concentration is above the minimum inhibitory concentration of the pathogen and that drug elimination is adequate. | 34–43 μmol/L for moderately severe infections and 43–51 μmol/L for life threatening infections or less sensitive pathogens. Trough amikacin concentrations of 1.7–6.8 μmol/L for less severe infections and 8.6–14 μmol/L for life-threatening infections usually ensure that the concentration is above the minimum inhibitory concentration of the pathogen and that drug elimination is adequate. |  |  |  | Emit Amikacin Calibrators (included in reagent kit) (0, 2.5, 5, 10, 20, 50) | 10445383 | 84 days | Lyophilized | Reconstitute with 1 mL of distilled or deionized water. After reconstitution, allow the calibrators to equilibrate at a room temperature for 2 hours | 3 | 6 | Human Serum | 19 | 19 |  | No |  | No |  | 10445383 | [10 empty] | n/a | n/a | n/a | n/a | n/a | n/a | 5.5 μg/mL | 9.4 μmol/L | 2.2 | 6.7 | 15.2 μg/mL | 26.0 μmol/L | 1.2 | 5 | 21.3 μg/mL | 36.4 μmol/L | 1.3 | 6.5 | Syva Amikacin assay on the Viva-ProE y=0.98x + 0.2 μg/mL n = 50 |  | No clinically significant interference has been found in samples with up to 800 mg/dL hemoglobin | No clinically significant interference has been found in samples with up to 30 mg/dL bilirubin | No clinically significant interference has been found in samples with up to 1000 mg/dL triglycerides |  |  |  | n/a | n/a | n/a | 6X314UL.12DS_US_E | IFU Document |  |  |  |  | 0.23
Atellica® CH | 10445455 |  | dr20 | 00842768020315 | Global | Drugs of Abuse/Toxicology | No | Emit II Lysergic Acid Diethylamide (LSD) Assay (Alliance Application) | LSD | 7045 | 240 | 1 x 240 |  |  | 9.7 | 30 | Mouse monoclonal antibodies reactive to lysergic acid diethylamide | ;#Urine;# | Unrefrigerated for up to 7 days following collection. After 7 days, specimens should be stored frozen. LSD degrades with prolonged exposure to light. Store specimens in opaque containers or wrap them in opaque paper or aluminum foil and store frozen. | 22.8 µL |  |  | EMIT | 0.5 ng/mL, Semi-quantitive 0–2.5 ng/mL | n/a | n/a | n/a | n/a | n/a |  |  |  | Emit® LSD 0 ng/mL Calibrator Emit® LSD 0.5 ng/mL Calibrator Emit® LSD 1.5 ng/mL Calibrator Emit® LSD 2.5 ng/mL Calibrator | 10445454 10445456 10445457 10445458 | Expiration date on label. | Liquid | n/a | 3 | 4 | Nonsterile human urine | 63 | 7 |  | No |  | No |  | n/a | [10 empty] | n/a | n/a | n/a | n/a | n/a | n/a | 119.22 mAU/min |  | 0.3 | 0.7 | 127.02 mAU/min |  | 0.3 | 0.6 | 138.76 mAU/min |  | 0.2 | 0.8 | Atellica CH compared to Viva-ProE System 100% agreement to positives (54/54) 100% agreement to negatives (52/52) n = 106 |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | n/a | 9L004UL.10DS_F | IFU Document |  |  |  |  | 0.2 | 14 | 9 | 8.9
Atellica® CH | 10445474 |  | dr21 | 00842768014437 | Global | Drugs of Abuse/Toxicology | No | Emit® II Plus Methaqualone (Alliance Application) | Mthaq | 7053 | 212 | 1 x 212 |  |  | 6.7 | 30 | Sheep polyclonal antibodies to methaqualone | ;#Urine;# | Stored unrefrigerated for up to 7 days, stored refrigerated for 30 days before analysis. After 7 days unrefrigerated or 30 days refrigerated, samples should be stored frozen. | 14 µL |  |  | EMIT | 300 ng/mL, Semi-quantitative 150–900 ng/mL | n/a | n/a | n/a | n/a | n/a |  |  |  | Emit® Calibrator/Control Level 0, 2, 3, 4, 5 | 10445406 10445408 10445409 10445410 10445411 | Once opened, the Emit® Calibrators/Controls are stable for 5 weeks when recapped and stored at 2–8°C when not in use | Liquid | n/a | 3 | 5 | Human Urine | 62 | 30 |  | No |  | No |  | n/a | [10 empty] | n/a | n/a | n/a | n/a | n/a | n/a | 147.94 mAU/min |  | 0.5 | 0.7 | 204.10 mAU/min |  | 0.3 | 0.5 | 176.81 mAU/min |  | 0.4 | 0.6 | Atellica CH compared to Viva-ProE System 100% agreement to positives (50/50) 100% agreement to negatives (50/50) n = 100 |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | n/a | 9Q022UL.7DS_E | IFU Document |  |  |  |  | 0.165 | 14 | 8.9 | 8.9
Atellica® CH | 10445287 |  | dr23 | 00842768017247 | Global | Drugs of Abuse/Toxicology | Yes | Syva® Creatinine Validity Test (Alliance Application) | SVTCR | 7047 | 462 | 3 x 154 |  |  | 6.3 | 7 | Reagent 1: NaOH with stabilizer; Reagent 2: 25 mM picric acid | ;#Urine;# | Stable for 1–2 days when stored at 2–10°C. | 8 µL |  |  | Colormetric | 2-20 mg/dL | n/a | n/a | n/a | SAMHSA classification of the specimen is based on creatinine and specific gravity results. See IFU Table 1. For countries outside the U.S., refer to country-specific guidelines. | n/a |  |  |  | Syva® Creatinine Validity Calibrator 2 Syva® Creatinine Validity Calibrator 20 | 10445273 10445269 | 21 days | Liquid | n/a | 3 | 2 | aqueous solution of creatinine | 80 | 2 |  | No |  | No |  | n/a | [10 empty] | n/a | n/a | n/a | n/a | n/a | n/a | 1.92 mg/dL |  | 1.3 | 5.8 | 19.31 mg/dL |  | 0.3 | 4.5 | 96.78 mg/dL |  | 0.3 | 4.5 | Syva® Creatinine Validity Test on the Viva-ProE analyzer y = 1.0x + 1.0 mg/dL n = 100 |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | n/a | 3T252UL.7DS_C | IFU Document |  |  |  |  | 0.265 | 14 | 9 | 11.5
Atellica® CH | 10445288 |  | vt1 | 00842768017254 | Global | Drugs of Abuse/Toxicology | Yes | Syva® Nitrite Validity Test (Alliance Application) | SVTNT | 7048 | 308 | 2 x 154 |  |  | 7.3 | 30 | Sulfanilamide in 0.3 M H3PO4 with surfactant | ;#Urine;# | Room temperature for 7 days, after 7 days store sample frozen. | 6.8 µL |  |  | Colormetric | 5-3500 μg/mL | n/a | n/a | n/a | A urine specimen is reported adulterated when the nitrite concentration is ≥500 μg/mL. A urine specimen is reported as invalid when the nitrite concentration is ≥200 μg/mL. | n/a |  |  |  | Syva® Validity Negative Calibrator/ Control Syva® Nitrite Validity Calibrator 500 | 10445268 10445275 | 21 days | Liquid | n/a | 3 | 2 | aqueous solution of NaNO2, NaCl, and preservative | 62 | 30 |  | No |  | No |  | n/a | [10 empty] | n/a | n/a | n/a | n/a | n/a | n/a | 222 µg/mL |  | 0.4 | 1.9 | 580 µg/mL |  | 0.3 | 1 | 493 µg/mL |  | 0.3 | 1 | Syva® Nitrite Validity Test on the Viva-ProE 100% Agreement to Positives 50/50 100% Agreement to Negatives 50/50 n = 100 |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | n/a | 3T162UL.6DS_US_C | IFU Document |  |  |  |  | 0.22 | 15.4 | 10.5 | 10.3
Atellica® CH | 10445300 |  | vt1 | 00842768017353 | Global | Drugs of Abuse/Toxicology | Yes | Syva® Specific Gravity Validity Test (Alliance Application) | SVTSG | 7051 | 648 | 4 x 162 |  |  | 5.1 | 30 | Poly (methyl vinyl ether-alt-maleic acid), isopropyl alcohol, Brij-35, sodium dodecyl sulfate | ;#Urine;# | Store at room temperature for 7 days. After 7 days, samples should be stored frozen. | 2 µL |  |  | Colormetric | 1.0010–1.0250 | n/a | n/a | n/a | Any specimen with a specific gravity result that is less than 1.0030 or greater than 1.0200 or has a creatinine value < 20 mg/dL should be tested on a refractometer in order to determine suitability for drug screening | n/a |  |  |  | Syva® Specific Gravity Validity Calibrator 1.0030 Syva® Specific Gravity Validity Calibrator 1.0200 | 10445290 10445291 | 21 | Liquid | n/a | 3 | 2 | Synthetic urine matrix | 52 | 2 |  | No |  | No |  | n/a | [10 empty] | n/a | n/a | n/a | n/a | n/a | n/a | 1.0027 |  | 0.03 | 0.1 | 1.0030 |  | 0.02 | 0.1 | 1.0200 |  | 0.02 | 0.1 | Atellica CH compared to Viva-ProE System (1.0030) 97% agreement to positives (76/80) 100% agreement to negatives (50/50) n = 130 Atellica CH compared to Viva-ProE System (1.0200) 100% agreement to positives (50/50) 100% agreement to negatives (80/80) n = 130 |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | n/a | 3T062UL.3DS_C | IFU Document |  |  |  |  | 0.226 | 14 | 9 | 11.7
Atellica® CH | 10445299 |  | vt1 | 00842768017346 | Global | Drugs of Abuse/Toxicology | Yes | Syva® pH Validity Test (Alliance Application) | SVTPH | 7050 | 810 | 5 x 162 |  |  | 4.5 | 30 | Aqueous solution of thymol blue | ;#Urine;# | Frozen samples must be thawed and mixed thoroughly before analysis. | 6.4 µL |  |  | Colormetric | 2.0 - 12.0 | n/a | n/a | n/a | The specimen will be characterized as adulterated if the pH is <3 or ≥11. The specimen will be characterized as invalid if the pH is ≥3.0 and <4.5 or ≥9.0 and <11.0. | n/a |  |  |  | Syva® pH Validity Calibrator 2.0, 3.0, 4.5, 9.0, 11.0, 12.0 | 10445279 10445280 10445286 10445285 10445283 10445284 | 90 | Liquid | n/a | 3 | 6 | Synthetic urine | 50 | 30 |  | No |  | No |  | n/a | [10 empty] | n/a | n/a | n/a | n/a | n/a | n/a | 3.20 |  | 0.5 | 4.9 | 4.46 |  | 0.5 | 2.1 | 9.00 |  | 1.8 | 2.5 | Atellica CH compared to Viva-ProE System (3.0 cutoff) 100% agreement to positives (52/52) 100% agreement to negatives (50/50) n=102 Atellica CH compared to Viva-ProE System (4.5 cutoff) 100% agreement to positives (52/52) 100% agreement to negatives (50/50) n=102 Atellica CH compared to Viva-ProE System (9.0 cutoff) 100% agreement to positives (52/52) 100% agreement to negatives (50/50) n=102 Atellica CH compared to Viva-ProE System (11.0 cutoff) 100% agreement to positives (50/50) 100% agreement to negatives (52/52) n=102 |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | n/a | 3T072UL.5DS_US_B | IFU Document |  |  |  |  | 0.218 | 14 | 11.5 | 9
Atellica® CH | 10445289 |  | vt1 | 00842768017261 | Global | Drugs of Abuse/Toxicology | Yes | Syva® Oxidant Validity Test (Alliance Application) | SVTOX | 7049 | 800 | 5 x 160 |  |  | 4.5 | 30 | 0.03% 2,2’-azino-bis(3-ethylbenzthiazoline-6-sulfonic acid) diammonium salt | ;#Urine;# | Store at room temperature for 7 days. After 7 days, samples should be stored frozen. | 6 µL |  |  | Colormetric | 1.0-50 μg/mL | n/a | n/a | n/a | An oxidant concentration assayed at >50 μg/mL is not normally observed in human urine; such samples should be suspected of adulteration. | n/a |  |  |  | Syva® Validity Negative Calibrator/Control Syva® Chromium (VI) Validity Calibrator 50 | 10445268 10445277 | 28 days | Liquid | n/a | 3 | 2 | Aqueous solution of 50 μg/mL Chromium | 62 | 21 |  | No |  | No |  | n/a | [10 empty] | n/a | n/a | n/a | n/a | n/a | n/a | 39.8 µg/mL |  | 0.3 | 1.4 | 72.1 µg/mL |  | 0.6 | 1.2 | 50.9 µg/mL |  | 0.3 | 1.5 | Atellica CH compared to Viva-ProE System 100% agreement to positives (50/50) 100% agreement to negatives (50/50) n = 100 |  | N/A | N/A | N/A | No Biotin in assay formulation |  |  | n/a | n/a | n/a | 3T674UL.5DS_US_C | IFU Document |  |  |  |  | 0.207 | 14.1 | 9 | 11.5
Atellica® IM | 10733004 |  | on10 | 00630414245751 | OUS | Oncology | No | Atellica IM Calcitonin (CALCT) | Atellica IM CALCT | CALCT | 100 |  | Atellica IM PW3 (SMN 10995666) (50.0 mL/pack ) | Atellica IM Multi-Diluent 13 | 38 | 28 | Recombinant anti-calcitonin antibody | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | 4 hours at RT, 24 hours at 2–8°C, up to 3 weeks at ≤ -20°C, up to 11 weeks at ≤ -70°C only 1 freeze/thaw cycle | 100 μL | No |  | 2-step sandwich immunoassay | 1.89–2000.00 pg/mL | 0.55–585.20 pmol/L |  |  |  |  | Males: ≤ 12.69 pg/mL ( ≤ 3.71 pmol/L) | Females: ≤ 5.89 pg/mL ( ≤ 1.72 pmol/L) |  | Atellica IM CALCT CAL | Included in the kit | Recon at RT 4 hours, recon at 2–8°C 3 days, recon ≤ -20°C 6 months thaw only once | Lyophilized | Recon with 2.0 mL of special reagent water. Invert vials 5 times, let stand for 30 minutes at RT | 3 for lot calibration, 2 for pack recalibration | 2 | Synthetic calcitonin antigen; casein perservative | 28 | 14 |  |  |  | Yes | Atellica IM Multi-Diluent 13 | 10995643 |  | 1.89 - 200,000.00 pg/mL | 0.55 - 58,520.00 pmol/L | No | [6 empty] | 1.89 pg/mL | 1.60 pg/mL | 1.89 pg/mL | 0.47 pmol/L | 0.55 pmol/L | 0.55 pmol/L | 5.47 pg/mL | 1.60 pmol/L | 3.8 | 5.7 | 346.65 pg/mL | 101.43 pmol/L | 2.2 | 3.2 | 1435.74 pg/mL | 420.10 pmol/L | 2.1 | 3.6 | ADVIA Centaur CALCT y = 0.95x + 0.74 pg/mL r= 0.99 | ADVIA Centaur CALCT y = 0.95x + 0.22 pmol/L r= 0.99 | ≤ 10% interference up to 500 mg/dL | ≤ 10% interference up to 40 mg/dL of conjugated bilirubin and up to 60 mg/dL of unconjugated bilirubin | ≤ 10% interference up to 3000 mg/dL (Intralipid) | ≤ 10% up to 40 ng/mL of biotin | 2,400,000 pg/mL | 702,240 pmol/L |  |  | WHO 2nd International Reference Preparation for Calcitonin (Human); NIBSC code: 89/620 | RPBL1331R01_EN Rev. 03, 2020-06 | IFU Document | 0.275 | 6.969 | 2.48 | 2.913 | 0.125 | 17.7 | 6.3 | 7.4
Atellica® IM | 10733001 |  | an05 | 00630414245720 | Global | Anemia | No | Atellica® IM Active-B12 | Atellica® IM AB12AB12 | AB12 | 100 | 1 x 100 | Atellica IM APW1 (probe wash) | Atellica IM AB12 MCM (master curve material) | 38 | 28 | Anti-transcobalamin (TC) monoclonal antibody and anti-holoTC monoclonal antibody | ;#Serum;# | 16 hours at RT, 3 days at 2–8°C. For longer storage, specimens may be frozen for up to 3 months at ≤ -20°C | 50 µL | No |  | 2-step sandwich immunoassay using direct | 4.25 - 146.00 pmol/L | 4.25 - 146.00 pmol/L | 4.25 - 292.00 pmol/L | 4.25 - 292.00 pmol/L | 27.24–169.62 pmol/L | 27.24–169.62 pmol/L |  |  |  | Atellica IM AB12 CAL | N/A (Kitted) | 8 hours at RT and 30 days at 2-8°C | Liquid | N/A | 3 for lot calibration, 2 for pack recalibration | 2 | Recombinant holotranscobalamin; BSA; sodium azide | 28 | 28 |  |  |  | Yes | Atellica IM Multi-Diluent 13 | 10995643 |  | 4.25 - 292.00 pmol/L | 4.25 - 292.00 pmol/L | Yes | Atellica IM AB12 QC | 10733002 | N/A | N/A | N/A | N/A | 0.83 pmol/L | 0.46 pmol/L | 4.25 pmol/L | 0.46 pmol/L | 0.83 pmol/L | 4.25 pmol/L | 19.18 pmol/L | 19.18 pmol/L | 2.7 | 4.9 | 57.37 pmol/L | 57.37 pmol/L | 2.5 | 4.8 | 108.29 pmol/L | 108.29 pmol/L | 2.1 | 4.5 | Atellica IM AB12 assay was compared to the ADVIA Centaur AB12 assay y = 1.05x - 1.21pmol/L r = 0.98 | Atellica IM AB12 assay was compared to the ADVIA Centaur AB12 assay y = 1.05x - 1.21pmol/L r = 0.98 | ≤10% change up to 500 mg/dL of hemoglobin | ≤10% change up to 40 mg/dL of bilirubin | ≤10% change up to 1000 mg/dL of triglycerides | ≤10% change up to 1,000,000 ng/mL of biotin | 91,200 pmol/L | 91,200 pmol/L | 91,200 pmol/L | 91,200 pmol/L | WHO NIBSC 03/178 | RPBL1359R2_EN Rev. 02, 2019-10 | IFU Document | 0.18 | 6.97 | 2.91 | 2.48 | 0.08 | 17.7 | 7.4 | 6.3
Atellica® IM | 10733006 |  | an04 | 00630414245775 | OUS | Anemia | No | Atellica® IM Erythropoietin | Atellica® IM EPOEPO | EPO | 100 | 1 x 100 | N/A | N/A | 19 | 28 | Monoclonal mouse anti-EPO antibody and Anti-EPO mouse monoclonal antibody | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#EDTA Plasma;# | 8 hours at RT, 7 days at 2–8°C. For longer storage, specimens may be frozen for up to 3 months at ≤ -20°C. | 100 | No |  | One-step sandwich immunoassay using direct | 0.98 - 750 mIU/mL | 0.98 - 750 mIU/mL | 0.98 - 7500 mIU/mL | 0.98 - 7500 mIU/mL | 5.45–28.35 mIU/mL | 5.45–28.35 mIU/mL |  |  |  | Atellica IM EPO CAL | N/A (Kitted) | 24 hours at RT or 90 days at 2-8°C | Liquid | N/A | 3 for lot calibration, 2 for pack recalibration | 2 | Recombinant human EPO; calf serum | 28 | 14 |  |  |  | Yes | Atellica IM Multi-Diluent 13 | 10995643 |  | 0.98 - 7500 mIU/mL | 0.98 - 7500 mIU/mL | Yes | Atellica IM EPO QC | 10733008 | N/A | N/A | N/A | N/A | 0.98 mIU/mL | 0.69 mIU/mL | 0.98 mIU/mL | N/A | N/A | N/A | 2.30 mIU/mL | 2.30 mIU/mL | 4.6 | 5.7 | 29.74 mIU/mL | 29.74 mIU/mL | 1.4 | 3.1 | 612.92 | 612.92 | 1.4 | 2.7 | Atellica IM EPO was compared to ADVIA Centaur EPO: y = 0.94x + 0.58 mIU/mL r = 1.00 | Atellica IM EPO was compared to ADVIA Centaur EPO: y = 0.94x + 0.58 mIU/mL r = 1.00 | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 40 mg/dL of bilirubin | ≤ 10% change up to 3000 mg/dL (intralipid) | ≤ 10% change up to 1,000,000 ng/mL of biotin | 192,334 mIU/mL | 192,334 mIU/mL | 192,334 mIU/mL | 192,334 mIU/mL | WHO 2nd IRP 67/343 and WHO 3rd IS 11/170 | RPBL1337R03_EN Rev. 03, 2019-04 | IFU Document | 0.28 | 6.97 | 2.91 | 2.48 | 0.126 | 17.7 | 7.4 | 6.3
Atellica® IM | 10732998 |  | au01 |  | OUS | Autoimmune | No | Atellica IM anti-CCP IgG (aCCP) | Atellica IM aCCP | aCCP | 100 | 1X100 |  | Atellica IM aCCP MCM SMN 10733000 | 33 | 35 | Monoclonal anti-human IgG (0.1 μg/mL) | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | 22 hrs at RT, 7 days at 2–8°C, > 7days store at -20°C or colder, avoid > 2 freeze/thaw cycles | 15 μL | No |  | 2-step sandwich immunoassay | 0.54–200.00 U/mL | NA |  |  |  |  | Healthly individuals median 0.74 U/mL | Samples ≤ 5.00 U/mL indicate that anti-CCP IgG antibodies were not detected | Samples ≥5.00 U/mL indicate that anti-CCP IgG antibodies were detected | Atellica IM aCCP CAL | Included in the kit | 2–8°C 90 days after opening, At room temperature 8 hours | Liquid | NA | 3 for lot calibration, 2 for pack recalibration | 2 | Processed human plasma positive for anti-CCP antibodies | 35 | 14 |  |  |  | Yes | Atellica IM Multi-Diluent 1 | 10995639 |  | 0.54–2000.00 U/mL | NA | Yes | Atellica IM aCCP QC | 10732999 | NA | NA | NA | NA | 0.54 U/mL | 0.20 U/mL | NA | NA | NA | NA | 4.44U/mL | NA | 4.4 | 4.7 | 33.18 U/mL | NA | 3.6 | 4.2 | 107.14 U/mL | NA | 4.5 | 4.7 | NA | NA | ≤ 10% change up to 1000 mg/dL of hemoglobin | ≤ 10% change up to 40 mg/dL of conjugated bilirubin and unconjugated bilirubin | ≤ 10% change up to 1500 mg/dL of Intralipid | ≤ 10% change up to 1500 ng/dL | 2145 U/mL | NA |  |  | Traceable to an internal standard | RPBL1372R02_EN Rev. 04, 2019-10 | IFU Document
Atellica® IM | 11202207 |  | id17 | 00630414246499 | OUS | ToRCH | No | Atellica IM CMV IgG (CMV IgG) | Atellica IM CMV IgG | CMVIgG | 100 | 1 x 100 | Atellica IM APW1 (10995458) | Atellica IM CMV IgG QC (11202209) | 40 min | 30 | anti‑human IgG mouse monoclonal antibody | ;#Serum;#Lithium Heparin Plasma;#Potassium EDTA Plasma;# | On the clot for up to 14 days at 2–8°C. Separated specimens: 48 hrs RT, 7 days 2–8°C. Frozen 2 months at -20°C to -70°C. Do not store in a frost-free freezer. | 20 μL | No |  | 2‑step sandwich immunoassay using indirect chemiluminescent technology | 0.05–30.00 Index | [5 empty] | Nonreactive / Reactive / No equivocal zone |  |  | Atellica IM CMV IgG CAL | Included | 60 days after opening at 2–8°C | Liquid | N/A | 3 for lot calibration, 2 for pack recalibration | 2 |  | 14 | 14 |  |  |  | No | [5 empty] | Yes | Atellica IM CMV IgG QC (quality control | 11202209 | Initial relative sensitivity 99.5% | Initial relative specificity 98.5% | Resoloved relative sensitivity 100.0% | Resoloved relative specificity 99.8% | N/A | N/A | N/A | N/A | N/A | N/A | 0.01 |  | 1.3 | 2.9 | 1.18 |  | 1.5 | 3.4 | 9.17 |  | 1.3 | 3 |  |  | 500 mg/dL of hemoglobin | 20 mg/dL of unconjugated bilirubin | 3000 mg/dL of intralipids (Triglycerides) | 4500 ng/mL of biotin |  |  |  |  | The Atellica IM CMV IgG assay standardization is based upon clinical agreement with a commercially available CMV IgG assay. | 11200619_EN Rev. 03, 2019-08 | IFU Document
Atellica® CH | 11097536 |  | di02 | 00630414220505 | Global | Diabetes | No | Atellica CH Enzymatic Hemoglobin A1c (A1c_E) | A1c_EA1c_E | 4039 | 600 | 2 x 300 | N/A | N/A | 10 | 63 | Peroxidase (Horseradish); fructosyl peptide oxidase (E. coli, recombinant); ofloxacin | ;#Whole blood;#Hemolysate;# | 48 hr at RT, 7 days at 2–8°C, or stored frozen for up to 21 months (with one freeze-thaw) at -70°C. | 4.5 μL | No |  | Enzymatic | 3.80% - 14.00% HbA1c | 18.03–129.50 mmol/mol HbA1c | N/A | N/A | < 5.7%: Normal 5.7–6.4%: Prediabetes ≥ 6.5%: Diabetic | < 39 mmol/mol: Normal 39–47 mmol/mol: Prediabetes ≥ 48 mmol/mol: Diabetic | N/A | N/A | N/A | Atellica CH A1c_E CAL | 11099338 | 11 days | Lyophilized | Level 1 is liquid and ready to use. Level 2 and 3: Add 1.0 mL of deionized water into each vial, let the vials stand for 30 minutes at RT to allow the lyophilized material to dissolve. Gently swirl and invert the vials. | 3 | 3 | Level 1: Horseradish peroxidase. Levels 2 and 3: human whole blood | 180 | 63 |  | No |  | No | N/A | 11099300 |  | N/A | N/A | No | N/A | N/A | NA | NA | NA | NA | 3.53% | 3.11% | N/A | 61.38 | 65.54 | N/A | 5.28 % | N/A | 0.5 | 1.6 | 6.49 % | N/A | 0.4 | 1.4 | 11.79 % | N/A | 0.3 | 1.2 | NGSP Reference Method y = 0.986x - 0.030 (%HbA1c) r=0.995 n=172 | NGSP Reference Method y = 0.986x - 0.664 (mmol/mol) r=0.995 n=172 | N/A | < 5% change up to 10 mg/dL of Bilirubin, conjugated and unconjugated | < 5% change up to 1000 mg/dL of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | IFCC reference calibrators | 11111592_EN Rev. 01, 2019-01 | IFU Document | 0.6 | 4.134 | 4.134 | 2.756 | 0.272 | 10.5 | 10.5 | 7
Atellica® IM | 10733019 |  | re11 | 00630414245904 | OUS | Reproductive Endocrinology | No | Atellica IM Pregnancy-Associated Plasma Protein-A (PAPP-A) | Atellica IM PAPP-A | PAPPA | 100 | 1 x 100 | NA | Atellica IM Multi-Diluent 13 | 33 | 28 | Anti-PAPP-A mouse monoclonal antibody labeled with acridinium ester | ;#Serum;# | 8 hours at room temperature (RT), 72 hours at 2–8°C, up to 6 months at -20°C Avoid >1 freeze/thaw cycle | 20 μL | No |  | 2-step sandwich immunoassay using direct chemiluminescent technology | 0.01–10.00 IU/L | 0.01–10.00 mIU/mL |  |  |  |  | < 0.01 IU/L (mIU/mL) healthy non-pregnant females | NA | NA | Atellica IM PAPP-A CAL | NA | 8 hrs at room temperature (RT), 30 days at 2–8°C | Liquid | NA | 3 for lot calibration, 2 for pack recalibration | 2 | Human serum; bovine serum | 28 | 14 |  |  |  | Yes | Atellica IM Multi-Diluent 13 | 10995643 |  | 0.01–100.00 IU/L | 0.01–100.00 mIU/mL | No |  |  | NA | NA | NA | NA | 0.005 IU/L | 0.004 IU/L | 0.01 IU/L | 0.004 mIU/mL | 0.005 mIU/mL | 0.01 mIU/mL | 0.32 IU/L | 0.32 mIU/mL | 2.7 | 3.4 | 4.17 IU/L | 4.17 mIU/mL | 2.9 | 4.1 | 9.03 IU/L | 9.03 mIU/mL | 3.6 | 4.4 | ADVIA Centaur PAPP-A y = 0.96x + 0.00 IU/L r = 1.00 | ADVIA Centaur PAPP-A y = 0.96x + 0.00 mIU/mL r = 1.00 | ≤ 10% interference up to 500 mg/dL (0.31 mmol/L) | ≤ 10% interference up to 40 mg/dL (684 μmol/L) conjugated and 60 mg/dL (1026 μmol/L) unconjugated | ≤ 10% interference up to 1000 mg/dL (11.3 mmol/L) | Not used in assay structure | 157 IU/L | 157 mIU/mL |  |  | Standardized to an internal standard | RPBL1308R03_EN Rev. 03, 2018-09 | IFU Document
Atellica® IM | 10733020 |  | re11 | 00630414245911 | OUS | Reproductive Endocrinology | No | Atellica IM Pregnancy-Associated Plasma Protein-A (PAPP-A) | Atellica IM PAPP-A | PAPPA | 500 | 5 X 100 |  | Atellica IM Multi-Diluent 13 | 33 | 28 | Anti-PAPP-A mouse monoclonal antibody labeled with acridinium ester | ;#Serum;# | 8 hours at room temperature (RT), 72 hours at 2–8°C, up to 6 months at -20°C Avoid >1 freeze/thaw cycle | 20 μL | No |  | 2-step sandwich immunoassay using direct chemiluminescent technology | 0.01–10.00 IU/L | 0.01–10.00 mIU/mL |  |  |  |  | < 0.01 IU/L (mIU/mL) healthy non-pregnant females | NA | NA | Atellica IM PAPP-A CAL | NA | 8 hrs at room temperature (RT), 30 days at 2–8°C | Liquid | NA | 3 for lot calibration, 2 for pack recalibration | 2 | Human serum; bovine serum | 28 | 14 |  |  |  | Yes | Atellica IM Multi-Diluent 13 | 10995643 |  | 0.01–100.00 IU/L | 0.01–100.00 mIU/mL | No |  |  | NA | NA | NA | NA | 0.005 IU/L | 0.004 IU/L | 0.01 IU/L | 0.004 mIU/mL | 0.005 mIU/mL | 0.01 mIU/mL | 0.32 IU/L | 0.32 mIU/mL | 2.7 | 3.4 | 4.17 IU/L | 4.17 mIU/mL | 2.9 | 4.1 | 9.03 IU/L | 9.03 mIU/mL | 3.6 | 4.4 | ADVIA Centaur PAPP-A y = 0.96x + 0.00 IU/L r = 1.00 | ADVIA Centaur PAPP-A y = 0.96x + 0.00 mIU/mL r = 1.00 | ≤ 10% interference up to 500 mg/dL (0.31 mmol/L) | ≤ 10% interference up to 40 mg/dL (684 μmol/L) conjugated and 60 mg/dL (1026 μmol/L) unconjugated | ≤ 10% interference up to 1000 mg/dL (11.3 mmol/L) | Not used in assay structure | 157 IU/L | 157 mIU/mL |  |  | Standardized to an internal standard | RPBL1308R03_EN Rev. 03, 2018-09 | IFU Document
Atellica® IM | 10733009 |  | re03 | 00630414245805 | OUS | Reproductive Endocrinology | No | Atellica IM Free Beta Human Chorionic Gonadotropin (FBHCG) | Atellica IM FBHCG | FBHCG | 100 | 1 X 100 | NA | Atellica IM Multi-Diluent 13 | 40 | 28 | Anti-free beta HCG mouse monoclonal antibody labeled | ;#Serum;# | 24 hours at room temperature, up to 6 days at 2–8°C, frozen for up to 6 months at -20°C. Avoid >1 freeze/thaw cycle | 20 μL | No |  | 2-step sandwich immunoassay using direct chemiluminescent technology | 0.14–200.00 IU/L | 0.14–200.00 ng/mL |  |  |  |  | < 0.28 IU/L (ng/mL) healthy non-pregnant females | NA | NA | Atellica IM FBHCG CAL | Included in the kit | Reconstituted at room temperature 8 hours, reconstituted at 2–8°C 30 days | Lyophilized | 2.0 mL of special reagent water. Invert 5 times and let stand for 30 minutes at room temperature | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, human free beta HCG; BSA | 28 | 14 |  |  |  | Yes | Atellica IM Multi-Diluent 13 | 10995643 |  | 0.14–2000.00 IU/L | 0.14–2000.00 ng/mL | No |  |  | NA | NA | NA | NA | 0.11 IU/L | 0.07 IU/L | 0.14 IU/L | 0.07 ng/mL | 0.11 ng/mL | 0.14 ng/mL | 8.31 IU/L | 8.31 ng/mL | 1.5 | 2.3 | 81.17 IU/L | 81.17 ng/mL | 1.3 | 2.6 | 177.70 IU/L | 177.70 ng/mL | 1.3 | 2.3 | ADVIA Centaur FBHCG y = 0.95x - 0.98 IU/L r = 1.00 | ADVIA Centaur FBHCG y = 0.95x - 0.98 ng/mL r = 1.00 | ≤ 10% interference up to 500 mg/dL (0.31 mmol/L) | ≤ 10% interference up to 40 mg/dL (684 μmol/L) conjugated, 60 mg/dL (1026 μmol/L) unconjugated | ≤ 10% interference up to 1000 mg/dL (11.3 mmol/L) | ≤ 10% interference up to 1500 ng/mL (6140 nmol/L) | 78,940 IU/L | 78,940 ng/mL |  |  | WHO 1st International Reference 75/551 | RPBL1319R04_EN Rev. 04, 2019-12 | IFU Document
Atellica® IM | 10733010 |  | re03 | 00630414245812 | OUS | Reproductive Endocrinology | No | Atellica IM Free Beta Human Chorionic Gonadotropin (FBHCG) | Atellica IM FBHCG | FBHCG | 500 | 5 x 100 | NA | Atellica IM Multi-Diluent 13 | 40 | 28 | Anti-free beta HCG mouse monoclonal antibody labeled | ;#Serum;# | 24 hours at room temperature, up to 6 days at 2–8°C, frozen for up to 6 months at -20°C. Avoid >1 freeze/thaw cycle | 20 μL | No |  | 2-step sandwich immunoassay using direct chemiluminescent technology | 0.14–200.00 IU/L | 0.14–200.00 ng/mL |  |  |  |  | < 0.28 IU/L (ng/mL) healthy non-pregnant females | NA | NA | Atellica IM FBHCG CAL | Included in the kit | Reconstituted at room temperature 8 hours, reconstituted at 2–8°C 30 days | Lyophilized | 2.0 mL of special reagent water. Invert 5 times and let stand for 30 minutes at room temperature | 3 for lot calibration, 2 for pack recalibration | 2 | After reconstitution, human free beta HCG; BSA | 28 | 14 |  |  |  | Yes | Atellica IM Multi-Diluent 13 | 10995643 |  | 0.14–2000.00 IU/L | 0.14–2000.00 ng/mL | No |  |  | NA | NA | NA | NA | 0.11 IU/L | 0.07 IU/L | 0.14 IU/L | 0.07 ng/mL | 0.11 ng/mL | 0.14 ng/mL | 8.31 IU/L | 8.31 ng/mL | 1.5 | 2.3 | 81.17 IU/L | 81.17 ng/mL | 1.3 | 2.6 | 177.70 IU/L | 177.70 ng/mL | 1.3 | 2.3 | ADVIA Centaur FBHCG y = 0.95x - 0.98 IU/L r = 1.00 | ADVIA Centaur FBHCG y = 0.95x - 0.98 ng/mL r = 1.00 | ≤ 10% interference up to 500 mg/dL (0.31 mmol/L) | ≤ 10% interference up to 40 mg/dL (684 μmol/L) conjugated, 60 mg/dL (1026 μmol/L) unconjugated | ≤ 10% interference up to 1000 mg/dL (11.3 mmol/L) | ≤ 10% interference up to 1500 ng/mL (6140 nmol/L) | 78,940 IU/L | 78,940 ng/mL |  |  | WHO 1st International Reference 75/551 | RPBL1319R04_EN Rev. 04, 2019-12 | IFU Document
Atellica® CH | 10329216 |  | ge46 | 08058056680150 | OUS | General Chemistry | Yes | Aldolase (Sentinel Alliance Application) | AldoAldo | 7068 | 668 | 4 x 167 | Atellica CH Empty Reagent Packs (11097534) | N/A | 8.9 | 30 per pack | collidin buffer, mono-iodoacetate, F-1.6-DP, sodium azide, NADH, GDH, TPI, LDH and ammonium sulphate | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | 15 days at 2-8 °C | 25 µL | No |  | Colorimetric | 1.0 - 30.0 U/L. | 1.0 - 30.0 U/L. | N/A | N/A | Serum/Plasma: ≤ 7.6 U/L (37 °C); ≤ 3.1 U/L (25 °C) | Serum/Plasma: ≤ 7.6 U/L (37 °C); ≤ 3.1 U/L (25 °C) | N/A | N/A | N/A | Clin Chem Calibrator (Sentinel) | 10313439 |  | Lyophilized | Add 3.0 mL of distilled water. Leave to stand for 30 minutes. Gently invert to ensure the homogeneity of the solution. | 3 | 2 | human based serum stabilized by adding preservatives | 30 per pack | 30 per pack |  | No |  | Yes | N/A | 11099300 |  | N/A | N/A | No | N/A | N/A | 1.0 U/L | 1.0 U/L | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 7.0 U/L | 7.0 U/L | 4.3 | 4.8 | 16.0 U/L | 16.0 U/L | 2.4 | 2.9 | 30.0 U/L | 30.0 U/L | 2.5 | 2.9 | Sentinel Aldolase on BC AU680 y=0.9641x + 0.0855 U/L r = 0.999 n=126 | Sentinel Aldolase on BC AU680 y=0.9641x + 0.0855 U/L r = 0.999 n=126 | Avoid hemolysed samples | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Internal standard | IFU: 17003-2.0/ 02 2015/07/16 and App. Sheet: 17003-1.0/02 F57 2019/04/24 |  | 0.6614 | 7.99 | 3.23 | 4.69 | 0.3 | 20.3 | 8.2 | 11.9
Atellica® CH | 11419175 |  | ge47 | 08058056682000 | OUS | General Chemistry | Yes | Total Bile Acids Liquid (Sentinel Alliance Application) | TBATBA | 7066 | 334 | 2 x 167 | Atellica CH Empty Reagent Packs (SMN 11097534) Sentinel Total Bile Acids Control Set (SMN 11532786) | N/A | 8.5 | 30 per pack | Thio-NAD⁺, Detergent, Sodium Azide, Good’s Buffer, NADH, 3α-HSD, Buffer, Conjugate Bile Acid | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#Potassium EDTA Plasma;# | Serum: 1 day at 20 °C, 1 weeks at 2-8 °C or 12 months at -20 °C. Plasma: 6 hr at RT, 48 hr at 2-8 °C, 2 months at -20 °C or 2 months at -70 °C. | 12 µL | No |  | Enzymatic Cycling Colorimetric | 1.0 – 180 μmol/L | 1.0 – 180 μmol/L | N/A | N/A | Male (fasting): 1.0 - 6.0 μmol/L; Female (fasting): 1.0 - 6.0 μmol/L | Male (fasting): 1.0 - 6.0 μmol/L; Female (fasting): 1.0 - 6.0 μmol/L | N/A | N/A | N/A | Total Bile Acids Standard | included in the kit | 28 | Liquid | N/A | 3 | 2 | sodium taurodeoxycholate (Sorachim cat # TAU-100) | 30 per pack | 30 per pack |  | No |  | No | N/A |  |  | N/A | N/A | Yes | Total Bile Acids Control Set | 11532786 | NA | NA | NA | NA | 0.19 μmol/L | 0.10 μmol/L | 0.4 μmol/L | 0.10 μmol/L | 0.19 μmol/L | 0.4 μmol/L | 3.0 µmol/L | 3.0 µmol/L | 2.1 | 3.2 | 18.0 µmol/L | 18.0 µmol/L | 0.8 | 1.3 | 180.0 µmol/L | 180.0 µmol/L | 0.7 | 1.2 | Randox TBA on ARCHITECT c8000 y = 1.07 x + 0.3032 µmol/L r = 0.999 n=144 | Randox TBA on ARCHITECT c8000 y = 1.07 x + 0.3032 µmol/L r = 0.999 n=144 | ≤ 10% change (or ± 0.6 μmol/L for samples < 6 μmol/L) up to 1000 mg/dL of hemoglobin | ≤ 10% change (or ± 0.6 μmol/L for samples < 6 μmol/L) up to 22.4 mg/dL of conjugated bilirubin and up to 29.9 mg/dL of unconjugated bilirubin | ≤ 10% change (or ± 0.6 μmol/L for samples < 6 μmol/L) up to 1703 mg/dL of Trig and up to 750 mg/dL of Intralipid |  | N/A | N/A | N/A | N/A | Internal Master Calibrator | IFU: 1700201 - 1.0/ 02 2018/04/13 and App sheet: F57 2019/04/24 |  | 0.833 | 4.37 | 2.42 | 4.68 | 0.378 | 11.1 | 6.15 | 11.9
Atellica® CH | 11533403 |  | ca03 | 08058056680549 | OUS | General Chemistry | Yes | CK-MB Liquid (Sentinel Alliance Application) | eCKMBeCKMB | 7067 | 320 | 1 x 320 | Atellica CH Empty Reagent Packs (SMN 11097534) | N/A | 8.5 | 30 per pack | monoclonal murine antibodies anti-CK-M | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | 3 days at 2-8 °C; 4 weeks at -20°C | 25 µL | No |  | Enzymatic | 3.0 - 1000 U/L | 3.0 - 1000 U/L. | 3.0 - 10000 U/L | 3.0 - 10000 U/L | ≤ 24 U/L | ≤ 24 U/L | N/A | N/A | N/A | CK-MB Calibrator (Sentinel) | 10284952 | 2 | Lyophilized | Add 1.0 mL of distilled water. Replace stopper and gently swirl. Allow reconstituted material to stand for at least 60 minutes. Gently invert just prior to use. Avoid foaming. | 3 | 2 | Human CK-MM and porcine brain CK-BB in bovine albumin matrix. | 30 per pack | 30 per pack |  | No |  | Yes | N/A | 11099300 |  | 3.0 - 10000 U/L | 3.0 - 10000 U/L | No | N/A | N/A | 3.0 U/L | 3.0 U/L | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 20 U/L | 20 U/L | 4.7 | 7 | 300 U/L | 300 U/L | 0.6 | 1.5 | 1000 U/L | 1000 U/L | 0.2 | 1.4 | Architect c16000 CK-MB y = 0.94 x + 2 U/L r = 1.000 n = 120 | Architect c16000 CK-MB y = 0.94 x + 2 U/L r = 1.000 n = 120 | ≤ 10% change up to 50 mg/dL of hemoglobin | ≤ 10% change up to 50 mg/dL of bilirubin | ≤ 10% change up to 600 mg/dL of intralipid | N/A | N/A | N/A | N/A | N/A | Inhibition method according to IFCC mod. recommendations | IFU: 17649H - 2.0/ 02 2015/04/14; App. Sheet: F57 2019/04/24 |  | 0.364 | 4.37 | 2.42 | 4.69 | 0.165 | 11.1 | 6.15 | 11.9
Atellica® CH | 10317329 |  | ge49 | 08058056680013 | OUS | General Chemistry | Yes | Zinc (Sentinel Alliance Application) | ZnZn | 7086 | 308 | 2 x 154 | Atellica CH Empty Reagent Packs (11097534) | N/A | 9.7 | 34 | good's buffer, specific masking agents for iron and copper, ascorbic acid, complexant 5-Br-PAPS | ;#Serum;#Heparinized Plasma;# | 14 days at 2-8 °C | 25 | No |  | Colorimetric | 5.0 – 2000 μg/dL | 5.0 – 2000 μg/dL | N/A | N/A | Adults; 70 - 120 μg/dL | Adults; 70 - 120 μg/dL | N/A | N/A | N/A | Zinc Standard | included in the kit |  | Liquid | N/A | 3 | 2 |  | 41 | 34 |  | No |  | Yes | N/A | 11099300 |  | N/A | N/A | No | N/A | N/A | 5.0 μg/dL | 5.0 μg/dL | NA | NA | N/A | N/A | N/A | N/A | N/A | N/A | 116.55 μg/dL | 116.55 μg/dL | 0.6 | 1.1 | 540.97 μg/dL | 540.97 μg/dL | 0.2 | 0.4 | 1068.23 μg/dL | 1068.23 μg/dL | 0.2 | 0.4 | Randox TBA on ARCHITECT c8000 y = 1.07 x + 0.3032 µmol/L r = 0.999 n=144 | Randox TBA on ARCHITECT c8000 y = 1.07 x + 0.3032 µmol/L r = 0.999 n=144 | ≤ 10% change up to 500 mg/dL of hemoglobin. | ≤ 10% change up to 40 mg/dL of bilirubin. | ≤ 10% change up to 1000 mg/dL of Trig. |  | N/A | N/A | N/A | N/A | NIST (National Institute of Standards and Technology). | IFU: 17640H - 2.0/ 02 2015/04/15. App Sheet: 17640H-1.0/02 F57 2019/07/25 |  | 0.708 | 5.59 | 3.23 | 4.69 | 0.321 | 14.2 | 8.2 | 11.9
Atellica® CH | 11533083 |  | ge48 | 08058056680020 | OUS | General Chemistry | Yes | Copper (Sentinel Alliance Application) | CopCop | 8011 | 567 | 3 x 189 | Atellica CH Empty Reagent Packs (11097534) | N/A | 9.7 | 30 | malate dehydrogenase (MDH) (mammalian); PEP: microorganism (PEPC); | ;#Serum;#Heparinized Plasma;# | 14 days at 2-8 °C | 20 | No |  | Colorimetric | 3.0 – 500 μg/dL | 3.0 – 500 μg/dL | N/A | N/A | Male: 70 - 140 μg/dL; Female : 80 - 155 μg/dL | Male: 70 - 140 μg/dL; Female : 80 - 155 μg/dL | N/A | N/A | N/A | Copper Standard | included in the kit |  | Liquid | N/A | 3 | 2 |  | 41 | 30 |  | No |  | Yes | N/A | 11099300 |  | N/A | N/A | No | N/A | N/A | 3.0 μg/dL | 3.0 μg/dL | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 97.59 μg/dL | 97.59 μg/dL | 1.4 | 2.4 | 317.70 μg/dL | 317.70 μg/dL | 0.5 | 0.7 | 414.32 μg/dL | 414.32 μg/dL | 0.3 | 0.6 | Beckman AU5800 y = 1.030x -5.466 μg/dL r = 0.999 n = 110 | Beckman AU5800 y = 1.030x -5.466 μg/dL r = 0.999 n = 110 | ≤ 10% change up to 300 mg/dL of hemoglobin | ≤ 10% change up to 40 mg/dL of bilirubin | ≤ 10% change up to 1500 mg/dL of lipids | N/A | N/A | N/A | N/A | N/A | NIST (National Institute of Standards and Technology) | IFU: 17638H - 2.0/ 02 2015/04/21. App. Sheet: 17638H-1.0/02 F57 2019/07/25 | IFU Document | 0.745 | 5.59 | 3.23 | 4.69 | 0.34 | 14.2 | 8.2 | 11.9
Atellica® CH | 11309910 |  | ge50 | 08058056681751 | Global | General Chemistry | Yes | Angiotensin Converting Enzyme (Sentinel Alliance Application) | ACEACE | 7069 | 308 | 2 x 154 | Atellica CH Empty Reagent Packs ( SMN 11097534) | N/A | 9.7 | 30 | Human specimens | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;# | 7 days at 2-8 °C or 1 year at -20 °C | 100 | No |  | Enzymatic | 8.4 - 120 U/L | 8.4 - 120 U/L | 8.4 - 1200 U/L | 8.4 - 1200 U/L | 13.3 – 63.9 U/L | 13.3 – 63.9 U/L | N/A | N/A | N/A | Sentinel ACE Calibrator | 11309911 | 7 | Lyophilized | Add 1.0 mL of distilled or deionized water. Leave to stand for 5 minutes. Swirl each vial just prior to use. | 3 | 2 | Human serum containing porcine ACE | 134 | 30 |  | No |  | Yes | N/A | N/A |  | 8.4 - 1200 U/L | 8.4 - 1200 U/L | Yes | Sentinel ACE Control Set | 11309912 | N/A | N/A | N/A | N/A | 4.5 U/L | 2.4 U/L | 8.4 U/L | 2.4 U/L | 4.5 U/L | 8.4 U/L | 20 U/L | 20 U/L | 7.1 | 8.6 | 75 U/L | 75 U/L | 1.4 | 4.7 | 120 U/L | 120 U/L | 1.1 | 4.4 | Sentinel ACE on Atellica CH vs. Beckman AU680 y = 0.91x + 4.27 r = 0.984 n = 115 | Sentinel ACE on Atellica CH vs. Beckman AU680 y = 0.91x + 4.27 r = 0.984 n = 115 | ≤ 10% change up to 100 mg/dL of hemoglobin | ≤ 10% change up to 13 mg/dL of unconjugated bilirubin; ≤ 10% change up to 26 mg/dL of conjugated bilirubin | ≤ 10% change up to 200 mg/dL of Trig | N/A | N/A | N/A | N/A | N/A | Internal master calibrator material. | IFU: 1767001-3.0/ 02 2020/03/04. App Sheet: 1767001-2.0/02 F57 2020/05/20 |  | 0.3 | 4.37 | 2.42 | 4.68 | 0.135 | 11.1 | 6.15 | 11.9
Atellica® IM | 10995648 |  | ca07 | 00630414599663 | Global | Cardiac | No | Atellica IM Myoglobin (MYO) | Atellica IM MYO | MYO | 250 | 1 x 100 | Atellica IM CAL U (SMN 10995519) | Atellica IM Multi-Diluent 10 (SMN 10995640) Atellica IM MYO MCM (SMN 10995650) | 10 | 28 | Goat polyclonal anti-myoglobin antibody/mouse monoclonal anti-myoglobin antibody | ;#Serum;#Heparinized Plasma;# | Up to 8 hrs at RT. 48hr 2-8C, freeze 1 time | 10 uL | No |  | direct chemiluminescent technology | 3.00 - 1000.00 ng/mL | 3.00 - 1000.00 ug/L |  |  |  |  | A myoglobin value of >110 ng/mL (ug/L) is highly suggestive of myocardial infarct |  |  | Atellica IM CAL U | 10995519 (2-pack) | 4 hrs at RT, 24 hrs recon at 2-8C, 60 days recon at -80 to -60C | Lyophilized | reconstitute 2.0 mL/Vial | 3 for lot calibration, 2 for pack recalibration | 2 |  | 33 | 14 |  |  |  | Yes | Atellica IM Multi-Diluent 10 | 10995640 |  | 3.00 - 20,000 ng/mL | 3.00 - 20,000.00 ug/L | No |  |  | 2.53 ng/mL | 2.53 ug/L | NA | NA | 3.57 ng/mL | 2.56 ng/mL | NA | 2.56 ug/L | 3.57 ng/mL | NA | 155.95 mg/dL (Control) | 155.95 ug/L (Control) | 2 | 2.9 | 361.51 mg/dL (Control) | 361.51 ug/L (Control) | 2.1 | 3 | 832.09 mg/dL | 832.09 ug/L | 2.2 | 2.9 | ADVIA Centaur Myoglobin y=0.96x + 2.80 ng/mL r=1.00 | ADVIA Centaur Myoglobin y=0.06x + 2.80 ug/L r=1.00 | < 5% change up to 1000 mg/dL of hemoglobin | < 5% change up to 40 mg/dL conjugated bilirubin 40 mg/dL of unconjugated bilirubin | <5% change up to 1000 mg/dL of triglycerides | No Biotin in assay formulation | 200,000 ng/mL | 200,000 ug/L |  |  | Traceable to an internal standard manufactured using highly purified material | 10995411_EN Rev. 02, 2019-07 | IFU Document | 0.176 | 6.969 | 2.087 | 1.457 | 0.08 | 17.7 | 5.3 | 3.7
Atellica® IM | 11203564 |  | id14 | 00630414285641 | OUS | ToRCH | No | Atellica IM Rubella G II (RbGII) | RbGII | RbGII | 100 |  | 10995458 Atellica IM APW1 (probe wash) | 11203576 Atellica IM RbGII QC (quality control material) 10995644 Atellica IM Multi‑Diluent 2 (diluent) 11203585 4 x 1.0 mL levels of master curve material | 42 | 60 | Mouse anti-human IgG monoclonal antibody | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#Potassium EDTA Plasma;# | 5 days at room temperature, and for up to 14 days at 2–8°C. Specimens devoid of red blood cells may be frozen for up to 7 months ≤ -20°C. Do not store in a frost-free freezer. | 15 μl | No |  | 2-step sandwich immunoassay that uses direct, chemiluminescent technology | 0.68–400.00 IU/mL | [5 empty] | Nonreactive: < 10.00 IU/mL; Reactive: ≥ 10.00 IU/mL |  |  | 1 vial Atellica IM RbGII CAL low calibrator 1 vial Atellica IM RbGII CAL high calibrator Atellica IM RbGII CAL lot-specific value sheet | Included in kit | 65 days 2–8°C, 8 hours room temperature | Liquid | N/A | 3 for lot calibration, 2 for pack recalibration | 2 | Processed human plasma negative and positive for anti-rubella IgG antibodies; preservatives | 70 | 60 |  |  |  | No | Atellica IM Multi‑Diluent 2 ReadyPack ancillary reagent pack | Included |  |  |  | Yes | Atellica IM RbGII QC (quality control material) | 11203576 | [10 empty] | 23.18 |  | 2.5 | 5.1 |  |  |  |  | 221.66 |  | 4.3 | 6.5 |  |  | ≤ 10% change in results up to 500 mg/dL of hemoglobin | ≤ 10% change in results up to 60 mg/dL of conjugated bilirubin, 40 mg/dL of unconjugated bilirubin | ≤ 10% change in results up to 1000 mg/dL of lipemia (Intralipid) | [5 empty] | Traceable to WHO 1st International Standard for human anti-rubella immunoglobulin (RUBI‑1‑94) | 11203565_EN Rev. 01, 2019-10 | IFU Document
Atellica® IM | 11200558 |  | re12 |  | Global | Reproductive Endocrinology | No | Atellica IM Androstenedione (ANDRO) | ANRO | ANDRO | 50 | 1x50 | NA | ANDRO Master Curve Material (SMN11200557) Atellica IM Multi-Diluent 1 | 14 | 30 | Biotinylated sheep monoclonal anti-androstenedione antibody | ;#Serum;#Lithium Heparin Plasma;#Potassium EDTA Plasma;# | up to 24 hours at room temperature, up to 31 days at 2–8°C, up to 12 months at ≤ -20°C. | 20 μL | No |  | 1-step, competitive immunoassay | 0.30 – 9.00 ng/mL | 1.05 – 31.42 nmol/L |  |  |  |  | Adult males < 0.30–1.65 ng/mL (1.05–5.76 nmol/L) (See IFU for pediatric ranges) | Females 22–39 years 0.37–2.25 ng/mL (1.29–7.86 nmol/L) | Females 40–54 years < 0.30–1.90 ng/mL (1.05–6.63 nmol/L) | ANDRO low and high calibrators | Included in the kit | 24 hrs recon at room temperature, 28 days recon at 2–8°C, 60 days recon at ≤ -20°C | Lyophilized | 2.0 mL of special reagent water, stand for 15–20 minutes at room temperature | 3 for lot calibration, 2 for pack recalibration | 2 | low and high levels of androstenedione | 90 | 30 |  |  |  | Yes | Atellica IM Multi‑Diluent 1 | 10995637 (2-pack),10995638 (6-pack),10995639 (vial) |  | 0.30 – 45.00 ng/mL | 1.05 – 157.10 nmol/L | No |  |  | NA | NA | NA | NA | 0.16 ng/mL | 0.13 ng/mL | 0.17 ng/mL | 0.45 nmol/L | 0.56 nmol/L | 0.59 nmol/L | 0.48 ng/mL | NA | 3.5 | 4.6 | 1.06 ng/mL | NA | 2.5 | 3.1 | 7.34 ng/mL | NA | 2.8 | 4 | ADVIA Centaur ANDRO assay (x) y = 0.93x + 0.03 ng/mL r= 0.992 | ADVIA Centaur ANDRO assay (x) y = 0.93x + 0.10 nmol/L r= 0.992 | ≤ 10% change up to 2 g/L of hemoglobin | ≤ 10% change up to .2 g/L of conjugated and unconjugated bilirubin, | ≤ 10% change up to 1000 mg/dL (intralipid) | no change up to > 500 ng/mL | NA | NA |  |  | traceable to an internal standard manufactured using highly purified material | 11203697_EN Rev.02, 2020-04 | IFU Document
Atellica® IM | 11200643 |  | id10 | 00630414246963 | OUS | Hepatitis | No | Atellica IM HBsAgII Quant | Atellica IM QHBsQHBs | QHBs | 200 | 1x200 | Atellica IM PW3 Atellica IM Multi-Diluent 14 | QHBs QC MCMs | 26 | 90 | monoclonal mouse anti-HBsAg | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#EDTA Plasma;# | 8 hrs at room temp,separated specimens. 7 days at 2-8C in primary tubes, 14 days at 2-8C in separated specimens, -20C for longer storage | 100 uL | No |  | sandwich, direct chemiluminescent | 0.017-125.0 IU/mL undiluted | 0.017-125.0 IU/mL undiluted | NA | NA | NA | NA | see IFU | N/A | N/A | Atellica IM QHBs CAL | provided in kit | 60 days at 2-8C | Liquid | N/A | 3 for lot calibration, 2 for pack recalibration | 2 | purified human HBsAg in buffer with sodium azide | 21 | 21 |  |  |  | Yes | Atellica IM Multi-Diluent 14 | 11200662 |  | 0.017-312,500 IU/mL | 0.017-312,500 IU/mL | Yes | Atellica IM QC | 11200644 | N/A | N/A | N/A | N/A | 0.014 IU/mL | 0.007 IU/mL | 0.017 IU/mL | 0.007 IU/mL | 0.014 IU/mL | 0.017 IU/mL | 0.659 | 0.659 | 1.4 | 2.6 | 305.6 | 305.6 | 2.7 | 3.6 | 1506 | 1506 | 3.6 | 5.7 | N/A | N/A | ≤ 10% interference up to 500 mg/dL of hemoglobin | ≤ 10% interference up to 40 mg/dL of bilirubin | ≤ 10% up to 1000 mg/dL of triglycerides | ≤ 10% up to 10 ng/mL of biotin | 3.5 mg/mL | 3.5 mg/mL |  |  | (WHO) 3rd International Standard for HBsAg, NIBSC code 12/226 | 11200655_EN Rev. 03, 2019-02 | IFU Document |  |  |  |  | 0.179 | 17.5 | 5.5 | 6.3
Atellica® CH | 11097543 |  | ge09 | 00630414279206 | Global | General Chemistry | No | Total Protein, CSF/Urine | UCFPUCFP | 4049 | 1480 | 4 x 370 | N/A | N/A | 5 | 30 | Pyrogallol red, methanol, sodium molybdate | ;#Urine;#Cerebrospinal fluid (CSF);# | Urine: 3 days when stored at 4°C with no additives; After 24 hours, store at 2–4°C for < 72 hours or frozen at -20°C for up to 1 year; CSF: Analyze fresh specimens or store at 4°C8 for less than 72 hours. Frozen specimens are stable at -20°C for 6 months. | 6.3 | No |  | pyrogallol red-molybdate | 6.0 – 250.0 mg/dL | 60 – 2500 mg/L | 6.0 – 2500 mg/dL | 60 – 25,000 mg/L | Urine: 1 -14 mg/dL CSF: 15 -45 mg/dL | Urine:10 – 140 mg/L CSF: 150 – 450 mg/L | N/A | N/A | N/A | Urinary/Cerebrospinal Fluid Protein Calibrator (UCFP CAL) | 11099339 | 60 | Liquid | No preparation required | 3 | 5 | Human serum albumin and IgG | 60 | 7 |  | No |  | Yes | N/A | 11099300 |  | 6.0 – 2500 mg/dL | 60 – 25,000 mg/L | No | N/A | N/A | 6.0 mg/dL | 60 mg/L | N/A | N/A | 6.0 mg/dL | 3.0 mg/dL | 6.0 mg/dL | 30 mg/L | 60 mg/L | 60 mg/L | 23.3 mg/dL | 233 mg/L | 2.4 | 3.4 | 70.8 mg/dL | 708 mg/L | 0.8 | 1.4 | 149.4 mg/dL | 1494 mg/L | 0.6 | 1.2 | Urine Dimension UCFP method y = 1.01x - 1.0 mg/dL r=0.999 n=109 CSF Dimension UCFP method y = 1.05x - 1.5 mg/dL r=0.999 n=118 | Urine Dimension UCFP method y = 1.01x - 10 mg/L r=0.999 n=109 CSF Dimension UCFP method y = 1.04x - 15 mg/L r=0.999 n=118 | N/A | N/A | N/A | No Biotin in assay formulation | N/A | N/A | N/A | N/A | Traceable to the NIST-SRM 927 | 11111705_EN Rev. 01, 2020-01 | IFU Document | 0.266 | 7.302 | 8.89 | 6.985 | 0.121 | 18.55 | 22.58 | 17.74
Atellica® IM | 11204647 |  | si02 |  | Global | Special ID | No | Atellica IM Zika Ab (ZikaAb) - Must be run with Atellica IM Zika IgM (ZikaM) | ZikaAb | ZikaAb | 100 |  | Atellica IM Zika IgM (ZikaM) 11204645 Atellica IM ZikaAb QC 11204648 Atellica IM ZikaM QC 11204646 |  | 39 | 28 | NS1 antigen | ;#Serum;#Lithium Heparin Plasma;#Potassium EDTA Plasma;# | After centrifugation, specimens in the primary collection device in contact with the cells or clot are stable for up to 24 hours at room temperature, and for up to 5 days at 2–8°C. • Separated samples are stable for up to 24 hours at room temperature, and for up to 5 days at 2–8°C. • Separated samples, devoid of red blood cells, are stable at ≤ -20°C for up to 150 days.Avoid more than 3 freeze-thaw cycles. Do not store in a frost-free freezer. | 15 μL | No |  | Antibody capture immunoassay using a 2-pass format | 0.00–9.00 Index | [5 empty] | < 0.80 Index (nonreactive), ≥ 0.80 Index (equivocal) using the Atellica IM ZikaAb, assay must be reflex tested using the Atellica IM ZikaM assay |  |  | Included |  | Opened at 2–8°C for 60 days | N/A | N/A | 3 for lot calibration, 2 for pack recalibration | 2 | Processed human plasma negative or positive for Zika antibodies; | 56 | 14 |  |  |  | No | NA |  |  |  |  | Yes | Atellica IM ZikaAb QC | 11204648 | [24 empty] | ≤ 10% change up to 1000 mg/dL of hemoglobin | ≤ 10% change up to 40 mg/dL of conjugated bilirubin | ≤ 10% change up to 3000 mg/dL of triglycerides | ≤ 10% change up to 3500 ng/mL of biotin |  |  |  |  | Based on agreement with known Zika positive samples. | 11204673_EN Rev. 01, 2019-11, 11205700_EN Rev. 01, 2019-11 | IFU Document
Atellica® IM | 11204645 |  | si02 |  | Global | Special ID | No | Atellica IM Zika IgM (ZikaM) - Must be run with Atellica IM Zika Ab (ZikaAb) | ZikaM | ZikaM | 50 | 50 tests | Atellica IM Zika Ab (ZikaAb) 11204647 Atellica IM ZikaAb QC 11204648 Atellica IM ZikaM QC 11204646 |  | 39 | 28 | Anti-human IgM monoclonal antibody | ;#Serum;#Lithium Heparin Plasma;#Potassium EDTA Plasma;# | • After centrifugation, specimens in the primary collection device in contact with the cells or clot are stable for up to 24 hours at room temperature, and for up to 5 days at 2–8°C. • Separated samples are stable for up to 24 hours at room temperature, and for up to 5 days at 2–8°C. • Separated samples, devoid of red blood cells, are stable at ≤ -20°C for up to 150 days. Avoid more than 3 freeze-thaw cycles. Do not store in a frost-free freezer. | 15 μL | No |  | IgM capture immunoassay using a 2-pass format | 0.00–9.00 Index | [5 empty] | Must be reflexed from ZikaAb. < 1.00 Index (nonreactive), ≥ 1.00 Index (reactive) using the Atellica IM ZikaM assay must be retested in duplicate (see IFU for more details). |  |  | Included |  | Opened at 2–8°C 60 days | N/A |  | 3 for lot calibration, 2 for pack recalibration | 2 | Processed human plasma negative or positive for Zika IgM antibodies | 49 | 14 |  |  |  | No | NA |  |  |  |  | Yes | Atellica IM ZikaM QC | 11204646 | [24 empty] | ≤ 10% change up to 1000 mg/dL of hemoglobin | ≤ 10% change up to 40 mg/dL of conjugated bilirubin | ≤ 10% change up to 3000 mg/dL of triglycerides | ≤ 10% change up to 3500 ng/mL of biotin |  |  |  |  | Based on agreement with known Zika positive samples. | 11205700_EN Rev. 01, 2019-11, 11204673_EN Rev. 01, 2019-11 | IFU Document
Atellica® IM | 11206923 |  | si03 | 00630414605586 | OUS | Special ID | No | Atellica IM SARS-CoV-2 Total | Atellica IM COV2T | COV2T | 500 | 5 x 100 tests | Atellica IM COV2T QC SMN 11206712 | NA | 10 | 28 | Streptavidin-coated paramagnetic microparticles preformed with biotinylated S1 RBD antigen, bovine serum albumin, goat serum | ;#Serum;#Lithium Heparin Plasma;#Potassium EDTA Plasma;# | stable onboard for 24 hrs. separated samples- 3 days at RT, 5 days at 2-8 C or freeze at -20 C for longer storage | 50 | No |  | 1-step antigen sandwich immunoassay using acridinium ester chemiluminescent technology | 0.60-75.0 Index | 0.60-75.00 Index |  |  |  |  | Nonreactive: <1.0 Index | Reactive: ≥ 1.0 Index | NA | Atellica IM COV2T CAL | included in reagent kit | 8 hrs at RT, 60 days opened at 2-8C | Liquid | Ready to use | 3 for lot calibration, 2 for pack recalibration | 2 | Processed human plasma nonreactive for antibodies to SARS-CoV-2 and processed human plasma spiked with antiboies to SARS-CoV-2 | 28 | 14 |  |  |  | Yes | ATellica IM Mult-Diluent 2 | 10995644 |  | 0.60 -375.00 Index | 0.60 -375.00 Index | Yes | Atellica IM COV2T | 11206712 | NA | NA | NA | NA | 0.5 | 0.4 | 0.6 | 0.4 | 0.5 | 0.6 | 2.72 Index | 2.72 Index | 2.1 | 7.6 | 10.84 Index | 19.84 Index | 2.6 | 9.2 | 65.28 Index | 65.78 Index | 2.7 | 10.1 | N/A | N/A | <10% change up to 1000 mg/dL of hemoglobin | <10% change up to 40 mg/dL of conjugated or unconjugated bilirubin | <10% change up to 2000 mg/dL of triglycerides (Intralipid) | <10% change up to 3500 ng/dL of biotin | NA | NA |  |  | Based on agreement with known positive and negative SARS-CoV-2 samples | 11206928_EN Rev. 02, 2020-10 | IFU Document |  |  |  |  | 0.5 | 17.7 | 5.5 | 17.5
Atellica® IM | 11206711 |  | si03 | 00630414605593 | OUS | Special ID | No | Atellica IM SARS-CoV-2 Total | Atellica IM COV2T | COV2T | 100 | 1 x 100 tests | Atellica IM COV2T QC SMN 11206712 | NA | 10 | 28 | Steptavidin-coated paramagnetic microparticles preformed with biotinylated S1 RBD antigen, bovine serum albumin, goat serum | ;#Serum;#Lithium Heparin Plasma;#Potassium EDTA Plasma;# | stable onboard for 24 hrs, separated samples - 3 days at RT, 5 days at 2-8C or freeze at -20C for longer storage | 50 | No |  | 1-step antigen sandwich immunoassay using acridinium ester chemiluminescent technology | 0.60-75.00 Index | 0.60-75.00Index |  |  |  |  | Nonreactive :<1.0 Index | Reactive: ≥1.0 Index | NA | Atellica IM COV2T CAL | included in reagent kit | 8 hrs at RT, 60 days opened at 2 - 8 C | Liquid | Ready to use | 3 for lot calibration, 2 for pack recalibration | 2 | Processed human plasma nin reactive for antibodies to SARS-CoV-2 and processed plasma spiked with antibodies to SARS-CoV-2 | 28 | 14 |  |  |  | Yes | Atellica IM Mult-Diluent 2 | 10995644 |  | 0.60-375 Index | 0.60-375 Index | Yes | Atellica IM COV2T QC | 11206712 | NA | NA | NA | NA | 0.5 | 0.4 | 0.6 | 0.4 | 0.5 | 0.6 | 2.72 Index | 2.72 Index | 2.1 | 7.6 | 10.84 Index | 10.84 Index | 2.6 | 9.2 | 65.28 | 65.28 | 2.7 | 10.1 | N/A | N/A | ≤ 10% change up to 1000 mg/dL of hemoglobin | ≤ 10% change up to 40 mg/dL of conjugated or unconjugated bilirubin | ≤ 10% change up to 2000 mg/dL of triglycerides (Intralipid) | ≤ 10% change up to 3500 ng/dL of biotin | NA | NA |  |  | Based on agreement with known postive and negative SARS-CoV-2 samples | 11206928_EN Rev.02, 2020-10 | IFU Document |  |  |  |  | 0.13 | 17.9 | 6.7 | 7.4
Atellica® CH | 11532568 |  | ge35 | 00630414287270 | Global | General Chemistry | No | Atellica CH Lactate_2 (Lac_2) (Plasma) | Lac_2Lac_2 | 8010 | 380 | 4 x 95 | N/A | N/A | 4 | 30 days per pack | PIPES buffer, TOOS, Peroxidase (plant), lactate oxidase and ascorbate oxidase (Microorganism);4-aminoantipyrine | ;#Lithium Heparin Plasma;#Potassium EDTA Plasma;#Sodium fluoride/potassium oxalate Plasma;# | Plasma: 1 day at 2–8°C; 30 days at -20°C. | 4 μL | No |  | lactate oxidase to pyruvate and hydrogen peroxide | Plasma: 0.6–110.0 mg/dL | Plasma: 0.07–12.21 mmol/L | Plasma: 0.6 - 1100.0 mg/dL | Plasma: 0.07 - 122.10 mmol/L | Plasma:4.5–19.8 mg/dL | Plasma: 0.50–2.20 mmol/L | N/A | N/A | N/A | Atellica CH SPCL CHEM CAL | 11099438 | 7 | Lyophilized | 5.0 mL of distilled water; Stand for 30 minutes at room temperature; Invert to mix Prior to use mix the contents by gently inverting the vials | 3 | 2 | Human serum | 62 | 30 |  | No |  | Yes | N/A | 11099300 |  | Plasma: 0.6 - 1100.0 mg/dL | Plasma: 0.07 - 122.10 mmol/L | No | N/A | N/A | 0.6 mg/dL | 0.07 mmol/L | N/A | N/A | 0.6 mg/dL | 0.4 mg/dL | N/A | 0.04 mmol/L | 0.07 mmol/L | N/A | 12.5 mg/dL | 1.39 mmol/L | 1.9 | 4.9 | 48.7 mg/dL | 5.41 mmol/L | 0.9 | 4.8 | 99.0 mg/dL | 10.99 mmol/L | 0.8 | 4.6 | ADVIA Chemistry 1800 Lac (plasma) y = 1.01x - 0.2 mg/dL r=0.995 n =101 | ADVIA Chemistry 1800 Lac (plasma) y = 1.01x - 0.02 mmol/L r=0.995 n =101 | ≤ 10% change up to 1000 mg/dL of hemoglobin | <10% change up to 9.2 mg/dL of conjugated bilirubin; <10% change up to 7.5 dL mg/dL of unconjugated bilirubin | < 10% change up to 1000 mg/dL of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | Internal standard using highly purified material | 11417138_EN 11417138_EN Rev. 02, 2020-09 | IFU Document | 0.442 | 5.625 | 3.5 | 2.75 | 0.201 | 14.287 | 8.89 | 6.985
Atellica® IM | 10733011 |  |  |  | OUS | Inflamation | No | Atellica IM Interleukin-6 (IL6) | Atellica IM IL6 | IL6 | 100 | 1 x 100 | NA | Atellica IM IL6 QC Atellica IM Multi-Diluent 13 | 14 | 42 | Monoclonal mouse anti-IL-6 antibody | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | 5 hours at room temperature, 24 hours at 2–8°C,frozen for 1 month at ≤ -20°C. Avoid more than 3 freeze/thaw cycles. | 50 μL | No |  | one-step direct immunoassay using chemiluminescent technology | 2.7 – 5500.0 pg/mL | NA |  |  |  |  | up to 4.4 pg/mL | NA | NA | Atellica IM IL6 CAL | Included in the kit | Reconstituted at 2–8°C 30 days, at room temperature 8 hours | Lyophilized | 2.0 mL of special reagent water. Invert vials 5 times, let stand for 30 minutes at room temperature. | 3 for lot calibration, 2 for pack recalibration | 2 | recombinant human IL-6; bovine serum albumin | 42 | 14 |  |  |  | Yes | Atellica IM Multi‐Diluent 13 | 10995643 |  | 2.7 – 55,000 pg/mL | NA | Yes | Atellica IM IL6 | 10733014 | NA | NA | NA | NA | 1.0 pg/mL | 0.6 pg/mL | 2.7 pg/mL | NA | NA | NA | 5.1 pg/mL | NA | 7.9 | 9.6 | 249.3 pg/mL | NA | 2.7 | 3.1 | 3981.6 pg/mL | NA | 2.1 | 2.5 | ADVIA Centaur IL6 (x) y = 0.97x + 0.08 pg/mL r= 1.00 | NA | ≤ 10% interference up to 1000 mg/dL (0.62 mmol/L) | ≤ 10% interference up to 25 mg/dL (427.5 μmol/L) conjugated, 60 mg/dL (1026 μmol/L) unconjugated | ≤ 10% interference up to 1500 mg/dL (17.0 mmol/L) | NA | 299,586 pg/mL | NA |  |  | WHO 1st IS (IL-6, Human rDNA derived), NIBSC code 89/548 | RPBL1348R02_EN_Rev. 02, 2020-07 | IFU Document
Atellica® IM | 10733015 |  |  |  | OUS | Inflamation | No | Atellica IM Lipopolysaccharide Binding Protein (LBP) | Atellica IM LBP | LBP | 50 | 1 x 50 | NA | Atellica IM LBP Quality Control Material Atellica IM Multi‐Diluent 1 | 14 | 28 | Monoclonal mouse anti-LBP antibody | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | up to 12 hours at room temperature, for up to 5 days at 2–8°C, 1 month at ≤ -20°C. Avoid more than 1 freeze/thaw cycle | 20 μL | No |  | one-step sandwich immunoassay using direct chemiluminescent technology | 1.1 – 120.0 μg/mL | NA |  |  |  |  | 2.3 – 18.3 μg/mL | NA | NA | Atellica IM LBP CAL | Included in the kit | Opened at 2–8°C 28 days, At room temperature 4 hours | Liquid | NA | 3 for lot calibration, 2 for pack recalibration | 2 | Recombinant LBP antigen; BSA; preservatives | 28 | 14 |  |  |  | No | Atellica IM Multi-Diluent 1 | 10995639 |  | 1.1 – 240.0 μg/mL |  | Yes | Atellica IM LBP QC | 10733018 | NA | NA | NA | NA | 1.1 μg/mL | 0.7 μg/mL | 1.1 μg/mL | NA | NA | NA | 14.4 μg/mL | NA | 4.4 | 7.1 | 60.4 μg/mL | NA | 6.2 | 8.1 | 103.2 μg/mL | NA | 5.3 | 7.2 | ADVIA Centaur® (X) Atellica IM (y) = 0.90 (x) + 0.70 μg/mL (intercept) r= 0.98 | NA | ≤ 10% interference up to 1000 mg/dL (0.62 mmol/L) | ≤ 10% interference up to 60 mg/dL (1026 μmol/L) Conjugated, and Unconjugated | ≤ 10% interference up to 1500 mg/dL (17.0 mmol/L) | NA | 1093 μg/mL | NA |  |  | traceable to an internal standard. | RPBL1320R03_EN Rev. 03, 2019-10 | IFU Document
Atellica® CH | 11097579 |  | ge19 | 00630414238364 | OUS | General Chemistry | No | Atellica CH Pancreatic Amylase_2 (Urine) | PAMY_2PAMY_2 | 4053 | 1720 | 4 x 430 | N/A | N/A | 9 | 31 | monoclonal mouse antibodies | ;#Urine;# | 10 days at RT; 31 days at 2-8C; Urine amylase is unstable in acidic urine. Adjust urine to a pH ≥ 7 before storage | 16 | No |  | Enzymatic reaction | 20–1500 U/L | 20–1500 U/L | 20–7500 U/L | 20–7500 U/L | Urine: Spontaneously voided urine: Male 7–356 U/L Female 13–319 U/L | Urine: Spontaneously voided urine: Male 7–356 U/L Female 13–319 U/L | N/A | N/A | N/A | Atellica CH SPCL CHEM CAL | 11097579 | 7 | Lyophilized | 5.0 mL of distilled water; Stand for 30 minutes at room temperature; Invert to mix Prior to use mix the contents by gently inverting the vials | 3 | 2 | Human serum | 62 | 31 |  | No |  | Yes | N/A | 11099300 |  | 20–7500 U/L | 20–7500 U/L | No | N/A | N/A | 5 U/L | 5 U/L | N/A | N/A | 5 U/L | 2 U/L | 17 U/L | 2 U/L | 5 U/L | 17 U/L | 48 U/L | 48 U/L | 3.1 | 3.3 | 105 U/L | 105 U/L | 0.9 | 1.1 | 1137 U/L | 1137 U/L | 0.2 | 0.5 | Roche cobas AMY-P (urine) y = 1.13x + 2.3 U/L r=0.994 n=106 | Roche cobas AMY-P (urine) y = 1.13x + 2.3 U/L r=0.994 n=106 | ≤ 10% change up to 500 mg/dL of hemoglobin. | ≤ 10% change up to 30 mg/dL of bilirubin | ≤ 10% change up to 1000 mg/dL of Intralipid |  | N/A | N/A | N/A | N/A | IRMM/IFCC-456 Pancreatic Alpha- Amylase Reference Material and traceable to JCCLS-SOP through the CRM-001c referen | 11111708_EN Rev. 03, 2020-01 | IFU Document | 0.886 | 5.625 | 3.5 | 2.75 | 0.402 | 14.287 | 8.89 | 6.985
Atellica® CH | 10470528 |  |  | 8058056680747 | OUS | General Chemistry | Yes | Glucose-6-Phosphate Dehydrogenase (Sentinel Alliance Application) | G6PDHG6PDH | 4082 | 448 | 1 x 448 | Atellica CH Empty Reagent Packs (11097534) | N/A | 8.5 | 30 | Non Biological G6P-DH; triethanolamine buffer; EDTA; NADP; G6PDH | ;#Whole blood;# | Unprepared: 8 hours at 2-8 °C Prepared sample: 2 hours at 2-8 °C | 4 | Yes |  | Enzymatic / Rate Reaction | 39 to 3700 U/L | 39 to 3700 U/L | N/A | N/A | In erythrocytes: 290 - 412 mU/109 erythrocytes (37 °C) 10.10 - 14.19 U/g Hb (37 °C) | In erythrocytes: 290 - 412 mU/109 erythrocytes (37 °C) 10.10 - 14.19 U/g Hb (37 °C) | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 62 | 30 |  | No |  | Yes | N/A | 11099300 |  | N/A | N/A | No | Sentinel G6P-DH Control Set | 11532787 | 11.9 U/L | 11.9 U/L | N/A | N/A | 11.9 U/L | 6.5 U/L | 35.1 U/L | 6.5 U/L | 11.9 U/L | 35.1 U/L | 218.9 | 218.9 | 7.2 | 8.3 | 990.3 | 990.3 | 1.9 | 2.5 | 3639.7 | 3639.7 | 0.9 | 1.6 | Abbott Alinity c8000: y = 18.23 + 1.097x U/L r-0.996 U/L | Abbott Alinity c8000: y = 18.23 + 1.097x U/L r-0.996 U/L | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | IFU: 17005-4.0/02 https://intranet.sentinel.it/sentineldoc</a> | IFU Document
Atellica® CH | 11532568 |  |  | 00630414287270 | Global | General ChemistryPlease select a value... | No | Atellica CH Lactate_2 (Lac_2) CSF | Lac_2Lac_2 | 8010 | 380 | 4 x 95 | N/A | N/A | 4 | 30 days per pack | PIPES buffer, TOOS, Peroxidase (plant), lactate oxidase and ascorbate oxidase (Microorganism);4-aminoantipyrine | ;#Cerebrospinal fluid (CSF);# | CSF: 24 hr at 2–8°C; 30 days at -20°C | 4 μL | No |  | lactate oxidase to pyruvate and hydrogen peroxide | CSF: 0.9–70 mg/dL | CSF: 0.10–7.77 mmol/L | CSF: 0.9 –140 mg/dL | CSF: 0.10 - 15.54 mmol/L | CSF: 5.4–19.8 mg/dL | CSF: 0.60–2.20 mmol/L | N/A | N/A | N/A | Atellica CH SPCL CHEM CAL | 11099438 | 7 | Lyophilized | 5.0 mL of distilled water; Stand for 30 minutes at room temperature; Invert to mix Prior to use mix the contents by gently inverting the vials | 3 | 2 | Human serum | 62 | 30 |  |  |  | Yes | N/A | 11099300 |  | CSF: 0.9 –140 mg/dL | CSF: 0.10 - 15.54 mmol/L | No | N/A | N/A | 0.6 mg/dL | 0.07 mmol/L | N/A | N/A | 0.6 mg/dL | 0.1 mg/dL | N/A | 0.04 mmol/L | 0.07 mmol/L | N/A | 12.5 mg/dL | 1.39 mmol/L | 1.9 | 4.9 | 48.7 mg/dL | 5.41 mmol/L | 0.9 | 4.8 | 99.0 mg/dL | 10.99 mmol/L | 0.8 | 4.6 | Dimension Vista LA method (CSF) y = 0.96x - 0.6 mg/dL r=0.998 n =124 | Dimension Vista LA method (CSF) y = 0.96x - 0.07 mmol/L r=0.998 n =124 | ≤ 10% change up to 1000 mg/dL of hemoglobin | <10% change up to 9.2 mg/dL of conjugated bilirubin; <10% change up to 7.5 dL mg/dL of unconjugated bilirubin | < 10% change up to 1000 mg/dL of Intralipid | No Biotin in assay formulation | N/A | N/A | N/A | N/A | Internal standard using highly purified material | 11417138_EN 11417138_EN Rev. 02, 2020-09 | IFU Document
Atellica® IM | 10995541 |  |  | 630414600413 | Global | DiabetesDiabetes | No | Atellica IM C-peptide (CpS) | Atellica IM CpSCpS | CpS | 100 | 1 x 100 | Atellica IM CpS CAL SMN 10995542 | Atellica IM Multi-Diluent 10 SMN 10995640 Atellica IM CpS MCM SMN 10995543 | 14 | 28 | Monoclonal mouse anti-C-peptide antibody | ;#Urine;# | Store at 2 - 8 C during the collection process. If not run within 2hours of collection completion aliquot and freeze at -20 C | 50 μL | No |  | 2-site sandwich chemiluminescent with acridinium-ester technology | 0.50 -300 ng/mL | 0.50 -300 ng/mL | All urines must be diluted 1:10 | All urines must be diluted 1:10 | <156.46 µg/24 hours | <156.46 µg/24 hours | NA | NA | NA | Atellica IM CpS CAL | 10995542 | 4 hours at RT | Liquid | Let the calibrators stand for 5 minutes at RT, Gently mix | 3 | 2 | citric acid buffer with casein and perservatives | 50 | 14 |  |  |  | No | Multi-Diluent 10 | 10995640 |  | All urines must be diluted 1:10 | All urines must be diluted 1:10 | No | Use comercially available control | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 18.95 ng/mL | 18.95 ng/mL | 2.4 | 5.1 | 97.28 ng/mL | 97.28 ng/mL | 2 | 9.3 | 248.40 ng/mL | 248.40 ng/mL | 3.7 | 8.5 | ADVIA Centaur CpS; y = 1.05x + 0.10 ng/mL r = 1.00 | ADVIA Centaur CpS; y = 1.05x + 0.10 ng/mL r = 1.00 | NA | NA | NA | NA | NA | NA | NA | NA | WHO International Standard 84/510 | 10997742_EN Rev .01, 2017-10 | IFU Document
Atellica® IM | 10995538 |  |  | 00630414598659 | OUS | N/AMetabolic | No | Atellica IM Cortisol (Cor) | Atellica IM CorCor | Cor | 50 | 1 x 50 | Atellica IM CAL E | Cortisol Urine Reconstitution Buffer (SMN 10995540) | 10 | 10 | Rabbit anti-cortisol antibody&nbsp; bound to mouse monoclonal anti-rabbit IgG antibody | ;#Urine;# | >8hrs at 2–8°C >48hrs freeze at ≤ -20°C for up to one month | 20 μL | No |  | Competitive immunoassay using direct chemiluminescent | 0.50- 53.00 μg/dL | 13.80– 1462.27 nmol/L | NA | NA | Direct Urine : 20.9 - 292.3 μg/24 hours Extracted Urine 9.5 - 136.2 μg/24 hours | Direct Urine: 57.7 - 806.8 nmol/24 hours Extracted urine 26.2 - 375.9 nmol/24 hours | NA | NA | NA | Atellica IM CAL E | 10995512 (2‑pack) 10995513 (6‑pack) | 14 days recon at 2–8°C 4 hours recons at RT | Lyophilized | Add 2.0 mL of special reagent water into each vial. Let vials stand for 15-20 minutes at room temperature. | 3 for lot calibration, 2 for pack recalibration | 2 | defibrinated human plasma | 19 | 10 |  |  |  | Yes | Atellica IM Multi‑Diluent 3 | 10995645 |  | NA | NA | No | Use commercially available controls | NA | NA | NA | NA | NA | Direct Urine ≤ 0.45 μg/dL, Extracted urine ≤ 0.44 μg/dL | Direct Urine ≤ 0.19 μg/dL, Extracted urine ≤ 0.18 μg/dL | Direct Urine ≤ 0.48 μg/dL, Extracted urine ≤ 0.44 μg/dL | Direct urine ≤ 5.24 nmol/L Extracted urine ≤4.97 nmol/L | Direct urine ≤ 12.42 nmol/L Extracted urine ≤12.14 nmol/L | Direct urine ≤ 13.24 nmol/L Extracted urine ≤12.14 nmol/L | 7.05 μg/dL | 194.51 nmol/L | 2.6 | 8.9 | 23.09 μg/dL | 637.05 nmol/L | 5.5 | 9.6 | 35.99 μg/dL | 992.96 nmol/L | 4.2 | 8.9 | ADVIA Centaur Cortisol y = 1.10x + 0.59 μg/dL r = 0.98 | ADVIA Centaur Cortisol y = 1.10x + 16.28 nmol/L r = 0.98 | NA | NA | NA | NA | NA | NA | NA | NA | Internal standards manufactured analytically which are traceable to gas chromatography-mass spectrometry (GCMS) | 10995336_EN Rev. 03, 2020-03 | IFU Document
Atellica® IM | 10995537 |  |  | 00630414600413 | OUS | N/AMetabolic | No | Atellica IM Cortisol (Cor) | Atellica IM CorCor | Cor | 250 | 5 x 50 | Atellica IM CAL E | Cortisol Urine Reconstitution Buffer (SMN 10995540) | 10 | 10 | Rabbit anti-cortisol antibody bound to mouse monoclonal anti-rabbit IgG antibody | ;#Urine;# | >8hrs at 2–8°C >48hrs freeze at ≤ -20°C for up to one month | 20 μL | No |  | Competitive immunoassay using direct chemiluminescent | 0.50- 53.00 μg/dL | 13.80– 1462.27 nmol/L | NA | NA | Direct Urine : 20.9 - 292.3 μg/24 hours Extracted Urine 9.5 - 136.2 μg/24 hours | Direct Urine: 57.7 - 806.8 nmol/24 hours Extracted urine 26.2 - 375.9 nmol/24 hours | NA | NA | NA | Atellica IM CAL E | 10995512 (2‑pack) 10995513 (6‑pack) | 14 days recon at 2–8°C, 4 hours recons at RT | Lyophilized | Add 2.0 mL of special reagent water into each vial. Let vials stand for 15-20 minutes at room temperature. | 3 for lot calibration, 2 for pack recalibration | 2 | defibrinated human plasma | 19 | 10 |  |  |  | Yes | Atellica IM Multi‑Diluent 3 | 10995645 |  | NA | NA | No | Use commercially available controls | NA | NA | NA | NA | NA | Direct Urine ≤ 0.45 μg/dL, Extracted urine ≤ 0.44 μg/dL | Direct Urine ≤ 0.19 μg/dL, Extracted urine ≤ 0.18 μg/dL | Direct Urine ≤ 0.48 μg/dL, Extracted urine ≤ 0.44 μg/dL | Direct urine ≤5.24 nmol/L Extracted urine 4.97 nmol/L | Direct urine ≤ 12.42 nmol/L Extracted urine ≤12.14 nmol/L | Direct urine ≤ 13.24 nmol/L Extracted urine ≤12.14 nmol/L | 7.05 μg/dL | 194.51 nmol/L | 2.6 | 8.9 | 23.09 μg/dL | 637.05 nmol/L | 5.5 | 9.6 | 35.99 μg/dL | 992.96 nmol/L | 4.2 | 8.9 | ADVIA Centaur Cortisol y = 1.10x + 0.59 μg/dL r = 0.98 | ADVIA Centaur Cortisol y = 1.10x + 16.28 nmol/L r = 0.98 | NA | NA | NA | NA | NA | NA | NA | NA | Internal standards manufactured analytically which are traceable to gas chromatography-mass spectrometry (GCMS) | 10995336_EN Rev. 03, 2020-03 | IFU Document
Atellica® IM | 11200450 |  |  |  | OUS | N/AThyroid | No | Atellica IM Thyroglobulin (Tg) | Atellica IM TgTg | Tg | 100 | 1 x 100 | NA | Atellica IM Multi-Diluent 11 SMN 109995642 (2-pk) SMN 10995641 (1 x 10 mL) | 18 | 28 | murine monoclonal anti-thyroglobulinantibody | ;#Serum;#EDTA Plasma;#Lithium Heparin Plasma;# | up to 3 days at RT, up to 7 days at 2-8C and freeze at -20 C for up to 12 months | 100 µL | No |  | 1-step sandwich chemiluminescent immunoassay | 0.050 - 150 ng/mL | 0.076 - 227 pmol/L | 1:10 0.50 - 1500 ng/mL, 1:20 0.50 - 3000 ng/mL, 1:50 0.50 - 7500 ng/mL | 1:10 0.076 - 2270pmol/L, 1 :20 0.076 - 4540 pmol/L, 1:50 0.076 - 11,350 pmol/L | 1.82 - 111 ng/mL | 2.76 - 168 pmol/L | NA | NA | NA | Atellica IM Tg CAL | included in reagent kit | aliquot and freeze -20C for up to 2 months | Lyophilized | Reconstitute with 4.0 mL distilled of deionized water. Mix by gentle inversion or swirling. | 2 | 2 | human thyroglobulin | 21 | 21 |  |  |  | Yes | Atellica IM Multi-Diluent 11 | 10995642/10995641 |  | 1:10 0.50 - 1500 ng/mL, 1:20 0.50 - 3000 ng/mL, 1:50 0.50 - 7500 ng/mL | 1:10 0.076 - 2270pmol/L, 1 :20 0.076 - 4540 pmol/L, 1:50 0.076 - 11,350 pmol/L | No | Use commercially available controls | NA | NA | NA | NA | NA | 0.036 ng/mL | 0.026 ng/mL | 0.050 ng/mL | 0.039 pmol/L | 0.036 pmol/L | 0.076 pmol/L | 0.157 ng/mL | 0.238 pmol/L | 5.1 | 8.3 | 78.9 ng/mL | 120 pmol/L | 1.2 | 2.7 | 136 ng/mL | 206 pmol/L | 1.9 | 2.9 | Commercial Tg method : y = 1.26 x - 0.215 ng/mL r = 0.979 | Commercial Tg method: y + 1.26x - 0.326 pmol/L r = 0.979 | ≤ 10% change up to1000 mg/dL of hemoglobin | ≤ 10% change up to 1900 mg/dL of intralipid | ≤ 10% change up to 60 mg/dL of bilirubin | ≤ 10% change up to 3500 ng/mL of biotin | none up to 80,000 ng/mL | none up to 121, 200 pmol/L | none up to 80,000 ng/mL | none up to 121, 200 pmol/L | CRM 457 | 11200496_EN Rev. 01 2020-11 | IFU Document
Atellica® IM | 11201759 |  |  |  | OUS | N/AThyroid | No | Atellica IM Anti‑Thyroglobulin II (aTgII) | Atellica IM aTg IIaTgII | aTgII | 100 | 1 x 100 | NA | Atellica IM aTgII QC SMN 1120752 Atellica IM aTgII MCM SMN 11201751 | 14 | 28 | acridinium-ester-labeled human thyroglobulin | ;#Serum;#Potassium EDTA Plasma;#Lithium Heparin Plasma;#Sodium Heparin Plasma;# | 8hr at room temperature, 7 days at 2-8C, freeze at -20c up to 6 months | 25 | No |  | 1-step analyte-bridging immunoassay using acridinium ester chemiluminescence | 1.3 - 1000 IU/mL | 1.3 - 1000 IU/mL | Not recommended | Not recommended | Cutoff 4.5 IU/mL | Cutoff 4.5 IU/mL | NA | NA | NA | Atellica IM aTgII CAL | included in reagent kit | 8 hrs | Lyophilized | Add 1.0 mL of reagent water . Let stand 20-30 minutes at RT. Gently mix and invert vials to ensure homogeneity | 2 | 2 | monoclonal mouse anti-human thyroglobulin | 90 | 28 |  |  |  | NA | NA | NA |  | Not recommended | Not recommended | Yes | Atellica IM aTgII QC | 11201752 | NA | NA | NA | NA | 0.9 IU/ML | 0.7 IU/mL | 0.9 IU/mL | 0.9 IU/mL | 0.7 IU/mL | 0.9 IU/mL | 7.56 IU/mL | 7.56 IU/mL | 1.9 | 2.3 | 50.9 IU/mL | 50.9 IU/mL | 1.5 | 2.3 | 915 IU/mL | 915 IU/mL | 1.9 | 2.4 | ADVIA Centaur; y =; 0.94x + 0.01 IU/mL; r = 0.997 | ADVIA Centaur; y = 0.94x + 0.01 IU/mL r = 0.997 | ≤ 10% change up to 1000 mg/dL of hemoglobin | ≤ 10% change up to 60 mg/dL of bilirubin | ≤ 10% change up to 3500 mg/dL of intralipid | ≤ 10% change up to 3500 mg/dL of biotin | none up to 50,000 IU/mL | none up to 50,000 IU/mL | none up to 50,000 IU/mL | none up to 50,000 IU/mL | WHO IRP, anti-Thyroglobulin Serum, Human (NIBSC 65/093 | 11204720_EN Rev. 01, 2020-02 | IFU Document
Atellica® IM | 11201761 |  |  |  | OUS | N/AThyroid | No | Atellica IM Anti‑Thyroglobulin II (aTgII) | Atellica IM aTg IIaTgII | aTgII | 500 | 5 x 100 | NA | Atellica IM aTgII QC SMN 1120752 Atellica IM aTgII MCM SMN 11201751 | 14 | 28 | acridinium-ester-labeled human thyroglobulin | ;#Serum;#Lithium Heparin Plasma;#Sodium Heparin Plasma;#Potassium EDTA Plasma;# | 8hr at room temperature, 7 days at 2-8C, freeze at -20c up to 6 months | 25 | No |  | 1-step analyte-bridging immunoassay using acridinium ester chemiluminescence | 1.3 - 1000 IU/mL | 1.3 - 1000 IU/mL | Not recommended | Not recommended | Cutoff 4.5 IU/mL | Cutoff 4.5 IU/mL | NA | NA | NA | Atellica IM aTgII CAL | included in reagent kit | 8 hrs | Lyophilized | Add 1.0 mL of reagent water . Let stand 20-30 minutes at RT. Gently mix and invert vials to ensure homogeneity | 2 | 2 | monoclonal mouse anti-human thyroglobulin | 90 | 28 |  |  |  | NA | NA | NA |  | Not recommended | Not recommended | Yes | Atellica IM aTgII QC | 11201752 | NA | NA | NA | NA | 0.9 IU/ML | 0.7 IU/mL | 0.9 IU/mL | 0.9 IU/mL | 0.7 IU/mL | 0.9 IU/mL | 7.56 IU/mL | 7.56 IU/mL | 1.9 | 2.3 | 50.9 IU/mL | 50.9 IU/mL | 1.5 | 2.3 | 915 IU/mL | 915 IU/mL | 1.9 | 2.4 | ADVIA Centaur; y =; 0.94x + 0.01 IU/mL; r = 0.997 | ADVIA Centaur; y = 0.94x + 0.01 IU/mL r = 0.997 | ≤ 10% change up to 1000 mg/dL of hemoglobin | ≤ 10% change up to 60 mg/dL of bilirubin | ≤ 10% change up to 3500 mg/dL of intralipid | ≤ 10% change up to 3500 mg/dL of biotin | none up to 50,000 IU/mL | none up to 50,000 IU/mL | none up to 50,000 IU/mL | none up to 50,000 IU/mL | WHO IRP, anti-Thyroglobulin Serum, Human (NIBSC 65/093 | 11204720_EN Rev. 01, 2020-02 | IFU Document
Atellica® IM | 10995585 |  |  |  | Global | N/AThyroid | No | Atellica IM Free Triiodothyronine (FT3) | Atellica IM FT3FT3 | FT3 | 60 | 1 x 60 | Atellica IM CAL A 2- pack SMN 10995500 Atellica IM CAL A 6- pack SMN 10995501 | Atellica IM FT3 MCM SMN 10995586 | 14 | 28 | acridinium-ester-labeled mouse monoclonal anti-T3 antibodies | ;#Serum;#EDTA Plasma;#Lithium Heparin Plasma;# | 8hr at RT; 48 hr at 2-8C; Freeze samples at or below -20°C if the sample is not assayed within 48 hr | 35 | No |  | competitive immunoassay using direct chemiluminescent technology | 0.20 - 20.00 pg/mL | 0.31 - 30.80 pmol/L | NA | NA | 2.3 - 4.2 pg/mL | 3.5 - 6.5 pmol/L | NA | NA | NA | Atellica IM CAL A | 10995500 (2-pack) 10995501 (6-pack) | 28 days recon at 2–8°C, 4 hours recon at RT | Lyophilized | Add 5.0 mL of reagent water, let the calibrators stand for 15–20 minutes at RT. Mix by gentle inversion until homogenous | 3 for lot calibration. 2 for pack calibration | 2 | human plasma with sodium azide,preservatives and protein stabilizers | 82 | 28 |  |  |  | No | NA | NA |  | NA | NA | No | Use commecially available controls | NA | 0.06 pg/mL | 0.09 pmol/L | NA | NA | 0.02 pg/mL | 0.00 pg/mL | NA | 0.03 pmol/L | 0.00 pmol/L | NA | 2.87 pg/mL | 4.42 pmol/L | 0.94 | 1.78 | 4.96 pg/mL | 7.64 pmol/L | 0.67 | 1.07 | 18.13 pg/mL | 27.92 pmol/L | 1.96 | 3.05 | ADVIA Centaur FT3 y = 0.99 x + 0.007 pg/mL r = 0.99 | ADVIA Centaur FT3 y = 0.99x + 0.011 pmol/L r = 0.99 | <10% change up to 500 mg/dL of hemoglobin | <10% change up to 20 mg/dL of bilirubin | <10% change up to 1000 mg/dL Intralipid | No Biotin in assay formulation | NA | NA | N/A | N/A | Internal standard manufactured using U.S.P.(United States Pharmacopeia) material | 10995347_EN Rev.03, 2020-06 | IFU Document
Atellica® IM | 10995584 |  |  |  | Global | N/AThyroid | No | Atellica IM Free Triiodothyronine (FT3) | Atellica IM FT3FT3 | FT3 | 300 | 5 x 60 | Atellica IM CAL A 2- pack SMN 10995500 Atellica IM CAL A 6- pack SMN 10995501 | Atellica IM FT3 MCM SMN 10995586 | 14 | 28 | acridinium-ester-labeled mouse monoclonal anti-T3 antibodies | ;#Serum;#EDTA Plasma;#Lithium Heparin Plasma;# | 8hr at RT; 48 hr at 2-8C; Freeze samples at or below -20°C if the sample is not assayed within 48 hr | 35 | No |  | competitive immunoassay using direct chemiluminescent technology | 0.20 - 20.00 pg/mL | 0.31 - 30.80 pmol/L | NA | NA | 2.3 - 4.2 pg/mL | 3.5 - 6.5 pmol/L | NA | NA | NA | Atellica IM CAL A | 10995500 (2-pack) 10995501 (6-pack) | 28 days recon at 2–8°C, 4 hours recon at RT | Lyophilized | Add 5.0 mL of reagent water, let the calibrators stand for 15–20 minutes at RT. Mix by gentle inversion until homogenous | 3 for lot calibration. 2 for pack calibration | 2 | human plasma with sodium azide,preservatives and protein stabilizers | 82 | 28 |  |  |  | No | NA | NA |  | 0.02 pg/mL | 0.03 pmol/L | No | Use commecially available controls | NA | 0.06 pg/mL | 0.09 pmol/L | NA | NA | 0.02 pg/mL | 0.00 pg/mL | NA | 0.03 pmol/L | 0.00 pmol/L | NA | 2.87 pg/mL | 4.42 pmol/L | 0.94 | 1.78 | 4.96 pg/mL | 7.64 pmol/L | 0.67 | 1.07 | 18.13 pg/mL | 27.92 pmol/L | 1.96 | 3.05 | ADVIA Centaur FT3 y = 0.99 x + 0.007 pg/mL r = 0.99 | ADVIA Centaur FT3 y = 0.99x + 0.011 pmol/L r = 0.99 | <10% change up to 500 mg/dL of hemoglobin | <10% change up to 20 mg/dL of bilirubin | <10% change up to 1000 mg/dL Intralipid | No Biotin in assay formulation | NA | NA | N/A | N/A | Internal standard manufactured using U.S.P.(United States Pharmacopeia) material | 10995347_EN Rev.03, 2020-06 | IFU Document
Atellica® IM | 10995703 |  |  | 00630414600161 | OUS | N/AThyroid | No | Atellica IM Thyroid Stimulating Hormone 3-Ultra (TSH3‑UL) | Atellica IM TSH3‑ULTSH3UL | TSH3UL | 110 | 1 x 110 | NA | Atellica IM TSH3‑UL MCM SMN10995705 Multi-Diluent 15 SMN 11200798 | 14 | 63 | anti-FITC monoclonal antibody | ;#Serum;#Lithium Heparin Plasma;#EDTA Plasma;# | RT 24 hours, 2–8°C 48 hours, serum and EDTA ≤ -20°C; up to 30 days, Lithium heparin ≤ -20°C; up to 14 days; Freeze samples only 1 time | 75 μL | No |  | 3rd generation immunoassay using direct chemiluminescent technology | 0.008–150.000 μIU/mL | 0.008–150.000 mIU/L | NA | NA | 0.55 - 4.78 μIU/mL | 0.55 - 4.78 mIU/L | NA | NA | NA | Atellica IM TSH3-UL CAL | Included in reagent kit | 28 days recon at 2–8°C,&nbsp; 4 hours recon at RT | Lyophilized | Add 2.0 mL of reagent water, let the calibrators stand for 15–20 minutes at RT. Gently swirl and invert the vials until homogeneous | 3 for lot calibration, 2 for pack recalibration | 2 | Equine serum | 49 | 63 |  |  |  | Yes | Multi-Diluent 15 | 11200798 |  | 1:2 0.008 – 300.000 μIU/mL, 1:5 0.008 - 750.000 μIU/mL | 1:2 0.008 – 300.000 mIU/L 1;5 0.008 - 750.000 mIU/L | No | Use commercially available Controls | NA | NP | NP | NP | NP | 0.007 μIU/mL | 0.003 μIU/mL | 0.007 μIU/mL | 0.003 mIU/L | 0.007 mIU/L | 0.007 mIU/L | 0.157 μIU/mL | 0.157 mIU/L | 2.1 | 3.7 | 8.995 μIU/mL | 8.995 mIU/L | 1.5 | 3.4 | 118.735 μIU/mL | 118.735 mIU/L | 2.1 | 3.3 | ADVIA Centaur TSH3‑UL y = 1.05x + 0.001 μIU/mL r = 1.000 | ADVIA Centaur TSH3‑UL y = 1.05x + 0.001 mIU/L r = 1.000 | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 40 mg/dL of conjugated bilirubin; ≤ 10% change up to 40 mg/dL of unconjugated bilirubin | ≤ 10% change up to 1000 mg/dL of Intralipid | No Biotin in assay formulation | 3000 μIU/mL | 3000 mIU/L | 3000 mIU/L | 3000 mIU/L | WHO 3rd International Standard for human TSH (IRP 81/565) | 10995434_EN Rev. 03, 2019-08(OUS) | IFU Document
Atellica® IM | 10995704 |  |  | 00630414600178 | OUS | N/AThyroid | No | Atellica IM Thyroid Stimulating Hormone 3-Ultra (TSH3‑UL) | Atellica IM TSH3‑ULTSH3UL | TSH3UL | 550 | &nbsp;5 x 110 | NA | Atellica IM TSH3‑UL MCM SMN10995705 Multi-Diluent 15 SMN 11200798 | 14 | 63 | anti-FITC monoclonal antibody | ;#Serum;#Heparinized Plasma;#EDTA Plasma;# | RT 24 hours, 2–8°C 48 hours, serum and EDTA ≤ -20°C; up to 30 days, Lithium heparin ≤ -20°C; up to 14 days; Freeze samples only 1 time | 75 μL | No |  | 3rd generation immunoassay using direct chemiluminescent technology | 0.008–150.000 μIU/mL | 0.008–150.000 mIU/L | NA | NA | 0.55 - 4.78 μIU/mL | 0.55 - 4.78 mIU/L | NA | NA | NA | Atellica IM TSH3-UL CAL | NA-included in reagent kit | 28 days recon at 2–8°C,&nbsp;&nbsp; 4 hours recon at RT | Lyophilized | Add 2.0 mL of reagent water, let the calibrators stand for 15–20 minutes at RT. Gently swirl and invert the vials until homogeneous | 3 for lot calibration, 2 for pack recalibration | 2 | Equine Serum | 49 | 63 |  |  |  | Yes | Multi-Diluent 15 | 11200798 |  | 1:2 0.008 – 300.000 μIU/mL, 1:5 0.008 - 750.000 μIU/mL | 1:2 0.008 – 300.000 mIU/L 1;5 0.008 - 750.000 mIU/L | No | Use commercially available Controls | NA | NP | NP | NP | NP | 0.007 μIU/mL | 0.003 μIU/mL | 0.007 μIU/mL | 0.003 mIU/L | 0.007 mIU/L | 0.007 mIU/L | 0.157 μIU/mL | 0.157 mIU/L | 2.1 | 3.7 | 8.995 μIU/mL | 8.995 mIU/L | 1.5 | 3.4 | 118.735 μIU/mL | 118.735 mIU/L | 2.1 | 3.3 | ADVIA Centaur TSH3‑UL y = 1.05x + 0.001 μIU/mL r = 1.000 | ADVIA Centaur TSH3‑UL y = 1.05x + 0.001 mIU/L r = 1.000 | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 40 mg/dL of conjugated bilirubin, ≤ 10% change up to 40 mg/dL of unconjugated bilirubin | ≤ 10% change up to 1000 mg/dL of Intralipid | No Biotin in assay formulation | 3000 μIU/mL | 3000 mIU/L | 3000 mIU/L | 3000 mIU/L | WHO 3rd International Standard for human TSH (IRP 81/565) | 10995434_EN Rev. 03, 2019-08(OUS) | IFU Document

## CH 완성 (hidden)
Atellica® CH
Test Name | AMR | CRR | Dilution | Reference Range | Calibrator Name | Calibrator Open Vial Stability | Levels used in Calibration curve | Lot Calibration Interval (days) | Pack Calibration Interval (days) | Automatic Dilution | Sample Diluent | Dilution Factor2 | Number of Tests | Kit Configuration | TTFR (min) | OBS | Assay Principle | Hemolysis | Icterus | Lipemia | Biotin | Hook Effect (Common Units) | High-Dose Hook Effect (Common Units) | Standardization Reference Method/Material
Alb | 1.0 - 6 g/dL | 12.0 g/dL | 2 | 3.2–4.8 g/dL | Atellica CH CHEM CAL | 2 days | 2 | 97 | 60 | Yes | N/A |  | 6800 | 4 x 1700 | 10 | 60 | Dye Binding: bromocresol green | ≤ 10% at 400 mg/dL (0.25 mmol/L) of hemoglobin | ≤ 10% at 30 mg/dL (513 µmol/L) of bilirubin, conjugated ≤ 10% at 30 mg/dL (513 µmol/L) of bilirubin, unconjugated | ≤ 10% at 1000 mg/dL (11.3 mmol/L) of Intralipid | No Biotin in assay formulation | N/A | N/A | BCG reference method NIST SRM 927reference materials
ALP_2c | 10 - 1000 U/L | 10 - 2300 U/L | 2.3 | 46–116 U/L | Atellica CH ALP_2 CAL | 30 days | 2 | 60 | 17 | Yes | N/A |  | 4800 | 4 x 1200 | 8 | 30 | IFCC Standardization | ≤ 10% at 1000 mg/dL of hemoglobin | ≤ 10% at 80 mg/dL of bilirubin | ≤ 10% at 600 mg/dL of Intralipid | No Biotin in assay formulation | N/A | N/A | traceable to IFCC primary reference procedure for the measurement of catalytic activity of alkaline phosphatase at 37°C
ALT | 7 - 1100 U/L | 7 - 3300 U/L | 3 | 10–49 U/L | Atellica CH ENZ 2 CAL | 30 days at 2-8 °C | 2 | 131 | 30 | Yes | N/A |  | 2550 | 3 x 850 | 10 | 90 | Modified IFCC | ≤ 10% at 450 mg/dL (0.28 mmol/L) of hemoglobin | ≤ 10% at 30 mg/dL (513 µmol/L) of bilirubin, conjugated ≤ 10% at 30 mg/dL (513 µmol/L) of bilirubin, unconjugated | ≤ 10% at 450 mg/dL (5.09 mmol/L) of Intralipid | No Biotin in assay formulation | N/A | N/A | IFCC reference method, IFCC-454
AMY_2 | 20 – 1500 U/L | 20 – 4500 U/L | 3 | 30–118 U/L | Atellica CH SPCL CHEM CAL | 7 days | 2 | 62 | 31 | Yes | N/A |  | 1050 | 3 x 350 | 9 | 31 | Ethylidene Blocked-pNPG7 | < 10% change up to 500 mg/dL of hemoglobin. | < 10% change up to 30 mg/dL of bilirubin. | < 10% change up to 650 mg/dL of Intralipid® | No Biotin in assay formulation | N/A | N/A | Traceable to the IRMM/IFCC-456 reference material and commutable to the IFCC Alpha-Amylase Primary Reference Procedure
APO A1 | 15 - 200 mg/dL | 400 mg/dL | 2 | Males: 79–169 mg/dL Females: 76–214 mg/dL | Atellica CH (APO A1 & B CAL) | 28 days | 5 | 60 | 14 | Yes | Atellica CH Diluent |  | 300 | 2 x 150 | 10 | 14 | PEG-enhanced immunoturbidimetric | ≤ 10% change up to 525 mg/dL of hemoglobin. | ≤ 10% change up to 30 mg/dL of bilirubin | ≤ 10% change up to 650 mg/dL of Intralipid. | No Biotin in assay formulation | 2,205 mg/dL | 2,205 mg/dL | WHO-IFCC reference material SP1-01
APO B | 15 - 200 mg/dL | 400 mg/dL | 2 | Males: 46–174 mg/dL Females:46–142 mg/dL | Atellica CH (APO A1 & B CAL) | 28 days | 5 | 60 | 30 | Yes | Atellica CH Diluent |  | 300 | 2 x 150 | 10 | 30 | PEG-enhanced immunoturbidimetric | ≤ 10% change up to 525 mg/dL of hemoglobin. | ≤ 10% change up to 30 mg/dL of bilirubin | ≤ 10% change up to 650 mg/dL of Intralipid. | No Biotin in assay formulation | 5600 mg/dL | 5600 mg/dL | WHO-IFCC reference material SP3-07
ASO_2 | 25.0 - 1000 IU/mL | 2000 IU/mL | 2 | Children (preschool) ≤ 100.0 IU/mL Children (school age)≤ 250.0 IU/mL Adults≤ 200.0 IU/mL Siemens-determined adult reference range ≤ 194.0 IU/mL | Atellica CH (LSP CAL) | 28 days | 2 | 60 | 30 | Yes | Atellica CH Diluent |  | 1200 | 3 x 400 | 10 | 30 | latex enhanced immunoturbidimetric | ≤ 10% change up to 1000 mg/dL of hemoglobin. | ≤ 10% change up to 60 mg/dL of bilirubin | ≤ 10% change up to 1000 mg/dL of Intralipid. | No Biotin in assay formulation | 8500 IU/mL. | 8500 IU/mL. | WHO 1st international reference material for ASO
AST | 8 - 1000 U/L | 6000 U/L | 6 | <34 U/L | Atellica CH ENZ 2 CAL | 30 days | 2 | 131 | 30 | Yes | N/A |  | 2550 | 3 x 850 | 10 | 90 | Modified IFCC | Avoid hemolyzed samples | ≤ 10% at 30 mg/dL (513 µmol/L) of bilirubin, conjugated ≤ 10% at 30 mg/dL (513 µmol/L) of bilirubin, unconjugated | ≤ 10% at 400 mg/dL (4.52 mmol/L) of Intralipid | No Biotin in assay formulation | N/A | N/A | IFCC reference method
C3 | 0.5 - 500.0 mg/dL | 0.5 - 1125.0 mg/dL | 2.25 | Newborns 58 –108 mg/dL 3 months 67 – 124 mg/dL 6 months 74 – 138 mg/dL 9 months 78 – 144 mg/dL 12 months 80 – 150 mg/dL 2–10 years 80 – 150 mg/dL 12–18 years 85 – 160 mg/dL 20 years 82–160 mg/dL 30 years; 84–160 mg/dL 40–70 years 90–170 mg/dL | Atellica CH (LSP CAL) | 28 days | 6 | 60 | 14 | Yes | Atellica CH Diluent |  | 400 | 2 x 200 | 10 | 30 | PEG-enhanced immunoturbidimetric | ≤ 10% at 1000 mg/dL of hemoglobin | ≤ 10% at 25 mg/dL of bilirubin, conjugated ≤ 10% at 18.8 mg/dL of bilirubin, unconjugated | ≤ 10% at 1000 mg/dL of Lipemia (avian triglycerides) | No Biotin in assay formulation | 2100 mg/dL | 2100 mg/dL | IRMM reference material CRM 470 from IFCC
C4 | 0.4 - 140.0 mg/dL | 0.4 - 420.0 mg/dL | 3 | 12.0–36.0 mg/dL | Atellica CH (LSP CAL) | 28 days | 6 | 60 | 14 | Yes | Atellica CH Diluent |  | 400 | 2 x 200 | 10 | 30 | PEG-enhanced immunoturbidimetric | ≤ 10% at 750 mg/dL of hemoglobin | ≤ 10% at 18.8 mg/dL of bilirubin, conjugated ≤ 10% at 18.8 mg/dL of bilirubin, unconjugated | ≤ 10% at 1000 mg/dL of Lipemia (avian triglycerides) | No Biotin in assay formulation | 300 mg/dL | 300 mg/dL | e IRMM reference material CRM 470 from IFCC
Ca (Serum/plasma) | 1 - 15 mg/dL | 30.0 mg/dL | 2 | 8.3–10.6 mg/dL | Atellica CH CHEM CAL | 2 days | 2 | 33 | 3 | Yes | N/A |  | 3560 | 4 x 890 | 8 | 30 | CPC | ≤ 10% at 1000 mg/dL (0.62 mmol/L) of hemoglobin | ≤ 10% at 30 mg/dL (513 µmol/L) of bilirubin, conjugated ≤ 10% at 30 mg/dL (513 µmol/L) of bilirubin, unconjugated | ≤ 10% at 650 mg/dL (7.35 mmol/L) of Intralipid | No Biotin in assay formulation | N/A | N/A | NIST atomic absorption reference method, NIST SRM 915 and SRM 909b reference materials
Chol_2 | 25 – 618 mg/dL | 1236 mg/dL | 2 | Low-risk levels (desirable); 200 mg/dL Moderate-risk levels (borderline) 200–239 mg/dL High-risk levels ≥ 240 mg/dL | Atellica CH CHEM CAL | 2 days | 2 | 50 | 7 | Yes | N/A |  | 8400 | 4 x 2100 | 10 | 26 | Colorimetric: cholesterol esterase and cholesterol oxidase conversion followed by a Trinder endpoint | ≤ 10% at 500 mg/dL (0.311 mmol/L) of hemoglobin | ≤ 10% at 60 mg/dL (1026 µmol/L) of bilirubin, conjugated ≤ 10% at 20 mg/dL (342 µmol/L) of bilirubin, unconjugated | ≤ 10% at 1000 mg/dL (11.3 mmol/L) of Intralipid | No Biotin in assay formulation | N/A | N/A | NCEP/CDC reference method; which uses NIST SRM 909 reference materials
Cl (serum/plasma) | 50 - 200 mmol/L | N/A | - | 98–107 mmol/L | IMT Reaget | 30 days onboard | 2 | automatic calibration every 4 hours | automatic calibration every 4 hours | No | N/A |  | 20000 | 4 x 5000 | 18 sec | 14 days or 5000 samples per multisensor | Indirect Integrated Multisensor Technology (IMT) | ≤ 10% at 750 mg/dL (0.47 mmol/L)) of hemoglobin | ≤ 10% at 60 mg/dL of bilirubin, conjugated ≤ 10% at 60 mg/dL of bilirubin, unconjugated | ≤ 10% at 3000 mg/dL of Intralipid | No Biotin in assay formulation | N/A | N/A | coulometric reference method, NIST Reference Serum and materials
CO2_c | 10.0 - 40.0 mEq/L | N/A | - | 20.0–31.0 mEq/L | Atellica CH CO2 CAL | 30 days | 2 | 30 | 6 | No | N/A |  | 7600 | 4 x 1900 | 2 | 27 | Enzymatic | ≤ 10% at 600 mg/dL of hemoglobin | ≤ 10% at 30 mg/dL of bilirubin | ≤ 10% at 750 mg/dL of Intralipid | No Biotin in assay formulation | N/A | N/A | Corning 965 method / NIST SRM 192
Crea_2 (serum/plasma) | 0.15 - 30.00 mg/dL | 60 mg/dL | 2 | Males 0.70–1.30 mg/dL Females 0.55–1.02 mg/dL | Atellica CH CHEM CAL | 2 days | 2 | 180 | 6 | Yes | N/A |  | 5888 | 4 x 1472 | 7 | 17 | Jaffe, alkaline picrate, kinetic with blank rate correction | Do not use hemolyzed samples ≤ 10% at 500 mg/dL (0.311 mmol/L) of hemoglobin | ≤ 10% at 20 mg/dL (342 µmol/L) of bilirubin, conjugated ≤ 10% at 10 mg/dL (171 µmol/L) of bilirubin, unconjugated | ≤ 10% at 500 mg/dL (5.65 mmol/L) of Intralipid | No Biotin in assay formulation | N/A | N/A | IDMS reference method via correlation ofpatient samples and NIST SRM967 reference material
Crea_2 (urine) | 3.00 -245.00 mg/dL | 735 mg/dL | 3 | Males 950–2490 mg/day Female 600–1800 mg/day | Atellica CH CHEM CAL | 2 days | 2 | 180 | 6 | Yes | N/A |  | 5888 | 4 x 1472 | 7 | 17 | Jaffe, alkaline picrate, kinetic with blank rate correction | N/A | N/A | N/A | No Biotin in assay formulation | N/A | N/A | IDMS reference method via correlation ofpatient samples and NIST SRM967 reference material
CRP_2 | 0.4 - 30.4 mg/dL | 91.2 mg/dL | 3 | < 1.0 mg/dL | Atellica CH (CRP_2 CAL) | 60 days | 6 | 60 | 30 | Yes | Atellica CH Diluent |  | 1000 | 2 x 500 | 8 | 30 | Latex-enhanced immunoturbidimetric assay | ≤ 10% at 1000 mg/dL of hemoglobin | ≤ 10% at 60 mg/dL of bilirubin, conjugated ≤ 10% at 60 mg/dL of bilirubin, unconjugated | ≤ 10% at 1000 mg/dL of Intralipid | No Biotin in assay formulation | 100.0 mg/dL | 100.0 mg/dL | IRMM reference material ERMDA-470.
CYSC_2 | 0.25–8.93 mg/L | 0.25–26.79 mg/L | 3 | 0.64–1.23 mg/L | Atellica® CH (CYSC_2 CAL) | 60 | 6 | 61 | 40 | Yes | Atellica CH Diluent |  | 800 | 4 x 200 | 10 | 40 | Latex-enhanced immunoturbidimetric | ≤ 10% change up to 1000 mg/dL of hemoglobin. | ≤ 10% change up to 60 mg/dL of conjugated bilirubin and 60 mg/dL of unconjugated bilirubin. | ≤ 10% change up to 1000 mg/dL of Intralipid. | No Biotin in assay formulation | 56.00 mg/L | 56.00 mg/L | IFCC Standard ERM–DA471
DBil_2 | 0.1–15.0 mg/dL | 0.1 - 22.5 mg/dL | 1.5 | ≤ 0.3 mg/dL | Atellica CH CHEM CAL | 8 hours | 2 | 60 | 30 | Yes | N/A |  | 1792 | 4 x 448 | 10 | 30 | Vanadate oxidation | ≤ 10% change up to 1000 mg/dL of hemoglobin. | N/A | ≤ 10% change up to 750 mg/dL of Lipemia (Triglycerides Concentrate) | No Biotin in assay formulation | N/A | N/A | AACC reference method NIST SRM 916 reference materials
D-HDL | 20.0 – 129.0 mg/dL | 258.0 mg/dL | 2 | Low (undesirable, high risk): 40.0 mg/dL High (desirable, low risk): ≥ 60.0 mg/dL | Atellica CH HDL/LDL CAL | 3 days | 2 | 60 | 30 | Yes | N/A |  | 1792 | 4 x 448 | 8 | 89 | Elimination/catalase | ≤ 10% at 500 mg/dL of hemoglobin | ≤ 10% at 30 mg/dL of bilirubin | ≤ 10% at 1000 mg/dL of Intralipid | No Biotin in assay formulation | N/A | N/A | NCEP Designated Comparison Method via patient sample correlation.
DLDL | 5.0 – 1000.0 mg/dL | N/A | - | Optimal; <100.0 mg/dL Near optimal: 100.0–129.0 mg/dL Borderline high: 130.0–159.0 mg/dL High: 160.0–189.0 mg/dL Very high: ≥ 190.0 mg/dL | Atellica CH HDL/LDL CAL | 3 days | 2 | 121 | 15 | No | N/A |  | 1600 | 4 x 400 | 10 | 15 | Elimination/catalase | ≤ 10% at 810 mg/dL of hemoglobin | ≤ 10% at 20 mg/dL of conjugated bilirubin ≤ 10% at 30 mg/dL of unconjugated bilirubin | ≤ 10% at 505 mg/dL of Intralipid | No Biotin in assay formulation | N/A | N/A | CDC Reference Method via patient sample correlation
GGT | 7 - 1200 U/L | 7 - 3600 U/L | 3 | males < 73 U/L females < 38 U/L | Atellica CH ENZ 1 CAL | 30 days | 2 | 60 | 22 | Yes | N/A |  | 1792 | 4 x 448 | 10 | 22 | Modified IFCC | ≤ 10% at 1000 mg/dL of hemoglobin | ≤ 10% at 10 mg/dL of bilirubin | ≤ 10% at 500 mg/dL of Intralipid | No Biotin in assay formulation | N/A | N/A | IFCC-452 reference material
GluH_3 (serum/ plasma) | 4 -700 mg/dL | 2100 mg/dL | 3 | Adults; 74–106 mg/dL Newborn; 40–60 mg/dL Newborn; 1 day 50–80 mg/dL Child 60–100 mg/dL | Atellica CH CHEM CAL | 2 days | 2 | 182 | 30 | Yes | N/A |  | 6240 | 4 x 1560 | 7 | 62 | Hexokinase | ≤ 10% change up to 1000 mg/dL of hemoglobin. | ≤ 10% change up to 30 mg/dL of bilirubin | ≤ 10% change up to 1000 mg/dL of Intralipid. | No Biotin in assay formulation | N/A | N/A | NIST SRM 965a
IgA_2 | 33.0 - 540.0 mg/dL | 15.0-5,400.0 mg/dL | 10 | 40.0-350.0 mg/dL | Atellica CH LSP CAL | 28 days | 6 | 60 | 30 | Yes | NA |  | 600 | 4 x 150 | 10 | 30 | PEG-enhanced immunoturbidimetric | <10% up to 1000 mg/dL of hemoglobin | <10% up to 50 mg/dL of bilirubin | <10% up to 922 mg/dL of Trig; <10% up to 500 mg/dL of Intralipid. | No Biotin in assay formulation | 9000.0 mg/dL | 9,000.0 mg/dL | IRMM reference material CRM-470
IgG_2 | 140 - 3400 mg/dL | 70-34,000 mg/dL | 10 | 650-1,600 mg/dL | Atellica CH LSP CAL | 28 days | 6 | 60 | 30 | Yes | NA |  | 720 | 4 x 180 | 10 | 30 | PEG-enhanced immunoturbidimetric | <10% up to 1000 mg/dL of hemoglobin | <10% up to 50 mg/dL of bilirubin | <10% up to 1000 mg/dL of Intralipid. | No Biotin in assay formulation | 70,000 mg/dL | 70,000 mg/dL | IRMM reference material CRM-470
IgM_2 | 21.0–330.0 mg/dL | 8.0-3,300.0 mg/dL | 10 | 50.0-300.0 mg/dL | Atellica CH LSP CAL | 28 days | 6 | 60 | 30 | Yes | NA |  | 720 | 4 x 180 | 10 | 30 | PEG-enhanced immunoturbidimetric | <10% up to 1000 mg/dL of hemoglobin | <10% up to 60 mg/dL of bilirubin | <10% up to 1000 mg/dL of Intralipid. | No Biotin in assay formulation | 7,000.0 mg/dL | 7,000.0 mg/dL | IRMM reference material ERMDA-470
IP (serum/plasma) | 0.3 -20 mg/dL | 0.3 - 40.0 mg/dL | 2 | 2.4-5.1 mg/dL | Atellica CH CHEM CAL | 2 days | 2 | 180 | 7 | Yes | N/A |  | 5100 | 3 x 1700 | 10 | 30 | Phosphomolybdate | ≤ 10% at 500 mg/dL of hemoglobin | ≤ 10% at 30 mg/dL of bilirubin | ≤ 10% at 650 mg/dL of Intralipid | No Biotin in assay formulation | N/A | N/A | NIST SRM 200
Iron_2 | 2 - 1000 µg/dL | 2000 μg/dL | 2 | Male 65-175 μg/dL Female 50-170 μg/dL | Atellica CH CHEM CAL | 2 days | 2 | 180 | 30 | Yes | N/A |  | 1792 | 4 x 448 | 10 | 30 | Ferrozine | N/A | <10% up to 50 mg/dL of bilirubin | <10% up to 500 mg/dL of Intralipid. | No Biotin in assay formulation | N/A | N/A | NIST SRM 937
K (serum/plasma) | 1 - 10 mmol/L | N/A | - | Serum-3.5–5.1 mmol/L Plasma-3.4–4.5 mmol/L | IMT Reaget | 30 days onboard | 2 | automatic calibration every 4 hours | automatic calibration every 4 hours | No | N/A |  | 20000 | 4 x 5000 | 18 sec | 14 days or 5000 samples per multisensor | Indirect Integrated Multisensor Technology (IMT) | N/A | <10% up to 60 mg/dL of bilirubin, conjugated and unconjugated | <10% up to 3000 mg/dL of Intralipid® | No Biotin in assay formulation | N/A | N/A | flame emission spectrophotometry reference method
LDLP | 14-750 U/L | 14 - 4500 U/L | 6 | 120 - 246 U/L | Atellica CH ENZ 1 CAL | 30 days | 2 | 60 | 28 | Yes | N/A |  | 1792 | 4 x 448 | 9 | 28 | Lactate / NAD | N/A | ≤ 10% at 30 mg/dL of bilirubin | ≤ 10% at 650 mg/dL of Intralipid | No Biotin in assay formulation | N/A | N/A | IFCC-453 reference material
Lip | 8 - 700 U/L | 9- 3500 U/L | 5 | 12 - 53 U/L | Atellica CH ENZ 1 CAL | 30 days | 2 | 61 | 9 | Yes | N/A |  | 1280 | 4 x 320 | 8 | 9 | Colorimetric rate | <10% up to 200 mg/dL of hemoglobin | <10% up to 25 mg/dL of bilirubin | <10% up to 1000 mg/dL of Intralipid. | No Biotin in assay formulation | N/A | N/A | Internal standard
Mg | 0.50 - 5.00 mg/dL | 10.00 mg/dL | 2 | 1.60 - 2.60 mg/dL | Atellica CH CHEM CAL | 2 days | 2 | 180 | 3 | Yes | Atellica CH Diluent |  | 1200 | 3 x 400 | 8 | 14 | Xylidyl Blue | <10% up to 750 mg/dL of hemoglobin | <10% up to 30 mg/dL of bilirubin | <10% up to 500 mg/dL of Intralipid. | No Biotin in assay formulation | N/A | N/A | Atomic absorption reference method/ NIST materials
Mg | Urine 1.00 - 14.00 mg/dL | 28.00 mg/dL | 2 | 24 - 255 mg/24 hours | Atellica CH CHEM CAL | 2 days | 2 | 180 | 3 | Yes | Atellica CH Diluent |  | 1200 | 3 x 400 | 8 | 14 | Xylidyl Blue | N/A | N/A | N/A | No Biotin in assay formulation | N/A | N/A | Atomic absorption reference method/ NIST materials
Na (serum/plasma) | 50 - 200 mmol/L | N/A | - | 50–200 mmol/L | IMT Reaget | 30 days onboard | 2 | automatic calibration every 4 hours | automatic calibration every 4 hours | No | N/A |  | 20000 | 4 x 5000 | 18 sec | 14 days or 5000 samples per multisensor | Indirect Integrated Multisensor Technology (IMT) | <10% up to 750 mg/dL of hemoglobin | <10% up to 60 mg/dL of bilirubin, conjugated and unconjugated | <10% up to 3000 mg/dL of Intralipid® | No Biotin in assay formulation | N/A | N/A | flame emission spectrophotometry reference method
RF | 3.5 - 90.0 IU/mL | 900 IU/mL | 10 | < 14 IU/mL | Atellica CH LSP CAL | 28 days | 6 | 60 | 14 | Yes | NA |  | 360 | 2 x 180 | 10 | 21 | Latex-enhanced immunoturbidimetric assay | ≤ 10% change up to 1000 mg/dL of hemoglobin. | ≤ 10% change up to 25 mg/dL of bilirubin | ≤ 10% change up to 1000 mg/dL of Intralipid. | No Biotin in assay formulation | 450 IU/mL | 450 IU/mL | Internal standard
TBil_2 | 0.15 - 35.0 mg/dL | 70.0 mg/dL | 2 | 0–1 day: <8.0 mg/dL 1–2 days: <12.0 mg/dL 3–5 days: <16.0 mg/dL > 5 days–60 years: 0.3–1.2 mg/dL 60–90 years: 0.2–1.1 mg/dL >90 years: 0.2–0.9 mg/dL | Atellica CH CHEM CAL | 8 hours | 2 | 60 | 30 | Yes | N/A |  | 1792 | 4 x 448 | 10 | 30 | Vanadate oxidation | <10% up to 1000 mg/dL of hemoglobin | N/A | ≤10% up to 750 mg/dL of Lipemia (Triglycerides concentrate) | No Biotin in assay formulation | N/A | N/A | AACC reference method which uses reference materials from the National Institute of Standards and Technology (NIST SRM 916
TIBC | 40 – 670 μg/dL | N/A | 0 | 250–425 μg/dL . | Atellica CH SPCL CHEM CAL | 7 days | 2 | 180 | 7 | No | N/A |  | 800 | 4 x 200 | 10 | 7 | Sequential release and uptake of iron | <10% up to 1000 mg/dL of hemoglobin | <10% up to 25 mg/dL of bilirubin | <10% up to 1000 mg/dL of Trig; <10% up to 375 mg/dL of Intralipid. | No Biotin in assay formulation | N/A | N/A | internal standard.
TP | 2.0 - 12.0 g/dL | 24.0 g/dL | 2 | 5.7-8.2 g/dL | Atellica CH CHEM CAL | 2 days | 2 | 181 | 30 | Yes | N/A |  | 7400 | 4 x 1850 | 10 | 90 | biuret reagent (cupric sulfate in an alkaline solution) | <10% up to 500 mg/dL of hemoglobin | <10% up to 25 mg/dL of bilirubin | <10% up to 500 mg/dL of Intralipid. | No Biotin in assay formulation | N/A | N/A | NIST SRM 927
Trig | 10 - 550 mg/dL | 1100 mg/dL | 2 | Normal:150 mg/dL Borderline high: 150-199 mg/dL High: 200-499 mg/dL Very high: ≥ 500 mg/dL | Atellica CH CHEM CAL | 2 days | 2 | 60 | 14 | Yes | N/A |  | 2000 | 4 x 500 | 4 | 30 | GPO | <10% up to 1000 mg/dL of hemoglobin | <10% up to 30 mg/dL of bilirubin | N/A | No Biotin in assay formulation | N/A | N/A | NIST reference material SRM909c
UA (serum) | 0.5 - 20 mg/dL | 100.0 mg/dL | 5 | Male: 3.7–9.2 mg/dL Female: 3.1–7.8 mg/dL | Atellica CH CHEM CAL | 2 days | 2 | 183 | 7 | Yes | N/A |  | 4800 | 4 x 1200 | 7 | 30 | Uricase/Peroxidase | <10% up to 250 mg/dL of hemoglobin | <10% up to 30 mg/dL of bilirubin | <10% up to 1000 mg/dL of Trig. | No Biotin in assay formulation | N/A | N/A | CDC candidate reference method which uses SRM 913 and SRM 909 reference materials from the National Institute of Standards and Technology
UN_c (serum/plasma) | 5 - 150 mg/dL | 300 mg/dL | 2 | 9-23 mg/dL | Atellica CH CHEM CAL | 2 days | 2 | 75 | 6 | Yes | N/A |  | 6240 | 4 x 1560 | 7 | 90 | Urease with GLDH | <10% up to 500 mg/dL of hemoglobin | <10% up to 30 mg/dL of bilirubin | <10% up to 650 mg/dL of Intralipid. | No Biotin in assay formulation | N/A | N/A | traceable to the CDC reference method, which uses SRM 912 and 909 reference materials from the National Institute of Standards and Technology (NIST).
wrCRP | 0.050–15.600 mg/dL | 0.050– 62.400 mg/dL | 4 | 0.000–0.500 mg/dL | Atellica CH (wrCRP CAL) | 28 days | 6 | 61 | 51 | Yes | NA |  | 2000 | 4 x 500 | 8 | 51 | Latex-enhanced immunoturbidimetric assay | ≤ 10% change up to 500 mg/dL of hemoglobin | ≤ 10% change up to 30 mg/dL of conjugated bilirubin and 30 mg/dL unconjugated bilirubin. | ≤ 10% change up to 500 mg/dL of Intralipid | No Biotin in assay formulation | 95.000 mg/dL | 95.000 mg/dL | IRMM reference material ERM-DA474/IFCC
Vanc | 3.0 - 50.0 µg/mL | 3.0 - 100.0 μg/mL | 2 | Peak 18 - 26 µg/mL Trough 5 - 10 µg/mL | Atellica CH DRUG CAL II | 30 per pack | 5 | 30 | 7 | Yes |  |  | 400 | 4 x 100 | 9

## 면역 참고치
Atellica® IM
항목 | 고신대학교복음병원 현재 참고치 | 신규 장비 참고치 | Expected Values #1 | Expected Values #2 | Expected Values #3 | N
Cor | 4.3-22.4 | 3.44–22.45 | NA | NA | NA | 120
CpS | 0.5-3 | 0.81 - 3.85 | NA | NA | NA | 100
Fer | 22-222 | Normal Males 22–322 Normal Females 10–291 |  |  |  | 100
Fol | 3-17 | 3.38–17 |  |  |  | 100
HCY | 4-15.4 | 3.7 – 13.9 | NA | NA | NA | 200
TSTII | M: 241-827 F: 14-76 | M: 187.72-684.19 F: 7-35.92 | See IFUs |  |  | 100
VB12 | 200-950 | 211-911 | NA | NA | NA | 450
FSH | - | 여포기: 2.5-10.2 배란기: 3.4-33.4 황체기: 1.5-9.1 임신: <0.3 폐경후: 23-116.3 남성: 13-70세: 1.4-18.1 | See IFU |  |  | 90
LH | - | 여포기: 1.9-12.5 배란기: 8.7-76.3 황체기: 0.5-16.9 임신: 0-1.1.5 남성: 20-70세: 1.5-9.3 | See IFU |  |  | 200
PRL | 0-25 | 비임신: 2.8-29.2 임신: 9.7-208.5 폐경후: 1.8-20.3 남성 2.1-17.7 | See IFU |  |  | 200

## 화학 참고치
검사 항목 | 변경 전 참고치 | 변경 후 참고치
Alb | 3.5-5 | 3.2–4.8
ALP | 25-100 | 46–116
ALT | 3-40 | 10–49
AMYLASE | 30-110 | 30–118
ASO | 0-240 | ≤ 200.0
AST | 5-40 | <34
C3 | 90-180 | 82–170
C4 | 10-40 | 12.0–36.0
Ca | 8-10 | 8.3–10.6
Cl (chloride) | 98-110 | 98–107
CO2 (Cabon dioxide) | 23-29 | 20.0–31.0
Creatinine | Males 0.80–1.30 Females 0.6–1.1 | Males 0.70–1.30 Females 0.55–1.02
Direct Bilirubin | 0-0.4 | 0-0.3
GGT | 0-50 | males < 73 U/L females < 38 U/L
IgA | 90-410 | 40.0-350.0
IgG | 700-1600 | 650-1,600
IgM | 40-230 | 50.0-300.0
Phosphorus | 3-4.5 | 2.4-5.1
Iron | Male 50-150 μg/dL Female 35-145 μg/dL | Male 65-175 Female 50-170
K | 3.5-5.3 | 3.5–5.1
Lipase | 5-50 | 12 - 53
Na | 136-150 | 136–145
RF | 0-12.5 | < 14
TIBC | 250-450 | 250–425
TP | 6.3-8.3 | 5.7-8.2
Uric acid | 3-8 | Male: 3.7–9.2 Female: 3.1–7.8
BUN | 5-23 | 9-23
hsCRP | 0-0.75 | 0.000–0.500
LDLP(LDH) | 200-450 | 120 - 246